| WORLD METEOROLOGICAL ORGANIZATION __________________ COMMISSION FOR INSTRUMENTS AND METHODS OF OBSERVATION EXPERT TEAM ON OPERATIONAL METROLOGY First Session Ljubljana, Slovenia 1 – 4 December 2015 |  | CIMO/ ET-OpMet-1/Doc. 7(2) (20.XI.2015) _______ ITEM: 7 Original: ENGLISH |
| --- | --- | --- |

UPDATE OF THE CIMO GUIDE

Draft update of Part I, Chapter 2 “Measurement of Temperature”

(Submitted by Tilman Holfelder)

Summary and purpose of document This document provides the proposed modification to the CIMO Guide Part I, Chapter 2 ““Measurement of Temperature” for review by the CIMO ET-OpMet members.

Action proposed

The Meeting is invited to review this document and to agree on the text to be submitted to the CIMO Editorial Board for the forthcoming update of the CIMO Guide.

________________

CHAPTER 2. Measurement of temperature

2.1	General

2.1.1	Definition

Thermodynamic temperature, T, is a physical quantity characterizing the average energy of random molecular motion within a substance. Direct measurement of T using so-called primary thermometers is experimentally very difficult, and is only intermittently carried out even at National Measurement Institutes. Instead, the Consultative Committee on the Thermometry (CCT) of the International Bureau of Weights and Measures (BIPM) recommend the use of  the International Temperature Scale of 1990 (ITS-90) to produce practical approximations to thermodynamic temperature (BIPM 1989,1990). ITS-90 summarises our knowledge of primary thermometry in 1990 and recommends the value of freezing points, melting points, or triple-points of pure substances. In the temperature range of meteorological interest, (-95 °C to +60 °C), ITS-90 specifies the way in which the electrical resistance of pure platinum varies in between these fixed point temperatures. The approximations to thermodynamic temperature produced by ITS-90 have been recently shown to be in error by less than ±0.01 °C over the entire range of meteorological interest.

For meteorological purposes, temperatures are measured for a number of media. The most common variable measured is air temperature (at various heights). Other variables are ground, soil, grass minimum and seawater temperature. WMO (1992) defines air temperature as “the temperature indicated by a thermometer exposed to the air in a place sheltered from direct solar radiation”. Although this definition cannot be used as the definition of the thermodynamic quantity itself, it is suitable for most applications.

2.1.2	Units and scales

The thermodynamic temperature (T), with units of kelvin (K) (also defined as “kelvin temperature”), is the basic temperature. The kelvin is the fraction 1/273.16 of the thermodynamic temperature of the triple point of water. The temperature (t), in degrees Celsius (or “Celsius temperature”) defined by equation 2.1, is used for most meteorological purposes (from the ice-point secondary reference in Table 2.2 in the annex):

t/°C = T/K – 273.15	(2.1)

A temperature difference of one degree Celsius (°C) is equal to one kelvin (K). Note that the symbol for the kelvin, K is used without the degree symbol.

In the thermodynamic scale of temperature, measurements are expressed as differences from absolute zero (0 K), the temperature at which the molecules of any substance possess no thermal energy. ITS-90 provides a practical approximation to thermodynamic temperature (see the annex), which is based on assigned values for the temperatures of a number of reproducible equilibrium states (see Table 2.1 in the annex) and on specified standard instruments calibrated at those temperatures.. In addition to the defining fixed points of the ITS-90, other secondary reference points are available (see Table 2.2 in the annex)..

2.1.3	Meteorological requirements

2.1.3.1	General

Meteorological requirements for temperature measurements primarily relate to the following:

(a)	The air near the Earth’s surface;

(b)	The surface of the ground;

(c)	The soil at various depths;

(d)	The surface levels of the sea and lakes;

(e)	The upper air.

These measurements are required, either jointly or independently and locally or globally, for input to numerical weather prediction models, for hydrological and agricultural purposes, and as indicators of climatic variability. Local temperature also has direct physiological significance for the day-to-day activities of the world’s population. Measurements of temperature may be required as continuous records or may be sampled at different time intervals. This chapter deals with requirements relating to (a), (b) and (c).

2.1.3.2	Accuracy requirements

The range, reported resolution and required uncertainty for temperature measurements are detailed in Part I, Chapter 1, of this Guide. In practice, it may not be economical to provide thermometers that meet the required performance directly. Instead, cheaper thermometers, calibrated against a laboratory standard, are used with corrections being applied to their readings as necessary. It is necessary to limit the size of the corrections to keep residual errors within bounds. Also, the operational range of the thermometer will be chosen to reflect the local climatic range. As an example, the table below gives an acceptable range of calibration and errors for thermometers covering a typical measurement range.

Example of possible thermometer characteristics

| Thermometer type | Ordinary | Maximum | Minimum |
| --- | --- | --- | --- |
| Span of scale (˚C) | –30 to 45 | –30 to 50 | – 40 to 40 |
| Range of calibration (˚C) | –30 to 40 | –25 to 40 | –30 to 30 |
| Maximum error | < 0.2 K | 0.2 K | 0.3 K |
| Maximum difference between maximum and minimum correction within the range | 0.2 K | 0.3 K | 0.5 K |
| Maximum variation of correction within any interval of 10 ˚C | 0.1 K | 0.1 K | 0.1 K |

All temperature-measuring instruments should be issued with a certificate confirming compliance with the appropriate uncertainty or performance specification, or a calibration certificate that gives the corrections that must be applied to meet the required uncertainty. This initial testing and calibration should be performed by an accredited calibration laboratory. Temperature-measuring instruments should also be checked subsequently at regular intervals, the exact apparatus used for this calibration being dependent on the instrument or sensor to be calibrated.

2.1.3.3	Response times of thermometers

For routine meteorological observations there is no advantage in using thermometers with a very small time-constant or lag coefficient, since the temperature of the air continually fluctuates up to one or two kelvin within a few seconds. Thus, obtaining a representative reading with such a thermometer would require taking the mean of a number of readings, whereas a thermometer with a larger time-constant tends to smooth out the rapid fluctuations. Too long a time constant, however, may result in errors when long-period changes of temperature occur. It is recommended that the time constant, defined as the time required by the thermometer to register 63.2% of a step change in air temperature, should be 20 s. The time constant depends on the airflow over the sensor.

2.1.3.4	Recording the circumstances in which measurements are taken

Temperature is one of the meteorological quantities whose measurements are particularly sensitive to exposure. For climate studies in particular, temperature measurements are affected by the state of the surroundings, by vegetation, by the presence of buildings and other objects, by ground cover, by the condition of, and changes in, the design of the radiation shield or screen, and by other changes in equipment (WMO, 2011). It is important that records should be kept not only of the temperature data, but also of the circumstances in which the measurements are taken. Such information is known as metadata (data about data; see Part I, Chapter 1, Annex 1.C).

2.1.4	Measurement methods

In order to measure the temperature of an object, a thermometer can be brought to the same temperature as the object (namely, into thermodynamic equilibrium with it), and the temperature of the thermometer itself can then be measured. Any physical property of a substance which is a function of temperature can be used as the basis of a thermometer. The properties most widely used in meteorological thermometers are thermal expansion and the change in electrical resistance with temperature. Radiometric thermometers operate in the infrared part of the electromagnetic spectrum and are used, among other applications, for temperature measurements from satellites. A special technique to determine the air temperature using ultrasonic sampling, developed to determine air speeds, also provides the average speeds of the air molecules, and as a consequence its temperature (WMO, 2002a).

Thermometers which indicate the prevailing temperature are often known as ordinary thermometers, while those which indicate extreme temperature over a period of time are called maximum or minimum thermometers.

There are various standard texts on instrument design and laboratory practice for the measurement of temperature, such as Jones (1992) and Middleton and Spilhaus (1960). Considering the concepts of thermometry, care should be taken that, for meteorological applications, only specific technologies are applicable because of constraints determined by the typical climate or environment.

2.1.4.1	Thermometer exposure and siting

Radiation from the sun, clouds, the ground and other surrounding objects passes through the air without appreciably changing its temperature, but a thermometer exposed freely in the open can absorb considerable radiation. As a consequence, its temperature may differ from the true air temperature, with the difference depending on the radiation intensity and on the ratio of absorbed radiation to dissipated heat. For some thermometer elements, such as the very fine wire used in an open-wire resistance thermometer, the difference may be very small or even negligible. However, with the more usual operational thermometers the temperature difference may reach 25 K under extremely unfavourable conditions. Therefore, in order to ensure that the thermometer is at true air temperature it is necessary to protect the thermometer from radiation by a screen or shield that also serves to support the thermometer. This screen also shelters it from precipitation while allowing the free circulation of air around it, and prevents accidental damage. Precipitation on the sensor will, depending on the local airflow, depress the sensor temperature, causing it to behave as a wet-bulb thermometer. Maintaining free circulation may, however, be difficult to achieve under conditions of rime ice accretion. Practices for reducing observational errors under such conditions will vary and may involve the use of special designs of screens or temperature-measuring instruments, including artificial ventilation. Nevertheless, in the case of artificial ventilation, care should be taken to avoid unpredictable influences caused by wet deposition in combination with evaporation during precipitation, drizzle, fog, and the like. An overview of concepts of temperature measurement applicable for operational practices is given by Sparks (1970).

In order to achieve representative results when comparing thermometer readings at different places and at different times, a standardized exposure of the screen and, hence, of the thermometer itself is also indispensable. For general meteorological work, the observed air temperature should be representative of the free air conditions surrounding the station over as large an area as possible, at a height of between 1.25 and 2 m above ground level. The height above ground level is specified because large vertical temperature gradients may exist in the lowest layers of the atmosphere. The best site for the measurements is, therefore, over level ground, freely exposed to sunshine and wind and not shielded by, or close to, trees, buildings and other obstructions. Sites on steep slopes or in hollows are subject to exceptional conditions and should be avoided. In towns and cities, local peculiarities are expected to be more marked than in rural districts. Temperature observations on the top of buildings are of doubtful significance and use because of the variable vertical temperature gradient and the effect of the building itself on the temperature distribution.

The siting classification for surface observing stations on land (see Part I, Chapter 1, Annex 1.B of this Guide) provides additional guidance on the selection of a site and the location of a thermometer within a site to optimize representativeness.

2.1.4.2	Measuring soil temperatures

The standard depths for soil temperature measurements are 5, 10, 20, 50 and 100 cm below the surface; additional depths may be included. The site for such measurements should be a level plot of bare ground (about 75 cm2) and typical of the surrounding soil for which information is required. If the surface is not representative of the general surroundings, its extent should not be less than 100 m2. When the ground is covered with snow, it is desirable to measure the temperature of the snow cover as well. Where snow is rare, the snow may be removed before taking the readings and then replaced.

When describing a site for soil temperature measurements, the soil type, soil cover and the degree and direction of the ground’s slope should be recorded. Whenever possible, the physical soil constants, such as bulk density, thermal conductivity and the moisture content at field capacity, should be indicated. The level of the water table (if within 5 m of the surface) and the soil structure should also be included.

At agricultural meteorological stations, the continuous recording of soil temperatures and air temperatures at different levels in the layer adjacent to the soil (from ground level up to about 10 m above the upper limit of prevailing vegetation) is desirable.

2.1.4.3	Temperature standards

Laboratory standards

Every determination of T is made in terms of W, the ratio of resistance R at T to that at the triple point of water: W(T) =R(T)/R(273.16 K). The ITS-90 defines a reference function Wr(T), derived from measurements of a group of high quality PRTs, and the functional form of deviation functions, which approximate the difference between a given PRT and the reference function. To ensure that the thermometric platinum is of sufficient purity that W(T) is close to Wr(T), the ITS-90 specifies a minimum value of W at the gallium melting point (29.7646 °C), and a maximum value at the mercury triple point (-38.8344 °C). Thermometers which fulfill these criteria are known as standard PRTs (SPRTs). Calibrating an SPRT involves measuring W(T) at a number of fixed-points and calculating the coefficients of the deviation function describing the difference from the ideal reference function Wr.. A national meteorological or other accredited calibration laboratory must have, as a working standard, a high- quality platinum resistance thermometer, traceable to the national standard (WMO 2010). The behaviour of this thermometer may be checked periodically in a water triple-point cell. Temperature of the the triple point of water is defined exactly and can be reproduced in a triple-point cell with an uncertainty of 0.001 K.

Travelling standards

The WMO reference psychrometer (WMO, 1992) is the reference instrument for determining the relationship between the air temperature measured by conventional surface instruments and the true air temperature. This instrument has been designed to be used as a free-standing instrument and not for deployment within a screen or shelter; it is the most accurate instrument available for evaluating and comparing instrument systems. It is not intended for continuous use in routine meteorological operations and is capable of providing a temperature measurement with an uncertainty of 0.04 K (at the 95% confidence level). See Part I, Chapter 4, for further information. Specification for modern travelling standards can be found in WMO 2010.

2.1.5	Implications of the Minamata Convention for the temperature measurement

The United Nations Environment Programme (UNEP) Minamata Convention on Mercury comes into force globally in 2020, and bans all production, import and export of mercury-in-glass thermometers (Part I, Chapter 1, 1.4.1). Therefore mercury-in-glass thermometers are no longer recommended and it is strongly encouraged to take appropriate measures to replace mercury-in-glass thermometers with modern alternatives.  Electrical resistance thermometers are state of the art and can provide an economical, accurate and reliable alternative to their dangerous, mercury-based precedents and offer significant advantages in terms of data storage and real-time data display.

2.2	Electrical thermometers

2.2.1	General description

Electrical instruments are in widespread use in meteorology for measuring temperatures. Their main virtue lies in their ability to provide an output signal suitable for use in remote indication, recording, storage, or transmission of temperature data. The most frequently used sensors are electrical resistance elements, semiconductor thermometers (thermistors) and thermocouples.

2.2.1.1	Electrical resistance thermometers

A measurement of the electrical resistance of a material whose resistance varies in a known manner with the temperature of the material can be used to represent the temperature.

For small temperature changes, the increase in resistance of pure metals is proportional to the change in temperature, as expressed in equation 2.2:

RT = R0 [1 + α (T – T0)]	(2.2)

where (T – T0) is small; RT is the resistance of a fixed amount of the metal at temperature T; R0 is its resistance at a reference temperature T0, and ( is the temperature coefficient of resistance in the vicinity of T0.

With 0 °C as the reference temperature, equation 2.2 becomes:

RT = R0 (1 + α · t)	(2.3)

For larger temperature changes and for certain metallic alloys, equation 2.4 expresses the relationship more accurately:

RT = R0 [1 + α (T – T0) + β (T – T0)2]	(2.4)

With 0 °C as the reference temperature, equation 2.4 becomes:

RT = R0 (1 + α · t + β · t2)	(2.5)

These equations give the proportional change in resistance of an actual thermometer, so that values for the coefficients α and β can be found by calibration of the thermometer concerned. Based on these results, the inverse function, namely, t as a function of R, can be derived. Such a function may be expressed in terms of a power series of (R0 – RT), namely, t = t (R0 – RT) = c1 (R0 – RT) + c2 (R0 – RT)2 + …

A good metal resistance thermometer will satisfy the following requirements:

(a)	Its physical and chemical properties will remain the same through the temperature measurement range;

(b)	Its resistance will increase steadily with increasing temperature without any discontinuities in the range of measurement;

(c)	External influences such as humidity, corrosion or physical deformations will not alter its resistance appreciably;

(d)	Its characteristics will remain stable over a period of two years or more;

(e)	Its resistance and thermal coefficient should be large enough to be useful in a measuring circuit.

Pure platinum best satisfies the foregoing requirements. Thus, it is used for the SPRTs needed for transferring the ITS-90 between instrument locations. Platinum with lower purity is also used for secondary standards and for operational sensors.

Practical thermometers are artificially aged before use and are commonly made from platinum alloys, nickel and occasionally tungsten for meteorological purposes. Usually they are hermetically sealed in a ceramic sheath. Their time constant is smaller than that of liquid-in-glass thermometers.

2.2.1.2	Semiconductor thermometers

Another type of resistance element in common use is the thermistor. This is a semiconductor with a relatively large temperature coefficient of resistance, which may be either positive or negative depending upon the actual material. Mixtures of sintered metallic oxides are suitable for making practical thermistors, which usually take the form of small discs, rods or spheres and are often glass-coated. The general expression for the temperature dependence of the resistance, R, of the thermistor is given in equation 2.6:

R = a exp (b/T)	(2.6)

(2.6)

where a and b are constants and T is the temperature of the thermistor in kelvin. For lower uncertainty the Steinhart-Hart equation is also commonly used

where A, B and C are determined by calibration.

The advantages of thermistors from a thermometric point of view are as follows:

(a)	The large temperature coefficient of resistance enables the voltage applied across a resistance bridge to be reduced while attaining the same sensitivity, thus reducing or even eliminating the need to account for the resistance of the leads and its changes;

(b)	The elements can be made very small, so their very low thermal capacities can yield a small time-constant. However, very small thermistors with their low thermal capacity have the disadvantage that, for a given dissipation, the self-heating effect is greater than for large thermometers. Thus, care must be taken to keep the power dissipation small.

A typical thermistor can have a resistance which varies by a factor of 100 or 200 over the temperature range –40 °C to 40 °C.

2.2.1.3	Thermocouples

The Seebeck effect describes the phenomenon whereby a temperature gradient in a metal gives rise to  an accompanying electric field. The magnitude of the accompanying electric field is always proportional to the temperature gradient but depends on a material-dependent and temperature-dependent Seebeck Coefficient.

Thus a wire of pure material in a temperature gradient spontaneously acquires a voltage across its ends, the magnitude of which is equal to the integrated Seebeck voltage along the wire. A thermocouple is made from two wires of different materials with differing Seebeck coefficients joined at one end – the so-called ‘hot’ junction. The open ends of the wires are connected to a sensitive voltage measuring device in one of a variety of configurations.

Typically in operation the open ends of the junctions are connected to the terminals of a high-resolution voltmeter.

Thermocouples are made by welding or soldering together wires of the metals concerned. These junctions can be made very small and with negligible heat capacity.

When used to measure temperature, a measurement is taken of the electromotive force set up when one junction is maintained at a standard known temperature – usually the ice point - and the other junction is allowed to take the temperature whose value is required. This electromotive force can be directly related to the difference in temperature between the two junctions by previous calibration of the system, and thus the unknown temperature is found by adding this difference algebraically to the known standard temperature.

In practice it is inconvenient to maintain an ice point and instead the measurements are referenced to the temperature of the terminals of the digital voltmeter. This technique – known as cold-junction compensation – requires a measurement of the temperature of the voltmeter terminals using a thermistor or platinum resistance thermometer. The additional thermocouple voltage that would have been expected if an ice-point had been used is then calculated and added to the measured voltage. The sum is then used to determine the temperature using interpolation of standard tables. Where cold-junction compensation is used, especial care must be taken close to voltmeter junctions where small temperature differences between the terminals can generate spurious voltages.

In meteorology, thermocouples are mostly used when a thermometer of very small time-constant, of the order of 1 or 2 s, and capable of remote reading and recording is required, usually for special research tasks. A disadvantage, if the absolute temperature is required, is the necessity for a constant-temperature enclosure for both the cold junction and ancillary apparatus for the measurements of the electromotive force that has been set up; thermocouples are best suited for the measurement of differential temperatures, since this complication does not arise. Very high accuracy can be achieved with suitably sensitive apparatus, but frequent calibration is necessary. Copper-constantan or iron-constantan combinations are suitable for meteorological work, as the electromotive force produced per degree Celsius is higher than with rarer and more expensive metals, which are normally used at high temperatures.

2.2.2	Measurements procedures

2.2.2.1	Electrical resistance thermometers

Electrical resistance thermometers may be connected to a variety of electrical measurement circuits, many of which are variations of resistance bridge circuits in either balanced or unbalanced form. In a balanced bridge, an accurate potentiometer is adjusted until no current flows in an indicator, with the position of the potentiometer arm being related to the temperature. In an unbalanced bridge, the out-of-balance current may be measured by a galvanometer; however, this current is not simply a function of the temperature and depends in part on other effects. An alternative which avoids this situation is to use a constant current source to power the bridge and to measure the out-of-balance voltage to obtain the temperature reading.

In the case of remote measuring, it should be taken into consideration that the wire between the resistance thermometer and the bridge also forms a resistance that alters depending on the temperature. Suitable precautions can be taken to avoid such errors.

Digital voltmeters can be used in conjunction with a constant current source to measure the temperature-dependent voltage drop across the thermometer element; the output can be scaled directly in temperature. Also, the digital output can be stored or transmitted without loss of accuracy and, thus, be available for further use. The digital output of the digital voltmeters can be subsequently converted back to an analogue voltage, if desired, to feed a recorder, for example.

2.2.2.2	Thermocouples

The voltage produced by a thermocouple is small, typical 40 µV °C-1 and may need to be amplified before being read by a digital voltmeter or analogue-to-digital converter (ADC). In addition to knowing the voltage produced by the thermocouple, it is also necessary to know the temperature of the reference junctions at which the thermocouple connects to the terminals of the digital voltmeter or ADC. This secondary temperature measurement is typically made by reading a resistance thermometer or thermistor positioned close to the terminals. Thermocouple reading devices typically combine this temperature measurement with knowledge of the thermocouple type to output a digital reading corresponding either to the temperature (in °C or K), or the voltage which would have been generated if the thermocouple reference junctions were at 0 °C.

Diodes and Bandgap devices:

2.2.3	Exposure and siting

The general requirements relating to the exposure and siting of electrical thermometers are described in 2.1.4.1. Additional requirements include the following:

(a)	The measurement of extreme values: Separate maximum and minimum thermometers may no longer be required if the electrical thermometer is connected to a continuously operating data recording system;

(b)	The measurement of surface temperatures: The radiative properties of electrical thermometers will be different from liquid-in-glass thermometers. Electrical thermometers exposed as grass minimum (or other surface) thermometers will, therefore, record different values from similarly exposed conventional thermometers. These differences may be minimized by placing the electrical thermometer within a glass sheath;

(c)	The measurement of soil temperatures: Electrical thermometers are deployed in brass plugs, inserted at the required depth into an undisturbed vertical soil face, the latter having been exposed by trenching. Electrical connections are brought out through plastic tubes via the trench, which is then refilled in such a way to restore, as far as possible, the original strata and drainage characteristics.

2.2.4	Sources of error

2.2.4.1	Electrical resistance thermometers

The main sources of error in a temperature measurement taken with electrical resistance thermometers are the following:

(a)	Self-heating of the thermometer element;

(b)	Inadequate compensation for lead resistance;

(c)	Inadequate compensation for non-linearities in the sensor or processing instrument;

(d)	Sudden changes in switch contact resistances.

Self-heating occurs because the passage of a current through the resistance element produces heat and, thus, the temperature of the thermometer element becomes higher than that of the surrounding medium. For a 1 mA (10 mA) current in a PT100 sensor the heating is approximately 0.1 mW (10 mW) and sensor heating varies with typical values between 0.01 °C and 0.05 °C.

The resistance of the connecting leads will introduce an error in the temperature reading. This will become more significant for long leads, for example, when the resistance thermometer is located at some distance from the measuring instrument; the reading errors will also vary as the temperature of the cables changes. These errors can be compensated for by using extra conductors, ballast resistors and an appropriate bridge network. To reduce errors, it is highly recommended to use four-wire platinum resistance thermometers.

PICTURE of 4-Wire configuration

Neither the electrical resistance thermometer nor the thermistor is linear over an extended temperature range but may approximate a linear output if the range is limited. Provision must, therefore, be made to compensate for such non-linearities. This is most likely to be required for thermistors, to achieve a usable meteorological range of measurement.

Sudden changes in switch contact resistance can occur as switches age. They may be variable and can go undetected unless regular system calibration checks are performed (see section 2.4.5).

2.2.4.2	Thermocouples

The main sources of error in the measurement of temperature using thermocouples are the following:

(a)	 Just wrong(b)	Conduction along the leads from the junction when there is a temperature gradient in the vicinity of the temperature measuring point;Correct, but generally a small effect

(c)	Stray secondary thermal electromotive forces due to the use of metals that are different from the thermocouple metals in the connecting circuit. The temperature differences in the remainder of the circuit must, therefore, be kept as low as possible; this is especially important when the electromotive forces to be measured are small (periodical recalibration will be necessary to allow for this);

(d)	Leakage currents can occur from neighbouring power circuits. This can be minimized by suitable screening of the leads;

(e)	Galvanic currents can be set up if any leads or junctions are allowed to get wet;

(f)	Changes in temperature in the galvanometer alter its characteristics (chiefly by changing its resistance). This will not affect the readings by the potentiometric method to any degree, but will affect direct-reading instruments. This effect can be minimized by keeping the temperature of the galvanometer as near as possible to that at which the circuit was calibrated;

(g)	In the potentiometric measurement, changes in the electromotive force of the standard cell against which the potentiometer current is adjusted and changes in the potentiometer current between adjustments will cause corresponding errors in the measured electromotive force. These errors will normally be small, provided that the standard cell is treated correctly, and that adjustments of the potentiometer current are made just before taking a temperature measurement.

Errors (a) and (f) emphasize the superiority of the potentiometric method when a very high degree of accuracy is required.

2.2.5	Comparison and calibration

2.2.5.1	Electrical resistance thermometers

Laboratory calibrations of thermometers should be carried out by accredited calibration laboratories. Thermometers should be compared against reference thermometers in either a dry-block calibrator or a stirred liquid bath..

Since the measurement instrument is an integral part of the electrical thermometer, its calibration should be checked by substituting the resistance thermometer by an accurate decade resistance box and by applying resistances equivalent to fixed 5 K temperature increments over the operational temperature range. The error at any point should not exceed 0.1 K.

Field inspections of the electrical thermometers should be performed as described in Part III, chapter 4, paragraph 4.3.5.

2.2.5.2	Thermocouples

The calibration and checking of thermocouples require the hot and cold junctions to be maintained at accurately known temperatures. The techniques and instrumentation necessary to undertake this work are generally very specialized and will not be described here.

2.2.6	Corrections

When initially issued, electrical thermometers (which have a serial number) should be provided with either:

(a)	A dated certificate confirming compliance with the appropriate standard; or

(b)	A dated calibration certificate giving the actual resistances at different temperatures in the temperature range. These resistances should be used when checking the uncertainty of the measuring instrument or system interface before and during operation. The magnitude of the resistance difference from the nominal value should not, in general, be greater than an equivalent temperature error of 0.1 or 0.2 K.

2.2.7	Maintenance

The regular field checks should identify any changes in system calibration. These may occur as a result of long-term changes in the electrical characteristics of the thermometer, degradation of the electrical cables or their connections, changes in the contact resistance of switches or changes in the electrical characteristics of the measuring equipment. Identification of the exact source and correction of such errors will require specialized equipment and training and should be undertaken only by a maintenance technician.

2.3	Mechanical thermographs

2.3.1	General description

The types of mechanical thermographs still commonly used are supplied with bimetallic or Bourdon-tube sensors since these are relatively inexpensive, reliable and portable. However, they are not readily adapted for remote or electronic recording. Such thermographs incorporate a rotating chart mechanism common to the family of classic recording instruments. In general, thermographs should be capable of operating over a range of about 60 K or even 80 K if they are to be used in continental climates. A scale value is needed such that the temperature can be read to 0.2 K without difficulty on a reasonably sized chart. To achieve this, provisions should be made for altering the zero setting of the instrument according to the season. The maximum error of a thermograph should not exceed 1 K.

2.3.1.1	Bimetallic thermograph

In bimetallic thermographs, the movement of the recording pen is controlled by the change in curvature of a bimetallic strip or helix, one end of which is rigidly fixed to an arm attached to the frame. A means of finely adjusting this arm should be provided so that the zero of the instrument can be altered when necessary. In addition, the instrument should be provided with a means of altering the scale value by adjusting the length of the lever that transfers the movement of the bimetal to the pen; this adjustment is best left to authorized personnel. The bimetallic element should be adequately protected from corrosion; this is best done by heavy copper, nickel or chromium plating, although a coat of lacquer may be adequate in some climates. A typical time-constant of about 25 s is obtained at an air speed of 5 m s–1.

2.3.1.2	Bourdon-tube thermograph

The general arrangement is similar to that of the bimetallic type but its temperature-sensitive element is in the form of a curved metal tube of flat, elliptical section, filled with alcohol. The Bourdon tube is less sensitive than the bimetallic element and usually requires a multiplying level mechanism to give sufficient scale value. A typical time-constant is about 6 s at an air speed of 5 m s–1.

2.3.2	Measurement procedures

In order to improve the resolution of the reading, thermographs will often be set, in different seasons, to one of two different ranges with corresponding charts. The exact date for changing from one set of charts to the other will vary according to the locality. However, when the change is made the instrument will need to be adjusted. This should be done either in the screen on a cloudy, windy day at a time when the temperature is practically constant or in a room where the temperature is constant. The adjustment is made by loosening the screw holding the pen arm to the pen spindle, moving the pen arm to the correct position and retightening, the screws. The instrument should then be left as is before rechecking, and any further adjustments made as necessary.

2.3.3	Exposure and siting

These instruments should be exposed in a large thermometer screen.

2.3.4	Sources of error

In the thermograph mechanism itself, friction is the main source of error. One cause of this is bad alignment of the helix with respect to the spindle. Unless accurately placed, the helix acts as a powerful spring and, if rigidly anchored, pushes the main spindle against one side of the bearings. With modern instruments this should not be a serious problem. Friction between the pen and the chart can be kept to a minimum by suitably adjusting the gate suspension.

2.3.5	Comparison and calibration

2.3.5.1	Laboratory calibration

There are two basic methods for the laboratory calibration of bimetallic thermographs. They may be checked by fixing them in a position with the bimetallic element in a bath of water. Alternatively, the thermograph may be placed in a commercial calibration chamber equipped with an air temperature control mechanism, a fan and a reference thermometer.

Comparisons should be made at two temperatures; from these, any necessary changes in the zero and magnification can be found. Scale adjustments should be performed by authorized personnel, and only after reference to the appropriate manufacturer’s instrument handbook.

2.3.5.2	Field comparison

The time constant of the instrument may be as low as one half that of the ordinary mercury-in-glass thermometer, so that routine comparisons of the readings of the dry bulb and the thermograph at fixed hours will, in general, not produce exact agreement even if the instrument is working perfectly. A better procedure is to check the reading of the instrument on a suitable day at a time when the temperature is almost constant (usually a cloudy, windy day) or, alternatively, to compare the minimum readings of the thermograph trace with the reading of the minimum thermometer exposed in the same screen. Any necessary adjustment can then be made by means of the setting screw.

2.3.6	Corrections

Thermographs would not normally be issued with correction certificates. If station checks show an instrument to have excessive errors, and if these cannot be adjusted locally, the instrument should be returned to an appropriate calibration laboratory for repair and recalibration.

2.3.7	Maintenance

Routine maintenance will involve an inspection of the general external condition, the play in the bearings, the inclination of the recording arm, the set of the pen, and the angle between the magnification arm and recording arm, and a check of the chart-drum clock timing. Such examinations should be performed in accordance with the recommendations of the manufacturer. In general, the helix should be handled carefully to avoid mechanical damage and should be kept clean. The bearings of the spindle should also be kept clean and oiled at intervals using a small amount of clock oil. The instrument is mechanically very simple and, provided that precautions are taken to keep the friction to a minimum and prevent corrosion, it should give good service.

2.4	Liquid-in-glass thermometers

Mercury-in-glass thermometers are indeed widespread, but no longer recommended taken into account the Minamata convention on mercury (Part I, Chapter 1, 1.4.2). NMHS are encouraged to take appropriate measures to replace mercury-in-glass thermometers with modern alternatives (e.g. electrical resistance thermometers). Due to genesis of thermometry and the widespread use of mercury-in-glass thermometers the following chapters include also mercury-in-glass thermometers.

2.4.1	General description

For routine observations of air temperature, including maximum, minimum and wet-bulb temperatures, liquid-in-glass thermometers are still commonly used. Such thermometers make use of the differential expansion of a pure liquid with respect to its glass container to indicate the temperature. The stem is a tube which has a fine bore attached to the main bulb; the volume of liquid in the thermometer is such that the bulb is filled completely but the stem is only partially filled at all temperatures to be measured. The changes in volume of the liquid with respect to its container are indicated by changes in the liquid column; by calibration with respect to a standard thermometer, a scale of temperature can be marked on the stem, or on a separate scale tightly attached to the stem.

The liquid used depends on the required temperature range; mercury is generally used for temperatures above its freezing point (–38.9 °C), while ethyl alcohol or other pure organic liquids are used for lower temperatures. The glass should be one of the normal or borosilicate glasses approved for use in thermometers. The glass bulb is made as thin as is consistent with reasonable strength to facilitate the conduction of heat to and from the bulb and its contents. A narrower bore provides greater movement of liquid in the stem for a given temperature change, but reduces the useful temperature range of the thermometer for a given stem length. The thermometer should be suitably annealed before it is graduated in order to minimize the slow changes that occur in the glass with ageing.

There are four main types of construction for meteorological thermometers, as follows:

(a)	The sheathed type with the scale engraved on the thermometer stem;

(b)	The sheathed type with the scale engraved on an opal glass strip attached to the thermometer tube inside the sheath;

(c)	The unsheathed type with the graduation marks on the stem and mounted on a metal, porcelain or wooden back carrying the scale numbers;

(d)	The unsheathed type with the scale engraved on the stem.

The stems of some thermometers are lens-fronted to provide a magnified image of the liquid thread.

Types (a) and (b) have the advantage over types (c) and (d) that their scale markings are protected from wear. For types (c) and (d), the markings may have to be reblackened from time to time; on the other hand, such thermometers are easier to make than types (a) and (b). Types (a) and (d) have the advantage of being less susceptible to parallax errors (see section 2.4.4). An overview of thermometers, designed for use in meteorological practices is given by HMSO (1980).

Whichever type is adopted, the sheath or mounting should not be unduly bulky as this would keep the heat capacity high. At the same time, the sheath or mounting should be sufficiently robust to withstand the normal risks associated with handling and transit.

For mercury-in-glass thermometers, especially maximum thermometers, it is important that the vacuum above the mercury column be nearly perfect. All thermometers should be graduated for total immersion, with the exception of thermometers for measuring soil temperature. The special requirements of thermometers for various purposes are dealt with hereafter under the appropriate headings.

2.4.1.1	Ordinary (station) thermometers

Usually it has been a very accurate mercury-in-glass-type thermometer. Its scale markings have an increment of 0.2 K or 0.5 K, and the scale is longer than that of the other meteorological thermometers.

The ordinary thermometer is used in a thermometer screen to avoid radiation errors. A support keeps it in a vertical position with the bulb at the lower end. The form of the bulb is that of a cylinder or an onion.

A pair of ordinary thermometers can be used as a psychrometer if one of them is fitted with a wet-bulb sleeve.

2.4.1.2	Maximum thermometers

The recommended type for maximum thermometers has been a mercury-in-glass thermometer with a constriction in the bore between the bulb and the beginning of the scale. This constriction prevents the mercury column from receding with falling temperatures. However, observers can reset by holding it firmly, bulb-end downwards, and swinging their arm until the mercury column is reunited. A maximum thermometer should be mounted at an angle of about 2° from the horizontal position, with the bulb at the lower end to ensure that the mercury column rests against the constriction without gravity forcing it to pass. It is desirable to have a widening of the bore at the top of the stem to enable parts of the column which have become separated to be easily united.

2.4.1.3	Minimum thermometers

As regards minimum thermometers, the most common instrument is a spirit thermometer with a dark glass index, about 2 cm long, immersed in the spirit. Since some air is left in the tube of a spirit thermometer, a safety chamber should be provided at the upper end which should be large enough to allow the instrument to withstand a temperature of 50 °C without being damaged. Minimum thermometers should be supported in a similar manner to maximum thermometers, in a near-horizontal position. Various liquids can be used in minimum thermometers, such as ethyl alcohol, pentane and toluol. It is important that the liquid should be as pure as possible since the presence of certain impurities increases the tendency of the liquid to polymerize with exposure to light and after the passage of time; such polymerization causes a change in calibration. In the case of ethyl alcohol, for example, the alcohol should be completely free of acetone.

Minimum thermometers are also exposed to obtain grass minimum temperature.

2.4.1.4	Soil thermometers

For measuring soil temperatures at depths of 20 cm or less, mercury-in-glass thermometers, with their stems bent at right angles, or any other suitable angle, below the lowest graduation, havebeen in common use. The thermometer bulb is sunk into the ground to the required depth, and the scale is read with the thermometer in situ. These thermometers are graduated for immersion up to the measuring depth. Since the remainder of the thermometer is kept at air temperature, a safety chamber should be provided at the end of the stem for the expansion of the mercury.

For measuring temperature at depths of over 20 cm, mercury-in-glass thermometers, mounted on wooden, glass or plastic tubes, with their bulbs embedded in wax or metallic paint, are recommended. The thermometer-tube assemblies are then suspended or slipped in thin-walled metal or plastic tubes sunk into the ground to the required depth. In cold climates, the tops of the outer tubes should extend above the ground to a height greater than the expected depth of snow cover.

The technique of using vertical steel tubes is unsuitable for measuring the diurnal variation of soil temperature, particularly in dry soil, and calculations of soil thermal properties based on such measurements could be significantly in error because they will conduct heat from the surface layer.

The large time-constant due to the increased heat capacity enables the thermometers to be removed from the outer tubes and read before their temperature has had time to change appreciably from the soil temperature.

When the ground is covered by snow, and in order that the observer may approach the line of thermometers without disturbing the snow cover, it is recommended that a lightweight bridge be constructed parallel to the line of thermometers. The bridge should be designed so that the deck can be removed between readings without affecting the snow cover.

2.4.2	Measurement procedures

2.4.2.1	Reading ordinary thermometers

Thermometers should be read as rapidly as possible in order to avoid changes of temperature caused by the observer’s presence. Since the liquid meniscus, or index, and the thermometer scale are not on the same plane, care must be taken to avoid parallax errors. These will occur unless the observer ensures that the straight line from his/her eye to the meniscus, or index, is at a right angle to the thermometer stem. Since thermometer scales are not normally subdivided to less than one fifth of a degree, readings to the nearest tenth of a degree, which are essential in psychrometry, must be made by estimation. Corrections for scale errors, if any, should be applied to the readings. Maximum and minimum thermometers should be read and set at least twice daily. Their readings should be compared frequently with those of an ordinary thermometer in order to ensure that no serious errors develop.

2.4.2.2	Measuring grass minimum temperatures

The grass minimum temperature is the lowest temperature reached overnight by a thermometer freely exposed to the sky just above short grass. The temperature is measured with a minimum thermometer such as that described in section 2.4.1.3. The thermometer should be mounted on suitable supports so that it is inclined at an angle of about 2° from the horizontal position, with the bulb lower than the stem, 25 to 50 mm above the ground and in contact with the tips of the grass. When the ground is covered with snow, the thermometer should be supported immediately above the surface of the snow, as near to it as possible without actually touching it.

Normally, the thermometer is exposed at the last observation hour before sunset, and the reading is taken the next morning. The instrument is kept within a screen or indoors during the day. However, at stations where an observer is not available near sunset, it may be necessary to leave the thermometer exposed throughout the day. In strong sunshine, exposing the thermometer in this way can cause the spirit to distil and collect in the top of the bore. This effect can be minimized by fitting a cotton sock on a black metal shield over the safety chamber end of the thermometer; this shield absorbs more radiation and consequently reaches a higher temperature than the rest of the thermometer. Thus, any vapour will condense lower down the bore at the top of the spirit column.

2.4.3	Thermometer siting and exposure

Both ordinary thermometers and maximum and minimum thermometers are always exposed in a thermometer screen placed on a support. Extreme thermometers are mounted on suitable supports so that they are inclined at an angle of about 2° from the horizontal position, with the bulb being lower than the stem.

The siting and exposure of grass minimum thermometers is as prescribed in section 2.4.2.2. At a station where snow is persistent and of varying depth, it is possible to use a support that allows the thermometers to be raised or lowered to maintain the correct height above the snow surface.

2.4.4	Sources of error in liquid-in-glass thermometers

The main sources of error common to all liquid-in-glass thermometers are the following:

(a)	Elastic errors;

(b)	Errors caused by the emergent stem;

(c)	Parallax and gross reading errors;

(d)	Changes in the volume of the bulb produced by exterior or interior pressure;

(e)	Capillarity;

(f)	Errors in scale division and calibration;

(g)	Inequalities in the expansion of the liquid and glass over the range considered.

The last three errors can be minimized by the manufacturer and included in the corrections to be applied to the observed values. Some consideration needs to be given to the first three errors. Error (d) does not usually arise when the thermometers are used for meteorological purposes.

2.4.4.1	Elastic errors

There are two kinds of elastic errors, namely reversible and irreversible errors. The first is of importance only when a thermometer is exposed to a large temperature range in a short period of time. Thus, if a thermometer is checked at the steam point and shortly afterwards at the ice point, it will read slightly too low at first and then the indicated temperature will rise slowly to the correct value. This error depends on the quality of the glass employed in the thermometer, and may be as much as 1 K (with glass of the highest quality it should be only 0.03 K) and would be proportionately less for smaller ranges of temperature. The effect is of no importance in meteorological measurements, apart from the possibility of error in the original calibration.

The irreversible changes may be more significant. The thermometer bulb tends to contract slowly over a period of years and, thus, causes the zero to rise. The greatest change will take place in the first year, after which the rate of change will gradually decrease. This alteration can be reduced by subjecting the bulb to heat treatment and by using the most suitable glass. Even with glass of the highest quality, the change may be about 0.01 K per year at first. For accurate work, and especially with inspector or check thermometers, the zero should be redetermined at the recommended intervals and the necessary corrections applied.

2.4.4.2	Errors caused by the emergent stem

A thermometer used to measure air temperature is usually completely surrounded by air at an approximately uniform temperature, and is calibrated by immersing the thermometer either completely or only to the top of the mercury column (namely, calibrated by complete or partial immersion). When such a thermometer is used to determine the temperature of a medium which does not surround the stem, so that the effective temperature of the stem is different from that of the bulb, an error will result.

For meteorological applications, the most likely circumstance where this might be encountered is when checking the calibration of an ordinary thermometer in a vessel containing another liquid at a temperature significantly different from ambient temperature and only the bulb or lower part of the stem is immersed.

2.4.4.3	Parallax and gross reading errors

If the thermometer is not viewed on the plane that is perpendicular to the stem of the thermometer, parallax errors will arise. The error increases with the thickness of the thermometer stem and the angle between the actual and the correct line of sight. This error can be avoided only by taking great care when making an observation. With mercury-in-glass thermometers suspended vertically, as in an ordinary screen, the thermometer must be viewed at the horizontal level of the top of the mercury column.

Errors can also occur because observers usually disturb the surroundings in some way when they approach to read the thermometer. It is, therefore, necessary for observers to take the readings to the nearest tenth of a degree as soon as possible. Gross reading errors are usually 1°, 5° or 10° in magnitude. Such errors will be avoided if observers recheck the tens and units figure after taking their initial reading.

2.4.4.4	Errors due to differential expansion

The coefficient of cubical expansion of mercury is 1.82 · 10–4 K–1, and that of most glass lies between 1.0 · 10–5 and 3.0 · 10–5 K–1. The expansion coefficient of the glass is, thus, an important fraction of that of mercury and cannot be neglected. As neither the coefficients of cubical expansion of mercury and glass nor the cross-sectional area of the bore of the stem are strictly constant over the range of temperature and length of the stem being used, the scale value of unit length of the stem varies along the stem, and the thermometer has to be calibrated by the manufacturer against a standard thermometer before it can be used.

2.4.4.5	Errors associated with spirit thermometers

The expansion coefficients of the liquids used in spirit thermometers are very much larger than those of mercury, and their freezing points are much lower (ethyl alcohol freezes at –115 °C). Spirit is used in minimum thermometers because it is colourless and because its larger expansion coefficient enables a larger bore to be used. Spirit thermometers are less accurate than mercury-in-glass thermometers of similar cost and quality. In addition to having the general disadvantages of liquid-in-glass thermometers, spirit thermometers have some peculiarities to themselves:

(a)	Adhesion of the spirit to the glass: Unlike mercury, organic liquids generally wet the glass. Therefore, when the temperature falls rapidly, a certain amount of the liquid may remain on the walls of the bore, causing the thermometer to read low. The liquid gradually drains down the bore if the thermometer is suspended vertically;

(b)	Breaking of the liquid column: Drops of the liquid often form in the upper part of the thermometer stem by a process of evaporation and condensation. These can be reunited with the main column, but errors may be caused at the beginning of the process before it is noticed. The column is also often broken during transport. This error is reduced during manufacture by sealing off the thermometer at its lowest temperature so that it contains the maximum amount of air in the stem;

(c)	Slow changes in the liquid: The organic liquids used tend to polymerize with age and exposure to light, with a consequent gradual diminution in liquid volume. This effect is speeded up by the presence of impurities; in particular, the presence of acetone in ethyl alcohol has been shown to be very deleterious. Great care has therefore to be taken over the preparation of the liquid for the thermometers. This effect may also be increased if dyes are used to colour the liquid to make it more visible.

The reduction of errors caused by breakage in the liquid column and the general care of spirit thermometers are dealt with later in this chapter.

2.4.5	Comparison and calibration in the field and laboratory

2.4.5.1	Laboratory calibration

Laboratory calibrations of thermometers should be carried out by accredited calibration laboratories. For liquid-in-glass thermometers, a liquid bath should be employed, within which it should be possible to maintain the temperature at any desired values within the required range. The rate of temperature change within the liquid should not exceed the recommended limits, and the calibration apparatus should be provided with a means of stirring the liquid. The reference thermometers and thermometers being calibrated should be suspended independently of the container and fully immersed, and should not touch the sides.

Sufficient measurements should be taken to ensure that the corrections to be applied represent the performance of the thermometer under normal conditions, with errors due to interpolation at any intermediate point not exceeding the non-systematic errors (see Part III, Chapter 4).

2.4.5.2	Field checks and calibration

All liquid-in-glass thermometers experience gradual changes of zero. For this reason, it is desirable to check them at regular intervals, usually about once every two years. The thermometers should be stored in an upright position at room temperature for at least 24 h before the checking process begins.

The ice point may be checked by almost filling a Dewar flask with crushed ice made from distilled water and moistening it with more distilled water. The space between the ice pieces as well as the bottom of the vessel should be free from air. The water should remain 2 cm beneath the ice surface. An ordinary Thermos flask will accommodate the total immersion of most thermometers up to their ice point. The thermometers should be inserted so that as little of the mercury or spirit column as possible emerges from the ice. An interval of at least 15 min should elapse to allow the thermometer to take up the temperature of the melting ice before a reading of the indicated temperature is taken. Each thermometer should be moved backwards and forwards through the mixture and immediately read to a tenth part of the scale interval. Further readings at 5 min intervals should be taken and a mean value computed.

Other points in the range can be covered by reference to a travelling standard or inspector thermometer. Comparison should be made by immersing the reference thermometer and the thermometer, or thermometers, to be checked in a deep vessel of water. It is generally better to work indoors, especially if the sun is shining, and the best results will be obtained if the water is at, or close to, ambient temperature.

Each thermometer is compared with the reference thermometer; thermometers of the same type can be compared with each other. For each comparison, the thermometers are held with their bulbs close together, moved backwards and forwards through the water for about 1 min, and then read. It must be possible to read both thermometers without changing the depth of immersion; subject to this, the bulbs should be as deep in the water as possible. Most meteorological thermometers are calibrated in the laboratory for total immersion; provided that the difference between the water and air temperature is not more than 5 K, the emergent stem correction should be negligible. Often, with the bulbs at the same depth, the tops of the columns of mercury (or other liquid) in the reference thermometer and the thermometer being checked will not be very close together. Particular care should therefore be taken to avoid parallax errors.

These comparisons should be made at least three times for each pair of thermometers. For each set of comparisons, the mean of the differences between readings should not exceed the tolerances specified in the table in section 2.1.3.2.

Soil thermometers may be checked in this manner, but should be left in the water for at least 30 min to allow the wax in which the bulbs are embedded to take up the temperature of the water. The large time-constant of the soil thermometer makes it difficult to conduct a satisfactory check unless the temperature of the water can be kept very steady. If the check is carefully carried out in water whose temperature does not change by more than 1 K in 30 min, the difference from the corrected reading of the reference thermometer should not exceed 0.25 K.

2.4.6	Corrections

When initially issued, thermometers identified by a serial number should be provided with either a dated certificate confirming compliance with the uncertainty requirement, or a dated calibration certificate giving the corrections that should be applied to the readings to achieve the required uncertainty.

In general, if the errors at selected points in the range of a thermometer (for example, 0 °C, 10 °C, 20 °C) are all within 0.05 K, no corrections will be necessary and the thermometers can be used directly as ordinary thermometers in naturally ventilated screens and as maximum, minimum, soil or grass minimum thermometers. If the errors at these selected points are greater than 0.05 K, a table of corrections should be available to the observer at the place of reading, together with unambiguous instructions on how these corrections should be applied.

Thermometers for which certificates would normally be issued are those:

(a)	For use in ventilated psychrometers;

(b)	For use by inspectors as travelling standards;

(c)	For laboratory calibration references;

(d)	For special purposes for which the application of corrections is justified.

For psychrometric use, identical thermometers should be selected.

2.4.7	Maintenance

2.4.7.1	Breakage in the liquid column

The most common fault encountered is the breaking of the liquid column, especially during transportation. This is most likely to occur in spirit (minimum) thermometers. Other problems associated with these thermometers are adhesion of the spirit to the glass and the formation by distillation of drops of spirit in the support part of the bore.

A broken liquid column can usually be reunited by holding the thermometer bulb-end downward and tapping the thermometer lightly and rapidly against the fingers or something else which is elastic and not too hard. The tapping should be continued for some time (5 min if necessary), and afterwards the thermometer should be hung, or stood, upright in a suitable container, bulb downward, for at least 1 h to allow any spirit adhering to the glass to drain down to the main column. If such treatment is not successful, a more drastic method is to cool the bulb in a freezing mixture of ice and salt, while keeping the upper part of the stem warm; the liquid will slowly distil back to the main column. Alternatively, the thermometer may be held upright with its bulb in a vessel of warm water, while the stem is tapped or shaken from the water as soon as the top of the spirit column reaches the safety chamber at the top of the stem. Great care must be taken when using this method as there is a risk of bursting the thermometer if the spirit expands into the safety chamber.

2.4.7.2	Scale illegibility

Another shortcoming of unsheathed liquid-in-glass thermometers is that with time their scale can become illegible. This can be corrected at the station by rubbing the scale with a dark crayon or black lead pencil.

2.4.8	Safety

Mercury, which has been the liquid most commonly used in liquid-in-glass thermometers, is poisonous if swallowed or if its vapour is inhaled. If a thermometer is broken and the droplets of mercury are not removed there is some danger to health, especially in confined spaces. More information on safety precautions for the  use of mercury is given in Part I, Chapter 3, in section 3.5.8. There are also restrictions on the carriage of mercury-in-glass thermometers on aircraft, or special precautions that must be taken to prevent the escape of mercury in the event of a breakage. The advice of the appropriate authority or carrier should be sought.

2.5	Radiation shields

A radiation shield or screen should be designed to provide an enclosure with an internal temperature that is both uniform and the same as that of the outside air. It should completely surround the thermometers and exclude radiant heat, precipitation and other phenomena that might influence the measurement.

Screens with forced ventilation, in which air is drawn over the thermometer element by a fan, may help to avoid biases when the microclimate inside the screen deviates from the surrounding air mass. Such a deviation only occurs when the natural wind speed is very low (< 1 m s–1). When such artificial ventilation is used, care should be taken to prevent the deposition of aerosols and rain droplets on the sensor which decrease its temperature towards the wet-bulb temperature. Artificially ventilated radiation shields should provide a clear indication of the fan status directly on the screen or on the control unit or data logger to allow maintenance staff to check whether the fan is functioning properly by visual inspection. Additionally, the fan status and preferably the fan speed should be provided in the data output for automatic monitoring purposes.

As a shield material, highly polished, non-oxidized metal is favourable because of its high reflectivity and low heat absorption. Nevertheless, thermally insulating plastic-based material is preferable because of its simple maintenance requirements. Thermally insulating material must be used if the system relies on natural ventilation.

The performance of a screen (response behaviour and microclimate effects introducing unwanted biases) depends predominantly on its design, in which care must be taken to ensure both radiation protection and sufficient ventilation. Since the start of meteorological temperature measurements, very diverse types of screens have been designed. Following the introduction of temperature measurements taken in automatic weather stations, the variety of these designs has increased significantly (see WMO, 1998a). Because of differences in specific applications, the degree of automation and climatology, it is difficult to recommend one specific type of design suitable for worldwide measurements. Nevertheless, many investigations and intercomparisons on designs and their performance have been carried out. A clear overview of screen designs is given by WMO (1972). Results of thermometer screen intercomparisons are reported by Andersson and Mattison (1991), Sparks (2001), WMO (1998b, 1998c, 1998d, 2000a, 2000b, 2002b, 2002c, 2002d, 2011) and Zanghi (1987).

An international standard (ISO/DIS 17714) defines most relevant screen types and describes the methods to determine or compare screen performances (ISO, 2007).

2.5.1	Louvred screens

Most of the numerous varieties of louvred screen rely on natural ventilation. The walls of such a screen should preferably be double-louvred and the floor should be made of staggered boards, but other types of construction may be found to meet the above requirements. The roof should be double-layered, with provisions for ventilation of the space between the two layers. In cold climates, owing to the high reflectivity of snow (up to 88%), the screen should also have a double floor. At the same time, however, the floor should easily drop or tilt so that any snow entering the screen during a storm can be removed.

The size and construction of the screen should be such that it keeps the heat capacity as low as practicable and allows ample space between the instruments and the walls. The latter feature excludes all possibility of direct contact between the thermometer sensing elements and the walls, and is particularly important in the tropics where insolation may heat the sides to the extent that an appreciable temperature gradient is caused in the screen. Direct contact between the sensing elements and the thermometer mounting should also be avoided. The screen should be painted both inside and outside with white, non-hygroscopic paint.

When double walls are provided, the layer of air between them serves to reduce the amount of heat that would otherwise be conducted from the outer wall to the inner enclosure, especially in strong sunshine. When the wind is appreciable, the air between the walls is changed continually so that the conduction of heat inwards from the outer walls is further decreased.

The free circulation of air throughout the screen helps the temperature of the inner wall adapt to ambient air changes. In this way, the influence of the inner wall upon the temperature of the thermometer is reduced. Also, the free circulation of air within the screen enables the thermometer to follow the ambient air changes more quickly than if radiative exchanges alone were operative. However, the air circulating through the screen spends a finite time in contact with the outer walls and may have its temperature altered thereby. This effect becomes appreciable when the wind is light and the temperature of the outer wall is markedly different from the air temperature. Thus, the temperature of the air in a screen can be expected to be higher than the true air temperature on a day of strong sunshine and calm wind, and slightly lower on a clear, calm night, with errors perhaps reaching 2.5 and –0.5 K, respectively, in extreme cases. Additional errors may be introduced by cooling due to evaporation from a wet screen after rain. All these errors also have a direct influence on the readings of other instruments inside the screen, such as hygrometers, evaporimeters, and the like.

Errors due to variations in natural ventilation can be reduced if the screen is fitted with a suitably designed forced ventilation system that maintains a constant and known ventilation rate, at least at low wind speeds. Care should be taken in the design of such systems to ensure that heat from the fan or an electrical motor does not affect the screen temperature.

In general, only one door is needed, with the screen being placed so that the sun does not shine on the thermometers when the door is open at the times of observation. In the tropics, two doors are necessary for use during different periods of the year. Likewise, in polar regions (where the sun is at a low angle) precautions should be taken to protect the inside of the screen from the direct rays of the sun either by a form of shading or by using a screen which is mounted so that it can be turned to an appropriate angle while the door is open for readings.

Although most screens are still made of wood, some recent designs using plastic materials offer greater protection against radiation effects because of an improved louvre design that provides a better airflow. In any case, the screen and stand should be constructed of sturdy materials and should be firmly installed so that errors in maximum and minimum thermometer readings caused by wind vibration are kept to a minimum. In some areas where wind vibration cannot be entirely damped, elastic mounting brackets are recommended. The ground cover beneath the screen should be grass or, in places where grass does not grow, the natural surface of the area.

The screen should be kept clean and repainted regularly; in many places, repainting the screen once every two years is sufficient, but in areas subject to atmospheric pollution it may be necessary to repaint it at least once a year.

2.5.2	Other artificially ventilated shields

The main alternative to exposure in a louvred screen, which is either naturally or artificially ventilated, is to shield the thermometer bulb from direct radiation by placing it on the axis of two concentric cylindrical shields and drawing a current of air (with a speed between 2.5 and 10 m s–1) between the shields and past the thermometer bulb. This type of exposure is normal in aspirated psychrometers (see Part I, Chapter 4). In principle, the shields should be made of a thermally insulating material, although in the Assmann psychrometer the shields are made of highly polished metal to reduce the absorption of solar radiation. The inner shield is kept in contact with a moving stream of air on both sides so that its temperature, and consequently that of the thermometer, can approximate very closely to that of the air. Such shields are usually mounted with their axes in a vertical position. The amount of direct radiation from the ground entering through the base of such shields is small and can be reduced by extending the base of the shields appreciably below the thermometer bulb. When the artificial ventilation is provided by an electrically driven fan, care should be taken to prevent any heat from the motor and fan from reaching the thermometers.

The design of the WMO reference psychrometer takes careful account of the effects of radiation and the use of artificial ventilation and shielding to ensure that the thermometer element is at equilibrium at the true air temperature (see Part I, Chapter 4).

_____________

Annex. Defining the fixed points of the international temperature scale of 1990

The fixed points of the International Temperature Scale of 1990 (ITS-90) of interest to meteorological measurements are contained in Table 2.1, while secondary reference points of interest to meteorological measurements are contained in Table 2.2.

The standard method of interpolating between the fixed points uses formulae to establish the relation between indications of the standard instruments and the values of the ITS-90 (BIPM, 1990). The standard instrument used from –259.34 °C to 961.78 °C is a platinum resistance thermometer.

An alternative practical method for ITS-90 approximation in platinum resistance thermometer calibration (determination of R0, A , B and C, see equation below) is to obtain resistance-temperature data by making a comparison with a calibrated standard platinum resistance thermometer at numerous temperatures in the range of interest and then fit a polynomial to the data by a least-squares technique.

The relationship between the resistance of the platinum resistance thermometer under calibration and the temperature measured with a reference thermometer is described with an interpolation equation. The Callendar–Van Dusen equation is generally accepted as the interpolation equation for industrial platinum resistance thermometers (defined in the IEC 60751: 2008) rather than for standard platinum resistance thermometers:

Rt = R0 (1 + A· t + B ·t2 + C· (t – 100) · t3)	(A.1)

where Rt is the resistance at temperature t of a platinum wire, R0 is its resistance at 0 °C (ice point) and A, B and C (C = 0 for t > 0 °C) are constants which are found using the least-squares method on the data acquired during the calibration.

Table 2.1. Defining fixed points on the ITS-90

| Equilibrium state | Assigned value of ITS |
| --- | --- |
|  | K | ˚C |
| Equilibrium between the solid, liquid and vapour phases of argon (triple point of argon) | 83.805 8 | –189.344 2 |
| Equilibrium between the solid, liquid and vapour phases of mercury (triple point of mercury) | 234.315 6 | –38.834 4 |
| Equilibrium between the solid, liquid and vapour phases of water (triple point of water) | 273.160 0 | 0.01 |
| Equilibrium between the solid and liquid phases of gallium (melting point of gallium) | 302.914 6 | 29.764 6 |
| Equilibrium between the solid and liquid phases of indium (freezing point of indium) | 429.748 5 | 156.598 5 |

Table 2.2. Secondary reference points and their temperatures on the ITS-90

| Equilibrium state | Assigned value of ITS |
| --- | --- |
|  | K | ˚C |
| Equilibrium between the solid and vapour phases of carbon dioxide (sublimation point of carbon dioxide) at standard atmospheric pressure p0 (1 013.25 hPa). The temperature t as a function of the vapour pressure of carbon dioxide is given by the equation: t = [1.210 36 · 10–2 (p – p0) – 8.912 26 · 10–6(p – p0)2 – 78.464] °C where p is the atmospheric pressure in hPa, in the temperature range 194 to 195 K | 194.686 | –78.464 |
| Equilibrium between the solid and liquid phases of mercury (freezing point of mercury) at standard atmospheric pressure | 234.296 | –38.854 |
| Equilibrium between ice and air-saturated water (ice-point) at standard atmospheric pressure. Variation of -100 hPa results in a shift of approximately + 0.001 K | 273.150 | 0.00 |
| Equilibrium between the solid, liquid and vapour phases of phenoxybenzene (diphenyl ether) (triple point of phenoxybenzene) | 300.014 | 26.864 |

_____________

References and further reading

Andersson, T. and I. Mattison, 1991: A Field Test of Thermometer Screens. SMHI Report No. RMK 62, Norrköping.

Bureau International des Poids et Mesures, 1989: Procès-Verbaux du Comité International des Poids et Mesures, 78th meeting, 1989, Paris (available from http://www.bipm.org/utils/common/pdf/its-90/ITS-90.pdf).

Bureau International des Poids et Mesures/Comité Consultatif de Thermométrie, 1990: The International Temperature Scale of 1990 (ITS-90) (H. Preston-Thomas). Metrologia, 27:3–10 (amended version) (available from http://www.bipm.org/utils/common/pdf/its-90/ITS-90_metrologia.pdf).

Her Majesty’s Stationery Office (HMSO)/Meteorological Office, 1980: Handbook of Meteorological Instruments, Volume 2: Measurement of temperature, London.

International Electrotechnical Commission, 2008: Industrial Platinum Resistance Thermometers and Platinum Temperature Sensors, IEC 60751:2008. Geneva.

International Organization for Standardization, 2007: Meteorology – Air Temperature Measurements – Test Methods for Comparing the Performance of Thermometer Shields/Screens and Defining Important Characteristics, ISO/DIS 17714:2007. Geneva.

Jones, E.B., 1992: Jones’ Instrument Technology. Volume 2: Measurement of temperature and chemical composition, Butterworths-Heinemann, Oxford.

Middleton, W.E.K. and A.F. Spilhaus, 1960: Meteorological Instruments. University of Toronto Press.

Sparks, W.R., 1970: Current concepts of temperature measurement applicable to synoptic networks. Meteorological Monographs, 11(33):247–251.

———, 2001: Field trial of Metspec screens. Technical Report TR19, Met Office/OD, Wokingham, United Kingdom of Great Britain and Northern Ireland.

World Meteorological Organization, 1972: The Effect of Thermometer Screen Design on the Observed Temperature (W.R. Sparks). (WMO-No. 315). Geneva.

———, 1992: Measurement of Temperature and Humidity: Specification, Construction, Properties and Use of the WMO Reference Psychrometer (R.G. Wylie and T. Lalas). Technical Note No. 194 (WMO-No. 759). Geneva.

———, 1998a: Recent Changes in Thermometer Screen Design and their Impact (A. Barnett, D.B. Hatton and D.W. Jones). Instruments and Observing Methods Report No. 66 (WMO/TD-No. 871). Geneva.

———, 1998b: An investigation of temperature screens and their impact on temperature measurements (J. Warne). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-98). Instruments and Observing Methods Report No. 70 (WMO/TD-No. 877). Geneva.

———, 1998c: A thermometer screen intercomparison (J.P. van der Meulen). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-98). Instruments and Observing Methods Report No. 70 (WMO/TD-No. 877). Geneva.

———, 1998d: Comparison of meteorological screens for temperature measurement (G. Lefebvre). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-98). Instruments and Observing Methods Report No. 70 (WMO/TD-No. 877). Geneva.

———, 2000a: A comparison of air temperature radiation screens by field experiments and computational fluid dynamics (CFD) simulations (A. Spetalen, C. Lofseik and P. Ø. Nordli). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-2000). Instruments and Observing Methods Report No. 74 (WMO/TD-No. 1028). Geneva.

———, 2000b: Temperature measurements: Some considerations for the intercomparison of radiation screens (J.P. van der Meulen). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-2000). Instruments and Observing Methods Report No. 74 (WMO/TD-No. 1028). Geneva.

———, 2002a: Measurement of temperature with wind sensors during severe winter conditions (M. Musa, S. Suter, R. Hyvönen, M. Leroy, J. Rast and B. Tammelin). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-2002). Instruments and Observing Methods Report No. 75 (WMO/TD-No. 1123). Geneva.

———, 2002b: Norwegian national thermometer screen intercomparison (M.H. Larre and K. Hegg). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-2002). Instruments and Observing Methods Report No. 75 (WMO/TD-No. 1123). Geneva.

———, 2002c: Results of an intercomparison of wooden and plastic thermometer screens (D.B. Hatton). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-2002). Instruments and Observing Methods Report No. 75 (WMO/TD-No. 1123). Geneva.

———, 2002d: Temperature and humidity measurements during icing conditions (M. Leroy, B. Tammelin, R. Hyvönen, J. Rast and M. Musa). Papers Presented at the WMO Technical Conference on Meteorological and Environmental Instruments and Methods of Observation (TECO-2002). Instruments and Observing Methods Report No. 75 (WMO/TD-No. 1123). Geneva.

———, 2010: Guidance on Instrumentation for Calibration Laboratories, Including RICs (D. Groselj). Instruments and Observing Methods Report No. 101 (WMO/TD-No. 1543). Geneva.

———, 2011: WMO Field Intercomparison of Thermometer Screens/Shields and Humidity Measuring Instruments (M. Lacombe, D. Bousri, M. Leroy and M. Mezred). Instruments and Observing Methods Report No. 106 (WMO/TD-No. 1579). Geneva.

Zanghi, F., 1987: Comparaison des Abris Météorologiques. Technical Memorandum No. 11, Météo-France/SETIM, Trappes.

Additional Reference

Thermodynamic analysis and experimental study of the effect of atmospheric pressure on the ice point A. H. Harvey, M. O. McLinden, and W. L. Tew

Citation: AIP Conference Proceedings 1552, 221 (2013); doi: 10.1063/1.4819543

View online: http://dx.doi.org/10.1063/1.4819543

# Footnotes

Advice concerning the safe use of mercury is given in Part I, Chapter 3, 3.5.8.

Wet-bulb temperatures are explained in Part I, Chapter 4.

# Comments

All variables should be italicized.

There should be a space between a number and its unit symbol

There should be no space between the degree symbol and the C i.e. °C not ° C

The degree symbol is ALT-2-4-8 in Windows

This is a rather confused . I suggest just deleting it and using the replacement text provided

Is that really common usage?

Italicise variables, get spaces correct. Long minus sign

MOle

They still possess zero-point energy.

Delete that table? And maybe just mention the ice point and its pressure dependence.

Calibration should be described later. It comes in two stages and this sentence is misleading.

Metrologically, we rarely speak of ‘accuracy’ Instead we specify the uncertainty of measurement. But in this case the use is fine.

Do we need to specify what kind of accreditation is appropriate? IS017025? WMO?

Weekly? Monthly? Annually?

The choice of recommended period is contentious. Ideally, an annual assessment would be used, but for many people this would disrupt measurements, cost money, and in most instances would be unnecessary – the calibration of a good quality PRT would not be expected to shift by much after standard meteorological exposures. Perhaps we could recommend a ‘check’ in an icepoint?

To remove any Fahrenheit ambiguity?

Should it be 20 s in still air? Or moving air. Or only ‘approximately  20 s’

Probably best just not mention radiometers here. They are a ‘bag of worms’!

Update the references

Add a mention about available electrical power – more relevant in an automated network.

IS there a better phrase? Perhaps ‘re-radiated heat’ or absorptions and emission of radiation.

Should we mention artificially ‘aspirated’ or ‘ventilated’ screens?

In the appendix Sparks 1970 is followed by several useful and more modern references. Perhaps this could just be changed to ‘Sparks (1970) and following references’. I have also written to Victor Venema asking for advice.

Comments from Victor Venema about Siting

Leroy, M., 1999: Classification d’un site. Note Technique no. 35.

Direction des Systèmes d’Observation, Météo-France, 12 pp.

Leroy, M., 2010: Siting Classification for Surface Observing Stations on Land, Climate, and Upper-air Observations JMA/WMO Workshop on Quality Management in Surface, Tokyo, Japan 27-30 July 2010 http://www.jma.go.jp/jma/en/Activities/qmws_2010/CountryReport/CS202_Leroy.pdf

Maybe Michel Leroy can help you further. These reports do not mention much literature to justify the choices.

For the larger-scale surrounding there are the Local Climate Zones:

Stewart, I.D. and Oke, T.R.  2012.  Local Climate Zones for urban temperature studies.  Bulletin of the American Meteorological Society,

93:  1879-1900.

Stewart, I.D., Oke, T.R., and E.S. Krayenhoff.  2014.  Evaluation of the ‘local climate zone’ scheme using temperature observations and model simulations.  International Journal of Climatology, 34:  1062-80.

http://www.wudapt.org/wudapt/

Is that correct? 100 m2 – is very much greater than 75 cm2 mentioned previously.

Do you mean high quality, or do you mean an SPRT

In operation an SPRT always measures W – the resistance ratio R(T) /R(TTPW).

It is curious that a psychrometer is the reference instrument for air temperature. IS this because it is aspirated?

Further discussion needed

It is metrological good practice to express uncertainties as standard uncertainties. Is there a reason that you are applying a coverage factor of 2?

Should we mention their disadvantages? Power requirements, and rapidly changing interface technology?

Variables should be italicised. I may miss out on this because it generates clutter on the page

Finding this inverse function is not always easy. I guess we want to avoid talking about the technicalities, but I just thought it was worth mentioning.

How stable? Why two years?

Does this mean that tungsten is used for meteorological purposes, but not nickel or platinum

And also with ceramic powder packing within a metal sheath.

Should we mention a typical figure?

Is b always negative then?

This entire section is just wrong. The seebeck effect is not a ‘contact effect’.

Is that really used? Iron is very corrodible!

The most common thermocouple combination is chromel alumel.

Thermocouples are given letter designations. Chromel Alumel is Type K, and Copper Constantan is Type T.

Lots more to say about thermocouples  e.g. extension wires.

Is it common to use bridges? In my field we use bridges for ultimate precision but I would have thought that it was much more common to use digital voltmeters or analogue to digital converters.

Absolutely >nobody< uses galvanometers. They are obsolete. Also they would detect the out-of-balance voltage (rather than the current)

SUITABLE PRECAUTIONS = 4 wire measurements.

This needs to be considered for all measurements – not just remote measurements. For a 100 Ω standard each degree corresponds to approximately 0.39 Ω so an addition 0.1 Ω maks an error of 0.25 °C.

Is this not nearly universal nowadays? Surely we should just mention alternatives to this as historical items.

Unless there is a universe that I don’t know about – and that is possible – none of these measurements are used.

Mention these? Are they used in cheap AWS devices.

Worth a picture?

This is true for PRTs and PT100s, but there is no regime in which the output of a thermistor is even approximately linear. Typically, a circuit takes the logarithm

The voltages are always small. It is especially important close to the voltmeter terminals which form a  secondary junction.

I think this means ‘pick up’ of mains hum and RF interference.

Indeed – or short circuits.

Is anybody using galvanometers to read thermocouples?

Standard cells are not commonly used except in metrology labs.

These are pretty much all irrelevant.

The most common problems with thermocouples which are not exposed to high temperatures are:

- The reference junction is not properly compensated.

That’s pretty much it.

I guess if the thermocouple is not hermetically sealed it could corrode.

‘potentiometric method’ when compared with what? Measurement uncertainty below 0.1

°C will be difficult in meteorological environments

As before: what do we mean by accredited?

What are the recommended calibration intervals? Every 10 °C? Every 5 °C Over what range?

I can’t see this text right now: what does it say?

Usually the temperature indicator includes a measurement of the reference junction temperature. The device and thermocouple would normally be calibrated a whole. I am not sure how that applies

Is it worth to re-state the convention that corrections are always ADDED to the reading? Thus if an instrument over reads the convention is to ADD a correction whose value is a negative number.

What is the meaning of ‘large’?

?? exposed how?

Mercury comment?

Would a diagram help?

A general question. Is this this the right level of detail for liquid in glass instruments? It just seems slightly overly detailed compared to the other techniques.

Needs to be discussed, which one are still necessary.

Needs to be checked

I am still looking up data on the pressure dependence. You might be interested to that the difference between the triple point of water (273.16 K exactly) and the melting temperature (273.150 K to within a mK) arises from the lowered pressure in the triple point of water cell (7 mK) and from the removal of dissolved air (3 mK).  I will find out the exact coefficients for these effects.