measurement of atmospheric composition

This chapter substitutes chapters 16 and 17 of the previous edition

16.1	General

The main purpose of this chapter is to introduce readers (particularly those who newly start the measurements) to methods and specific techniques used for measurements of various components of atmospheric composition and a number of physical parameters of the atmosphere related to its composition. Such measurements are often accompanied by measurement of basic meteorological variables, as introduced in the preceding chapters. Within WMO, the Global Atmosphere Watch (GAW) Programme was established to coordinate atmospheric composition and related physical parameter measurements taken by WMO Member countries. For further practical details of the measurement activities, the reader is referred to the GAW reports and other references listed at the end of the chapter.

The need to understand and to identify scientifically sound measures to control the increasing influence of human activity on the global atmosphere forms the rationale of the GAW Programme (GAW report No. 172). The grand challenges addressed by GAW include:

- Stratospheric ozone depletion and the increase of ultraviolet (UV) radiation at Earth’s surface;

- Changes in weather and climate related to human influence on atmospheric composition, particularly, greenhouse gases, and impact on ozone and aerosols as well as due to natural processes;

- Risk reduction of air pollution on human health and issues involving long-range transport and deposition of air pollution.

In addition, measurements of atmospheric composition are essential for understanding the radiation budget of the atmosphere and for improvement of NWP.

The GAW monitoring system focuses on six classes of variables:

- Ozone: column (total) ozone and ozone vertical profiles with a focus on the stratosphere and upper troposphere;

- Greenhouse gases: carbon dioxide CO2  (including Δ14C, δ13C and δ18O in CO2, and O2/N2 ratios), methane CH4 (including δ13C and δD in CH4), nitrous oxide N2O, halogenated compounds  SF6;

- Reactive gases: surface and tropospheric ozone, carbon monoxide (CO), Volatile Organic Compounds (VOCs), nitrogen oxides (NOx), sulphur dioxide (SO2) and molecular hydrogen (H2);

- Atmospheric Wet Deposition (focused largely on major ions);

- UV Radiation;

- Aerosols (including physical properties, size distribution and chemical composition).

A number of ancillary parameters are recommended for measurement at GAW stations:

1) 	Solar radiation;

2) 	Major meteorological parameters;

3) 	Natural radioactivity including krypton-85, radon, and some other radionuclides.

Due to the low mixing ratios of atmospheric trace constituents the instruments and methods used for the quantitative and qualitative determination of atmospheric constituents are complex and sometimes difficult to operate. Small errors, for example in spectral signatures or cross-sensitivities to other compounds can easily confound the accuracy of atmospheric composition measurements. Therefore, besides correct operation, regular calibration of the equipment, participation in inter-comparison exercises, station audits and personnel training are essential for accurate and reliable measurements. Obtaining reliable and high quality results for most of the measurements described here is not feasible without the close involvement of specialist staff at a professional level. Main principles of the Quality Assurance of atmospheric composition observations within GAW are described in the section 16.1.3

# 16.1.1	Definitions/descriptions

Depending on the measurement principle and instrument platform, three types of measurements are routinely performed and reported, namely:

(a)	Near-surface atmospheric content (from monitoring stations or mobile platforms such as ships, cars, or trains);

(b)	Total atmospheric column content (from ground or satellite based remote sensing);

(c)	Vertical concentration profiles (from aircraft, balloons, rockets, ground based remote sensing or satellite instruments);

Near-surface atmospheric content refers to the results of (continuous or discrete) measurements of a particular component’s quantity in the atmospheric layer of a few tens of metres above the ground at a particular location on Earth’s surface. Results of surface measurements are commonly given in units of partial pressure, concentration, mixing ratio or mole fraction, though use of non SI units is strongly discouraged.

Total atmospheric column refers to the total amount of particular substance contained in a vertical column in the atmosphere above the ground extending from Earth’s surface to the upper edge of the atmosphere.  Commonly used units of total ozone are (a) column thickness of a layer of pure ozone at standard temperature and pressure (STP) and (b) vertical column density (total number of molecules per unit area in an atmospheric column). For the other atmospheric constituents vertical column density or column-averaged abundances are used. It is also common to report partial column content of a substance, as for example the tropospheric column content of NOx. Here, the vertical column that is integrated extends from Earth’s surface to the tropopause.

The vertical concentration profile expresses the variation of the content of a trace compound in the atmosphere (given in the same units as surface content, namely partial pressure, concentration, number density, mixing ratio or mole fraction) as a function of height or ambient pressure.

Observations of atmospheric composition include gaseous composition, aerosol and precipitation chemistry. The characteristics of the precipitation chemical composition are given in the section 16.2.4. The variables describing aerosols (physical and chemical properties) are listed in the aerosol section (16.2.5).

The following units are used to express the results of atmospheric trace compound observations:

- Dobson unit (DU): A measure of total ozone equalling a thickness of 10–5 m of pure ozone at STP (commonly used but not a unit in the International System of Units).

- Milliatmosphere centimetre (m-atm-cm): A measure of total ozone equal to 10–3 cm of pure ozone at STP (1 m-atm-cm is equivalent to 1 DU).

- Mole fractions of substances in dry air (dry air includes ALL gaseous species except water vapor):

μmol/mol = 10-6 mole of trace substance per mole of dry air

nmol/mol = 10-9 mole of trace substance per mole of dry air

pmol/mol = 10-12 mole of trace substance per mole of dry air

Dry mole fraction requires either drying air samples prior to measurement or correcting the measurement for water vapor. When drying is impossible or the correction would add substantial uncertainty to the measurement, wet mole fractions can be reported instead. This must be clearly indicated in the metadata of the observational record.

The appropriate unit for amount of substance is dry-air mole fraction, reported as ppm (parts per million, i.e. µmol/mol), ppb (parts per billion, i.e. nmol/mol), or ppt (parts per trillion, i.e. pmol/mol). Often a “v” has been appended to these units to indicate mixing ratio by volume. When reporting mole fractions as volume mixing ratios, one assumes the atmosphere to be an ideal gas. Deviations from ideal under GAW conditions can be large (e.g., CO2), so the use of mole fraction is strongly preferred, because it does not require an implicit assumption of ideality of the gases and, more importantly, because it is applicable also to condensed-phase species. In general use of SI units is highly recommended.

- Isotope or molecular ratios:

Atmospheric molecules can be present in different isotopic configurations. Isotope ratio data are expressed as deviations from an agreed-upon reference standard using the delta notation:

δ = (Rsample/Rreference – 1), with R = [heavy isotope]/[light isotope].

δ-Values are expressed in multiples of 1000 (‰; per mil ‘units’).

The international reference scale (i.e., the primary scale) for δ13C is Vienna Pee Dee Belemnite (VPDB). NBS19 and LSVEC (since 2006) are the primary international reference materials defining the VPDB scale.  For δ18O, multiple scales are in use (VPDB, Vienna Standard Mean Ocean Water (VSMOW), air-O2).

The delta notation is also used to express relative abundance variations of O2/N2 (and Ar/N2) ratios in air:

δ(O2/N2) = (Rsample/Rstandard – 1) with R = O2/N2

The respective international air standard is not established, yet. The Scripps Institution of Oceanography (SIO) local O2/N2 scale, based on a set of cylinders filled at the Scripps Pier is the most widely used scale.

δ(O2/N2) values are expressed in multiples of 106 or per meg ‘units’.

Precipitation chemistry observations include measurements of several parameters which are described in more details in Section 16.2.4. The following units are used:

- pH measurements are expressed in units of acidity defined as:  pH = -log10 [H+], where [H+] is expressed in mole L-1  ;

- Conductivity is expressed in µS cm-1 (micro Siemens per centimeter), unit commonly used for measuring electric conductivity;

- Acidity/Alkalinity is expressed in µmole L-1  (micromole per liter);

- Major ions contents is expressed in mg L-1 (milligram per liter).

Aerosol observations of volumetric quantities, i.e., the amount of substance in a volume of air, are reported for standard temperature and pressure conditions of 273.15 K and 101.325 kPa.  These may refer to a number concentration of particles (cm-3), an area concentration (m2 m-3, or m-1), or a mass concentration (g m-3).  Aerosol optical depth is a dimensionless quantity.

# 16.1.2	Measurement principles and techniques

The existing measurement techniques for atmospheric chemical composition measurements can be separated into three main groups: passive sampling, active sampling and remote sensing techniques. Principally, active techniques draw the air sample through the detector or sampling device by a pump, whereas passive techniques use the diffusion of air to the sampling device. In remote sensing techniques, the analyzed air mass and the detector are at different locations. Total or partial column measurements are possible only with remote sensing techniques.

In the case of active sampling, measurements can either be done continuously (or at least quasi-continuously with short integration times) or samples can be collected or specially prepared  (in glass or stainless steel cylinders, on sorbent substrates or filters) and analysed off-line in specialized laboratories. Collection of discrete samples entails the storage of samples. During this time flask properties may influence the composition of the sample due to chemical or surface effects or permeation through sealing polymers during storage. This demands careful tests of the sampling containers.

The analytical techniques most commonly used (and recommended in the GAW program) for detecting and quantifying atmospheric trace constituents can be summarized as follows:

1) Spectroscopic methods refer to the measurement of changes in radiation intensity due to absorption, emission, photoconductivity or Raman scattering of a molecule or aerosol particle as a function of wavelength. Spectral measurement devices are refereed to as spectrometers, spectrophotometers, spectrographs or spectral analyzers. Spectral measurements can be performed in different parts of spectrum depending on the component to be measured or on several individual wavelengths. As absorption lines are different for molecules with different isotopic composition and line shapes depend on the bulk composition of the gas it should be taken care that reference gases have similar properties as the analyzed atmospheric air.

2) Gas chromatography (GC) is a physical method of separation that distributes components to separate between two phases, one stationary (stationary phase), the other (the mobile phase) moving in a definite direction. There are numerous chromatographic techniques and corresponding instruments. To be suitable for GC analysis, a compound must have sufficient volatility and thermal stability. GC involves a sample being vapourised and injected onto the head of the chromatographic column. The sample is transported through the column by the flow of inert, gaseous mobile phase. The column itself contains a liquid stationary phase which is adsorbed onto the surface of an inert solid. A chromatography detector is a device used to visualize components of the mixture being eluted off the chromatography column. There are two general types of detectors: destructive and non-destructive. The destructive detectors perform continuous transformation of the column effluent (burning, evaporation or mixing with reagents) with subsequent measurement of some physical property of the resulting material (plasma, aerosol or reaction mixture), e.g. Flame Ionization Detector (FID). The non-destructive detectors are directly measuring some property of the column effluent (for example UV absorption) and thus affords for the further analyte recovery, e.g. Electron Capture Detector (ECD).

3) Mass spectrometry (MS) is an analytical technique that produces spectra of the masses of the molecules comprising a sample of material. The spectra are used to determine the elemental composition of a sample, the masses of particles and of molecules, and to elucidate the chemical structures of molecules. Mass spectrometry works by ionizing chemical compounds to generate charged molecules or molecule fragments and measuring their mass-to-charge ratios. In a number of instruments mass spectrometry can be used as a detector method for gas chromatography.

Detection methods of gases and aerosols can be different and based on different physical phenomena. The details of the detection methods applicable to different gases and aerosol properties are summarized in the sections below.

# 16.1.3	Quality assurance

The objectives of the GAW Quality Assurance (QA) system are to ensure that data reported by stations are consistent, of known and adequate quality, supported by comprehensive metadata, and regionally or globally representative with respect to spatial and temporal distribution.

The principles of the GAW QA system apply to each measured variable and include:

- Defined data quality objectives (including tolerable levels of uncertainty in the data, completeness, compatibility requirements etc.).

- Establishment of harmonized recommendations on measurement techniques and quality control procedures to reach data quality objectives (Measurement Guidelines and Standard Operating Procedures).

- Network-wide use of only one reference standard or scale (primary standard). In consequence, there is only one institution that is responsible for this standard.

- Traceability to the primary standard of all measurements made by GAW stations.

- Use of detailed log books for each parameter containing comprehensive meta information related to the measurements, maintenance, and quality control actions.

- Regular independent assessments (including audits and comparison campaigns).

- Timely submission of data and associated metadata to the responsible World Data Centre as a means of permitting independent review of data by a wider community.

The following GCOS monitoring principles apply also to the GAW observations:

- The impact of new systems or changes to existing systems should be assessed prior to implementation.

- A suitable period of overlap for new and old observing systems should be required.

- Uninterrupted station operations and observing systems should be maintained.

The GAW QA system further recommends the adoption and use of internationally accepted methods and vocabulary to describe uncertainty in measurements.

Five types of central facilities (Annex 16.1) dedicated to the six groups of measurement variables (see 16.1) are operated by WMO Members and form the basis of quality assurance and data archiving system. Those include:

- Central Calibration Laboratories (CCLs) that host primary standards and scales;

- World/Regional Calibration Centres (WCCs/RCCs) that coordinate inter-comparison campaigns, help with the instruments calibration and perform stations/labs audits;

- Quality Assurance/Science Activity Centres (QA/SACs) provide technical and scientific support and coordinate cooperation between the central facilities and the GAW stations.

- World Data Centres (WDCs) mainly ensure dissemination and easy access of GAW data and secure the data through appropriate data archiving.

The work of the Central Facilities on the quality assurance of the GAW observations is supported by respective Scientific Advisory Groups those tasks include in particular assistance in development of the measurement procedures and guidelines, data quality objectives and, when applicable, standard operating procedures, review new measurement techniques and make recommendations about their applicability for the GAW observations.

16.2	Measurement of specific variables

The following sections will briefly introduce the measurement of the main compounds observed under the GAW program. This Section provides the summary of the measurement techniques, while comprehensive measurement guidelines can be found in the specialized GAW Reports, cited in individual sections. In the cases, where GAW Measurement Guidelines or Standard Operating Procedures are not available, the links are provided to the information, necessary to carry out respective measurements. The background for the measurements of individual components can be found in the WMO Global Atmosphere Watch (GAW) Strategic Plan: 2008-2015 (GAW Report No. 172) and its Addendum (GAW Report No. 197).

## 16.2.1 (Stratospheric) Ozone

### 16.2.1.1 Ozone Total Column

Measurements of total ozone are possible by remote sensing techniques only. Most precise information on total ozone and its changes at individual sites can be obtained by measurements from the ground, e.g. by solar spectroscopy in the wavelength region of 300-340 nm. Within the GAW Programme, Dobson spectrophotometers (designed for manual operation) and Brewer spectrophotometers (designed for automatic operation) are used as the instruments for routine total ozone observations, thus providing two independent networks.

Details of the total ozone measurements with the Dobson spectrometer and their quality assurance are provided in the GAW Report No. 183. Total ozone observations are made with this instrument by measuring the relative intensities of selected pairs of ultraviolet wavelengths, called the A, B*, C, C', and D wavelength pairs, emanating from the sun, moon or zenith sky.  The A wavelength pair, for example, consists of the 3055 A.U. (Angstrom Units, 1 A.U. = 10 nm) wavelength that is highly absorbed by ozone, and the more intense 3250 A.U. wavelength that is relatively unaffected by ozone.  Outside the earth's atmosphere the relative intensity of these two wavelengths remains essentially fixed.  In passing through the atmosphere to the instrument, however, both wavelengths lose intensity because of scattering of the light by air molecules and dust particles; additionally, the 3055 A.U. wavelength is strongly attenuated while passing through the ozone layer whereas the attenuation of the 3250 A.U. wavelength is relatively weak. The relative intensity of the A wavelength pair as seen by the instrument, therefore, varies with the amount of ozone present in the atmosphere since as the ozone amount increases the observed intensity of the 3055 A.U. wavelength decreases, whereas the intensity of the 3250 A.U. wavelength remains practically unaltered.  Thus, by measuring the relative intensities of suitably selected pair wavelengths with the Dobson instrument, it is possible to determine how much ozone is present in a vertical column of air extending from ground level to the top of the atmosphere in the neighbourhood of the instrument.  The result is expressed in terms of a thickness of a layer of pure ozone at standard temperature and pressure.

The measurement principle of the Brewer spectrometer is similar to the one of the Dobson instrument. The operating procedures are provided by the producing company at http://www.kippzonen.com/?productgroup/26142/Brewer+Spectrophotometer.aspx. The recommendations for the GAW network are available at the World Ozone and Ultraviolet Radiation Data Centre (WOUDC) at   ftp://ftp.tor.ec.gc.ca/pub/woudc/documentation/SOP_Documents/brewerspectrophotometer_sop-june2008.pdf

Results of comparisons of Brewer and Dobson instruments, as well as recommendations on the operation of the Brewer instruments are provided in the reports of the Biennial WMO consultation on Brewer ozone and UV spectrometer operation, calibration and data processing (e.g. GAW Report No. 176).

The world (primary) standard instruments of Brewer and Dobson networks are calibrated by the Langley plot method performed at the Mauna Loa Observatory at Hawaii (every 2-4 years); regional standards are calibrated against the primary standard every 2-3 years; and the station instruments are calibrated by side-by-side calibration with the standard instruments every 4 years. An extension of these calibration cycles up to 5-6 years for station instruments is currently in the pipeline. In addition three successful Langley plot campaigns at the Izaña Observatory on Tenerife with primary and regional standard Dobsons have proved the suitability of this location and facility for this absolute calibration method.

Complementary measurements of total ozone are provided by the Differential Optical Absorption Spectroscopy (DOAS) type UV/visible spectrometers that also allow detection of various minor trace gases (e.g. NO2 and BrO). The French instrument is called Systeme d'Analyse par Observation Zenithale (SAOZ) but it is based on the same principle as DOAS. These instruments are part of the measurement suites within the Network for the Detection of Atmospheric Composition Change (NDACC, http://www.ndsc.ncep.noaa.gov/instr/). Compared to the more established Brewer/Dobson network, the measurement and analysis procedures for DOAS type instruments are less standardized, but regular comparison campaigns have been carried out. Other instruments providing total ozone measurements from the ground (Russian filter instruments or those of the DOAS/SAOZ type) are not operated under the same data QA/QC programme as Dobson and Brewer instruments. The Russian and other filter instruments are not independently calibrated, but are tied to either Dobson or Brewer instruments. Data quality of all individual total ozone series deposited at WOUDC needs to be documented for the users.

# 16.2.1.2. Ozone Profile Measurements

Measurements of the vertical ozone distribution are possible by both active and remote sensing methods.

Dobson spectrometer can be used for the measurements of the vertical ozone distribution utilizing Umker method (GAW Report No. 183). The reduction of the Umkehr measurement to an ozone profile requires a complex algorithm that includes knowledge of the radiative properties of the real atmosphere.  As this knowledge changes, the algorithm will change. A standard Umkehr observation consists of a series of C-pair wavelength measurements made on the clear zenith sky during morning or afternoon.  The measurements are commenced a few minutes before sunrise and continued until the sun is at an elevation of not less than about 20 degrees, or commenced in the afternoon when the sun is at an elevation of not less than about 20 degrees and continued until shortly after sunset. The zenith sky must be free from clouds for a period of from one-half to one hour near sunrise or sunset to make Umkehr observations.  This is especially true at low latitude stations where the sun rises or sets rapidly.  At other times, it is desirable that the zenith sky be cloudless but permissible that clouds cross it periodically when measurements are not made.  Umkehr observations cannot be made at a polar station or at high latitude stations during summertime when the sun does not sink below the horizon.

To be able to compute the vertical distribution of ozone, it is necessary to know the total amount of ozone present at the time of observations.  Several total ozone measurements must, therefore, be made during the morning or afternoon, particularly if the ozone amount is changing fairly rapidly.

The resulting ozone profile derived from reduction of these measurements is quite dependent on the algorithm used.  The method of Umkehr data analysis was originally developed by Götz, Meetham, and Dobson in 1934.  Later the method has been refined by Ramanathan and Dave (in 1957) and Mateer and Dütsch (in 1964).  The Umkehr algorithm is described by Petropavlovskikh et. al (2005) and updated information can be found here: http://www.esrl.noaa.gov/gmd/ozwv/umkehr/

Ozone measurement from light balloons (ozonosondes) is an active method for the measurements of ozone vertical distribution in the atmosphere. Other active methods for ozone mole fraction measurements (that are used on aircraft platforms) are described in the section on reactive gases (Section 16.2.3).

Ozonesondes are small, lightweight and compact balloon borne instruments, developed for measuring the vertical distribution of atmospheric ozone up to an altitude of about 30-35 km. The sensing device is interfaced to a standard meteorological radiosonde for data transmission to the ground station and can be flown on a small rubber balloon. Three major different types of ozonesondes, i.e. Brewer-Mast, electrochemical concentration cell (ECC), and the carbon iodine cell, are nowadays in use. Each sonde type has its own specific design.

Total weight of the flight package is typically about 1 kg and can be flown on small weather balloons. Normally data are taken during ascent, at a rise rate of about 5 m/s, to a balloon burst altitude of 30-35 km. The inherent response time of the ozonesonde is 20-30s such that the effective height resolution of the measured vertical ozone profile is typically 100-150 m.

The principles of the ozonesondes operations and an overview of the different aspects of quality assurance and quality control for ozonesonde measurements in GAW are given in details in the GAW Report No. 201.

## Ozone profile measurements can also be obtained by other instruments operated under the umbrella of NDACC Lidar and microwave measurements are part of the NDACC suite of measurements and are valuable for assessing ozone trends in the upper stratosphere and for validating satellite measurements in the upper atmosphere. The disadvantage of microwave ozone measurements is the rather poor vertical resolution, but they have a potential to measure up to the mesopause region. The combination of sonde, Umkehr, lidar, and microwave data from the ground is important for assessing the quality of the ozone profile measurements from space.

16.2.1.3 Aircraft and Satellite Observations

Ozone in the atmosphere is also measured by instruments located on board aircraft and space satellites. The airborne observations are usually made by in situ photometers sampling the air in the troposphere and lower stratosphere during a flight. The measurements are mostly used in research campaigns on atmospheric chemistry, but there’ve been also long-term efforts using commercial aircraft, e.g. the MOZAIC (Measurement of ozone, water vapour, carbon monoxide and nitrogen oxides aboard Airbus in-service aircraft), CARIBIC (Civil Aircraft for the Regular Investigation of the atmosphere Based on an Instrument Container, www.caribic-atmospheric.com) and recently IAGOS (in-service Aircraft for a Global Observing System) projects.

Large-scale monitoring of atmospheric ozone is performed by remote-sensing instruments from satellites. These programmes can be divided according to lifetime into the long-term operational monitoring systems that generate large (global) data sets used both for trend analyses and for operational mapping of ozone, and into temporary experimental missions.

Satellite observations  can be grouped according to the radiation-detection technology used for the instruments and the retrieval schemes  applied for the derivation of ozone column density or concentration from the measured radiances While nadir viewing instruments are primarily used for column observations and coarse vertical profiling, limb sounding instruments are able to measure vertical profiles of ozone at high vertical resolution by solar, lunar or stellar occultation or by observing limb scatter and emission through the atmospheric limb (Tegtmeier et al., 2013; Sofieva et al., 2013).

In 1978 continuous ozone observations started with the Total Ozone Mapping Spectrometer and (TOMS) and Solar Backscatter UV (SBUV) aboard Nimbus 7 (Heath, Krueger and Park, 1978). The past decades saw a number of missions with nadir sounding instruments, thus allowing multi-decadal observations of the global ozone column (e.g., van der A et al., 2010). Limb sounder observations since 1978 starting with SAGE, SAGE II, and HALOE provided long-term time series of stratospheric ozone profiles at a high vertical resolution (2-4 km). However, with the end of the UARS mission in 2005 and ENVISAT mission in 2012 and many aging systems flying (AURA, ODIN), a substantial decline of ozone profile limb sounding started. In general, the technologies of these space-borne systems have been improved with respect to retrieval algorithms for processing the observations and by increasing number of atmospheric species being monitored.

Satellite monitoring of ozone is closely linked to ground based observations mainly for validation purposes. Due to their global coverage, satellite ozone observations are routinely used for large scale assimilation into numerical ozone mapping and prediction models (see, e.g., the European GMEAS MACC project). Setting up and integrating ground and satellite ozone monitoring systems is a strategic task of WMO, space agencies and other related scientific groups represented by the Committee on Earth Observation Satellites (CEOS). Information about on-going and planned satellite ozone monitoring missions and instruments can be found at: http://database.eohandbook.com/ and here: http://accsatellites.aeronomie.be/index.php/satellite-missions

## 16.2.2 Greenhouse Gases

All greenhouse gases are reported in dry mole fractions on the most recent scales summarized in GAW Report No.206 (as of September 2013) and reviewed every two years at the WMO/IAEA Meetings on Carbon Dioxide, Other Greenhouse Gases and Related Measurement Techniques (GGMT). The primary reference in the case of greenhouse gases is a set of the cylinders of natural air with known mole fractions of the studied gases. The primary scale is transferred to station working standards through secondary and tertiary gas standards in cylinders.

### 16.2.2.1. Carbon Dioxide (CO2) (including Δ14C, (13C and (18O in CO2, and O2/N2 Ratios)

Carbon dioxide is usually measured by active methods in the atmospheric boundary layer.

Most of historical background atmospheric CO2 measurements are made with Non-Dispersive Infra-Red (NDIR) gas analysers, but a few programmes use a gas chromatographic (GC) method.  The GC method requires separation of CO2 from other gases in the air sample, reduction of this CO2 over a catalyst with H2 to CH4, and detection of the CO2-derived CH4 by flame ionisation detector (FID).  Chromatographic peak responses from samples are compared to those from standards with known CO2 mole fractions to calculate the CO2 mole fraction in the sample.  GC techniques are limited to a measurement frequency of 1 sample every few minutes.  NDIR instruments are based on the same principle that makes CO2 a greenhouse gas: its ability to absorb IR radiation.  NDIRs measure the intensity of IR radiation passing trough a “sample” cell relative to radiation passing through a reference cell.  It is not necessary to know the CO2 mole fraction of the reference cell gas.  Sample air, pumped from inlets mounted well-away from the measurement building, and standard gas flow alternately through the sample cell.  A difference in CO2 concentration between sample and reference gases (or standard and reference gases) contained in the two cells results in a voltage that is recorded by the data acquisition system.

Most of the new installments are performing the measurements of CO2 with laser-based optical spectroscopic methods,  like Fourier Transform Infrared Absorption Spectroscopy or High Finesse Cavity Absorption Spectroscopy, which includes Cavity Ring-Down Spectroscopy (CRDS) and off‐axis Integrated Cavity Output Spectroscopy (ICOS). Advantageous properties for these techniques are reduced calibration demands due to better linearity and detector response stability.

Carbon dioxide abundances are reported in dry-air mole fraction, (mol mol-1, abbreviated ppm, on the WMO CO2 Mole Fraction scale (WMO X2007 scale in September, 2013). Water vapour affects the measurement of CO2 in two ways: 1. H2O also absorbs IR radiation and can interfere with the measurement of CO2; 2. H2O occupies volume in the sample cell, while standards are dry.  At warm, humid sites, 3% of the total volume of air can be H2O vapour.  The impact of water vapour on the CO2 measurement must be considered. Drying to a dewpoint of -50ºC is sufficient to eliminate interferences. The novel optical spectroscopic methods often allow simultaneous determination of the H2O vapour content, making it in principle possible to correct for dilution due to H2O and spectroscopic effects. However, current best-practice (see GAW Report No. 206) still recommends sample drying while the determination of dry-air mole fractions without sampling drying and subsequent correction is under review.

An alternate method of CO2 measurement that is generally applicable to many other trace gases is collection of discrete air samples in vacuum-tight flasks.  These flasks are returned to a central laboratory where CO2 is determined by NDIR, GC or other instrument.  This method is used where low-frequency sampling (e.g., once a week) is adequate to define CO2 spatial and temporal gradients, and for comparison with in situ measurements as a quality control step. This sampling strategy has the advantage that many species can be determined from the same sample.

Measurements of O2/N2 ratios and stable isotopes of CO2 ((13C and (18O) help to partition carbon sources and sinks between the ocean and biosphere.  Isotopic measurements are often made from the same discrete samples used for CO2 mole fraction measurements.  Isotopic standards are maintained by the International Atomic Energy Agency (IAEA), but measurement sites are part of the GAW CO2 network.

Measurement methods for stable isotope determinations is Isotope Ratio Mass Spectrometry (IRMS), a specialization of mass spectrometry, in which mass spectrometric methods are used to measure the relative abundance of isotopes in a given sample. Measurement setup is described by the GAW Central Calibration Laboratory for stable isotopes at Max-Planck-Institute for Biogeochemisty in Jena (http://www.bgc-jena.mpg.de/service/iso_gas_lab/pmwiki/pmwiki.php/IsoLab/Co2InAir).  In recent years, optical analysers that report mole fractions of individual isotopologues have become increasingly available and are now in routine use.  Many of these instruments can provide isotopic ratios with a repeatability of around 0.05‰ for δ13C of atmospheric CO2 and are valuable for continuous measurements. Unlike mass spectrometric techniques, δ values from such instruments are often calculated from the ratio of individual measured mole fractions using tabulated absorption line strengths and are not from direct measurements of a standard material.  The reference isotopic abundance is normally taken from a spectral parameter database (typically HITRAN) that is used in the analysis, and this does not provide a common scale such as VPDB or the Jena Reference Air Set (JRAS). Some corrections applicable to mass spectrometric methods, such as those for 17O and N2O are not required, but other corrections, depending on the method used to calculate the isotopic “δ” values from individual mole fractions, such as interference from other atmospheric components and instrument fluctuations, may be required. It is important to realize compatibility between the techniques before measurement results are made public.

Measurements of the changes in atmospheric O2/N2 ratio are useful for constraining sources and sinks of CO2 and testing land and ocean biogeochemical models. The relative variations in O2/N2 ratio are very small but can now be observed by at least six analytical techniques.  These techniques can be grouped into two categories: (1) those which measure O2/N2 ratios directly (mass spectrometry and gas chromatography), and those which effectively measure the O2 mole fraction in dry air (interferometric, paramagnetic, fuel cell, vacuum ultraviolet photometric). A convention has emerged to convert the raw measurement signals, regardless of technique, into equivalent changes in mole ratio of O2 to N2. For mole-fraction type measurements, this requires accounting for dilution due to variations in CO2 and possibly other gases. If synthetic air is used as a reference material, corrections may also be needed for differences in Ar/N2 ratio. There are currently about 10 laboratories measuring O2/N2 ratios.  The O2/N2 reference is typically tied to natural air delivered from high-pressure gas cylinders.  As there is no common source of reference material, each laboratory has employed its own reference. There is currently no CCL for O2/N2. Hence it has not been straightforward to report measurements on a common scale, but several laboratories report results on a local implementation of the Scripps scale. There are no named versions, yet.

The practice of basing O2/N2 measurements on natural air stored in high-pressure cylinders appears acceptable for measuring changes in background air, provided the cylinders are handled according to certain best practices, including orienting cylinders horizontally to minimize thermal and gravitational fractionation. Nevertheless, improved understanding of the source of variability of measured O2/N2 ratios delivered from high-pressure cylinders is an important need of the community.  An independent need is the development of absolute standards for O2/N2 calibration scales to the level of 5 per meg or better.

Atmospheric 14CO2 measurements are usually reported in Δ14C notation, the permil deviation from the absolute radiocarbon reference standard, corrected for isotopic fractionation and for radioactive decay since the time of collection.  For atmospheric measurements of Δ14C in CO2, two main sampling techniques are used: High-volume CO2 absorption in basic solution or by molecular sieve, and whole-air flask sampling (typically 1.5-5 L flasks). Two methods of analysis are used: conventional radioactive counting and accelerator mass spectrometry (AMS). The current level of measurement uncertainty for Δ 14C in CO2 is 2-5‰, with a few laboratories at slightly better than 2‰ uncertainty. Recommendations on calibration are provided in the GAW Report No. 206.

Recommendations on Quality Assurance of CO2 measurements (including Δ14C, (13C and (18O in CO2, and O2/N2 Ratios) are reviewed every two years at the WMO/IAEA Meeting on Carbon Dioxide, Other Greenhouse Gases and Relater Measurement Techniques (GGMT). The GAW Report No. 206 can be used as the most recent reference regarding calibration and measurements quality control.

### 16.2.2.2 Methane (CH4)

Methane is usually measured quasi-continuously or from discrete samples by active methods in the atmosphere. Recommendations for CH4 measurements are provided in GAW Report No. 185.

For CH4 measurements at GAW stations, gas chromatography (GC) with a flame ionization detector (FID) is typically used. The analytical set-up can vary greatly depending on details such as type (manufacturer) of GC, chromatographic separation scheme used, carrier gas (e.g., N2 or He), data acquisition, system control hardware and software, and peak integration system.  Consequently, operating procedures for the individual systems will vary.

New analyzers for atmospheric measurements of CH4 based on optical methods give better repeatability than GC methods, but their long-term reliability is still being assessed.  They are also difficult to repair in the field, and they often need to be returned to the factor for repairs. Although these instruments, which also measure water vapour, often come advertised as not needing calibration or sample drying, attendees at the 13th meeting of CO2 experts (GAW Report No. 168) strongly recommend that the analyzers are calibrated routinely and that air samples are dried to a dew point of ≤-40 °C.

# 16.2.2.3 Nitrous Oxide (N2O)

Nitrous Oxide is usually measured by active methods in the atmospheric boundary layer. Recommendations for N2O measurements are provided in GAW Report No. 185.

A gas chromatograph equipped with electron capture detector (ECD) is widely used to separate and detect N2O in ambient air. This technique offers good repeatability, but it can be difficult to implement.  Because the N2O lifetime is long and fluxes small, spatial gradients are small; therefore, their quantification requires very precise measurements. The digitized ECD signal is recorded and integrated to quantify peak heights and areas. Collecting discrete samples of air in flasks is an alternative method of monitoring N2O. Flasks would be returned to a central laboratory for analysis by GC. Typical sampling frequencies are weekly or bi-weekly sampling.

Updated recommendations on the measurements calibration and quality control are provided in GAW Report No. 206.

Very recently, optical analyzers including High Finesse Cavity Absorption Spectrometers with near-infrared laser sources, Fourier Transform Infrared (FTIR) Analyzers and off-axis ICOS analyzers with mid-infrared laser sources became commercially available for N2O that exceed the precision of gas-chromatography and should allow the target DQO to be reached. First experiments typically show excellent performance; however, no recommendations can be made since the evaluation and the assessment of their long-term applicability is still in progress.

### 16.2.2.4 Halocarbons and SF6

Halocarbons and SF6 are usually measured quasi-continuously or from discrete air samples by active methods in the atmospheric boundary layer. Measurement guidelines for these species are not formalized yet in the GAW Programme.

SF6 is typically measured using gas chromatography with electron capture detection (GC-ECD) techniques on the same channel as N2O.

Global measurements of halocarbons are currently performed by NOAA and AGAGE (Advanced Global Atmospheric Gases Experiment). The measurement histories for both NOAA and AGAGE extend back to the late 1970’s. Both groups measure halocarbons using GC-ECD and GC-MS techniques. Halocarbons measured include chlorofluorcarbons (CFCs), hydrochlorofluorocarbons (HCFCs), chlorinated solvents such as CCl4 and CH3CCl3, halons, hydrofluorocarbons (HFCs), methyl halides, and SF6. For many halocarbons, measurement of mole fractions in the background troposphere requires sample pre-concentration. The AGAGE group operates a network of in situ systems, while the NOAA group operates in situ systems (for a limited number of gases) and a flask-based program. For more information on instrumentation and sampling sites see: http://agage.eas.gatech.edu; http://www.esrl.noaa.gov/gmd/hats/.

# 16.2.2.5 Remote sensing of greenhouse gases

There are several techniques that are used for remote sensing of the greenhouse gases. Total Carbon Column Observational Network (TCCON, https://tccon-wiki.caltech.edu/) is a network of ground-based Fourier Transform Spectrometers recording direct solar spectra in the near-infrared spectral region. From these spectra, column-averaged abundances of CO2, CH4, N2O, HF, CO, H2O, and HDO are retrieved. Observations in the mid-infrared (in the NDACC network, http://www.acd.ucar.edu/irwg/) allow for accurate measurements of column-averaged abundances of CH4,N2O and CO.

Satellite remote sensing of greenhouse gases is treated separately in Chapter xxxxx.

16.2.3 Reactive Gases

The reactive gases considered in the GAW Programme include surface and tropospheric ozone (O3), carbon monoxide (CO), volatile organic compounds (VOCs), oxidized nitrogen compounds (NOx), and sulphur dioxide (SO2).  All of these compounds play a major role in the chemistry of the atmosphere and as such are heavily involved in inter-relations between atmospheric chemistry and climate, either through control of ozone and the oxidising capacity of the atmosphere, or through the formation of aerosols.  The global measurement base for most of them is entirely unsatisfactory, the only exceptions being surface ozone and carbon monoxide.

Different reference standards and methods are used in the group of reactive gases. For more stable gases the reference material can be prepared as a cylinder filled with the air/other matrix with known gas mole fraction (e.g., for CO, Non-Methane Hydrocarbons and terpenes), while for others (like ozone or oxidized nitrogen compounds) only reference methods/instruments are possible.

### 16.2.3.1. Surface Ozone

The detailed Measurement Guidelines for tropospheric ozone measurements are provided in GAW Report No. 209.

The mole fraction most appropriate to the chemical and physical interpretation of ozone measurements is the mole fraction of ozone in dry air. However, ozone measurements are usually made without sample drying, because an efficient system for drying air and leaving the ozone content of the air unchanged has not been developed. It is recommended that ozone measurements be accompanied by measurements of water vapour mole fraction of sufficient precision that the ozone measurements could be converted to mole fractions with respect to dry air without loss of precision.

A number of techniques are used for measurements of ozone in the background atmosphere. These include:

- UV absorption techniques

- Chemiluminescence techniques

- Electrochemical techniques

- Cavity Ring-Down Spectroscopy (CRDS) with NO titration

- Differential Optical Absorption Spectroscopy (DOAS)

- Multi-Axis Differential Optical Absorption Spectroscopy (MAxDOAS)

- Tropospheric ozone lidar

A review of each of these techniques along with information pertinent to their applicability for use at GAW stations is provided in GAW Report No. 209. Note that only the first four of these techniques (i.e. in-situ) can be traceable via a chain of calibrations to the primary standard as recommended by GAW.

The main principle of the UV method is based on the absorption of light in the UV region by the ozone molecule. The broad UV spectrum of ozone shows its maximum around 254 nm. This wavelength represents exactly the strongest emission line of an Hg lamp and the highest spectral sensitivity of a UV detector, which is a caesium-telluride vacuum UV diode or UV-sensitive photomultiplier tube (PMT). The instrument measures the relative light attenuation between an air sample which remains unchanged (i.e. containing ozone) and one in which ozone has been removed. The ozone mole fraction is calculated via the Lambert-Beer law. Though UV absorption is an absolute measuring method calibration is necessary, at least due to determination of scrubber efficiency. Details of this measurement technique are given in GAW Report No. 209.

Because of its high accuracy and precision, low detection limit, long-term stability, sufficient time resolution and ease of operation (almost no consumables), the UV absorption technique is recommended for use for routine surface ozone measurements at all GAW stations.

The advantages of the chemiluminescence methods (CLD) for ozone measurements are their fast response times and high sensitivity relative to the UV method. This makes the chemiluminescence suitable for ambient air measurements which require high time resolution (e.g. airborne measurements). Since CLD isn’t an absolute method, calibrations are necessary. Due to its relative complexity, chemiluminescence is not recommended for routine surface ozone measurements at GAW stations. However the chemiluminescence method is a method appropriate for experimental studies of ozone at GAW stations with extended programmes or additional as back-up or for QA/QC reasons, since both instruments have different artefacts.

All electrochemical techniques use the oxidation of Iodide to Iodine by O3:

2 Iˉ   +   O3    +   H2O   ->   I2   +   2 OHˉ   +   O2

Depending on approach the formed Iodine is stabilized by former reactions, or reduced at the surface of a cathode whereby the electrical current is measured. Like UV absorption it is an absolute measuring method in principle. However, due to some sources of error, e.g. overvoltage or zero current, calibration is necessary.

The CRDS with NO titration is an experimental method with promise for future observations and should be incorporated into experimental studies of ozone measurements at selected GAW stations where appropriate.

Remote sensing techniques (DOAS, MAxDoas, Lidar) would require a similar traceability chain as in-situ measurements which is theoretically possible via the knowledge of the ozone cross-sections at the particular wavelengths used in the instruments. This issue is currently under consideration by the ACSO committee (Absorption Cross-Section for Ozone).

Differential Optical Absorption Spectroscopy (DOAS) is a ground-based remote sensing method suitable for observations of several trace substances. The instrument consists of a light source, a long ambient air open optical path generally between 100 m and several km, a retro-reflector and a spectrometer with a telescope, housed with the light source. The spectrometer observes the light source via the retro-reflector. The DOAS system uses Beer’s law to determine the ozone concentration (averaged over the light path). In principle DOAS should be a sensitive technique, but this is confounded by the inability of the system to regularly measure a definitive zero and determine the contribution of other UV absorbing gases and aerosol to the observed signal. The DOAS may be used as an experimental technique.

Multi-Axis Differential Optical Absorption Spectroscopy (MAxDOAS) is a ground-based remote sensing method for observations of several trace substances. While this method is suitable for stratospheric monitoring, it is also possible to apply it for trace gas profiles measurements in the upper and lower troposphere. However, since the retrieval procedures as well as possible tropospheric interferences are more complicated in the lower troposphere, it needs highly experienced personnel for extracting and calculating the mole fractions for the respective trace gases out of the various spectra.  MAxDOAS measurements of ozone, nitrogen dioxide, formaldehyde, bromine monoxide and other species are recommended especially for providing a link between ground-based and satellite measurements at selected GAW stations with extended research programmes.

LIDAR (LIght Detection and Ranging) is a ground-based remote sensing method for observations of several trace substances. For tropospheric ozone measurements, a LIDAR typically uses two or more wavelengths between 266 nm and 295 nm. The chosen wavelengths are shorter than the ones used for stratospheric ozone detection (typically between 308 and 353 nm). Compared to the stratosphere, higher ozone absorption efficiency is necessary in the troposphere in order to get enough sensitivity because of the lower ozone mixing ratios in the troposphere compared to the stratosphere. Too much absorption means that most light is extinguished at lower elevations, making it difficult to collect measurement signals from higher elevations. The extreme dynamic range of the back-scattering signal over the troposphere (some decades over a few kilometers of height) is a major technical problem. LIDAR tropospheric ozone measurements are recommended especially for providing a link between ground-based and satellite measurements at selected GAW stations with extended research programmes.

### 16.2.3.2. Carbon Monoxide (CO)

The detailed Measurement Guidelines for carbon monoxide measurements are provided in GAW Report No. 192.  Carbon monoxide scale is evaluated every two year together with the scales of the major greenhouse gases. For most recent scale please consult GAW Report No. 206.

Measurements of CO are possible both in-situ and by flask collection with subsequent analysis in the laboratory.  In-situ continuous observations provide information about CO variability on a time scale ranging from seconds to one hour depending on the measurement technique. In contrast to flask sampling, near-real-time data delivery can be realized using continuous measurements.

In-situ observations can be made using a broad variety of analytical techniques. Non-dispersive infrared radiometry (NDIR) is based on spectral absorption at 4.7 μm. NDIR is frequently used for continuous measurements at remote locations; however instrument drift, limited precision and long averaging times are factors limiting the achievable data quality. Gas chromatography, when coupled with a number of different detectors (e.g. flame ionization (GC/FID) or hot mercuric oxide reduction/UV absorption (GC/HgO)) can provide high precision and adequate detection limits. The HgO-reduction detector tends to have a non-linear response over the range of atmospheric CO, and requires careful, repeated multipoint characterization of the detector response.  The GC/FID technique requires catalytic conversion of CO to CH4. For confidence in the results, the catalytic conversion efficiency must be determined on a regular basis. This type of issue complicates efforts to properly maintain instrument calibration and provide accurate measurements. GC measurements are quasi-continuous in nature and therefore may not detect fast changes of mole fractions that can be captured by high frequency measurements.

Several new measurement techniques have recently become available. The most established technique is based on resonance fluorescence of CO (induced by a high frequency discharge) in the vacuum ultraviolet (VURF). This method provides low detection limits with excellent precision in the range of atmospheric mixing ratios.  A commercially available instrument based on VURF is used in several research laboratories, field sites and on other platforms, e.g. in the CARIBIC passenger aircraft project.

Spectroscopic techniques based on Cavity Ring-Down Spectroscopy (CRDS) and cavity-enhanced quantum cascade laser spectroscopy (QCL) have become available.  The CRDS technique operates using lasers in the near-infrared and was previously mainly used for measurements of carbon dioxide, methane, and ammonia. The QCL technique measures in the mid-infrared and commercial instruments are available that can determine both CO and N2O with a single laser.  Both CRDS and QCL techniques provide CO measurements with low detection limits and excellent reproducibility. Tests of their long-term applicability for routine measurements at the GAW stations are still ongoing. An additional alternative with similar performance that has recently been commercialized is Fourier Transform Infrared Absorption spectroscopy.

Remote sensing of CO column from the ground is done in the TCCON network using ground-based Fourier Transform Spectrometers in the near-infrared spectral region.

# 16.2.3.3. Volatile Organic Compounds (VOCs)

The measurement of VOCs is complex due to the many different molecules present in the atmosphere.  While systematic surveying of many of these species is important for air quality purposes, the low concentrations of VOC away from their sources imply that only a few molecules can be measured routinely in the background atmosphere.  A core set of molecules recommended for measurements in the GAW Programme with recommended measurement methods is provided in Table 1.

Currently the Measurement Guidelines for VOCs are under development in collaboration with ACTRIS (Aerosols, Clouds and Trace gas Research InfraStructure) network (http://www.actris.net/). A Standard Operation Procedure (SOP) for taking air samples with stainless steel canisters is available in GAW Report No. 204. General recommendations on VOC measurements can be found in GAW Report No. 171. Regular GAW VOC workshops review the status of VOC measurements in the GAW Programme (http://www.wmo.int/pages/prog/arep/gaw/4VOC_expert_meeting2012.html) and provide further guidance of the development of measurement techniques, quality assurance and gas standards.

Table 1. List of the priory VOCs in the GAW Programme.

| Molecule | Lifetime (OH=106 cm-3) | Importance to GAW | Steel flask | Glass flask | Analysis Method |
| --- | --- | --- | --- | --- | --- |
| 1. Ethane | 1.5 months | sources of methane Natural sources Biomass burning Fossil fuel Ocean production (S. hemisphere) Trend in size of seasonal cycle Indicator of halogen chemistry) | √ | √ | GC/FID |
| 2. Propane | 11 days | source of methane Natural sources Biomass burning Fossil fuel Ocean production (S. hemisphere) | √ | √ | GC/FID |
| 3. Acetylene | 15 days | motor vehicle tracer biomass burning tracer ratios to the other hydrocarbons trends | √ | √ | GC/FID |
| 4. Isoprene | 3 hours | biosphere product sensitive to temperature/land use/climate change O3 precursor Oxidizing capacity precursor to formaldehyde | ? | ? | GC/FID PTR-MS |
| 5. Formaldehyde | 1 day | indicator of isoprene oxidation biomass burning comparison with satellites trends | - | - | DOAS |
| 6 Terpenes | 1-5 hours | precursors to organic aerosols | - | - | GC/MS PTR-MS |
| 7. Acetonitrile | 0.5-1 year | biomass burning indicator biofuel burning indicator | - | ? | GC/MS PTR-MS |
| 8. Methanol | 12 days | sources in the biosphere (methane oxidation) abundant oxidation product | - | ? | GC/FID PTR-MS |
| 9. Ethanol | 4 days | tracer of alternative fuel usage | - | ? | GC/FID PTR-MS |
| 10. Acetone | 1.7 months | abundant oxidation product free radical source in the upper troposphere | ? | ? | GC/FID PTR-MS |
| 11. DMS | 2 days | major natural sulphur source sulphate aerosol precursor tracer of marine bioproductivity | ? | ? | GC/FID PTR-MS |
| 12. Benzene | 10 days | tracer of combustion biomass burning indicator | √ | ? | GC/FID GC/MS |
| 13. Toluene | 2 days | ratio to benzene used for air mass age precursor to particulates | - | ? | GC/FID GC/MS |
| 14. Iso and normal Butane | 5 days | chemical processing indicator lifetime/ozone production | √ | √ | GC/FID GC/MS |
| 15. Iso/normal Pentane | 3 days | ratio provides impact of NO3 chemistry | √ | √ | GC/FID GC/MS |

GC/FID is Gas Chromatography - Flame ionization detection.

GC/MS is Gas Chromatography - Mass Spectrometry

DOAS is Differential Optical Absorption Spectroscopy

PTR-MS is Proton Transfer Reaction Mass Spectrometry

√ indicates state of current practice

Measurements of low molecular weight aliphatic and aromatic hydrocarbons (C2-C9) have been made successfully for many years, predominantly in short-term regional experiments. The preferred analytical method for these compounds, which include the molecules 1-4 and 12-15 of table 1, is gas chromatography with flame ionisation detection (GC/FID). Air samples, from flasks or in situ, are normally pre-concentrated using cryogenic methods or solid adsorbents. An alternative technique is gas chromatography-mass spectrometry (GC/MS). Although GC/MS is potentially the more sensitive method, it is typically subject to greater analytical uncertainties (changes in instrument response over time, detection of common, low mass fragments). However, GC/MS may be valuable for the detection of certain hydrocarbons in very remote locations where ambient levels may be below the detection limit of a typical GC/FID.

The recommended analytical technique for monoterpenes is GC/MS. Although it is possible to measure some terpenes using an FID detector, the complexity of the chromatographic analysis (co-eluting peaks, particularly with aromatics) makes peak identification and quantification difficult. GC/MS gives better sensitivity.

Oxygenated hydrocarbons, including the target compounds 8-10 (Table 1), can also be measured using GC/FID or GC/MS. Particular care should be taken with sample preparation (including water removal) and inlet systems must be designed to minimize artifacts and component losses commonly encountered with oxygenate analysis.  Acetone and methanol can also be measured using proton transfer reaction mass spectrometry (PTR-MS). An advantage of PTR-MS is that it is an online method that does not require the preconcentration of samples. However, it is less sensitive than GC methods and there are potential interferences from isobaric compounds, e.g. O2H+ and methanol. As the stability of oxygenated VOCs in grab samples (stainless steel or glass flasks) remains highly uncertain, it is suggested that these species be measured primarily by on-line methods at a selection of ground-based measurement stations. The successful storage of acetone in certain flasks has been reported, so the possibility of analyzing this compound in the glass or stainless steel flask network should be investigated.

Formaldehyde (HCHO) is not stable in flasks and has to be measured in situ.  Methods of analysis include the Hantzsch fluorimetric (wet chemical) method or DOAS (Differential Optical Absorption Spectroscopy). Both are relatively complex and would require specialist training for potential operators. It is unlikely therefore to be able to make measurements at more than a few ground stations. Formaldehyde is routinely detected by satellites. Satellite retrievals yield total vertical column amount and an important objective of the GAW Programme would be to provide periodic ground-based measurements at selected sites for comparison/calibration purposes (ground-truthing).

The feasibility of HCHO measurements with PTR-MS (Whistaler et al., 2008; Warneke et al., 2011) and QCL (Herndon et al., 2007) was shown during limited measurement campaigns. Their applicability for long-term routine HCHO measurements is not tested yet.

Acetonitrile is preferably measured with GC/MS, because this compound is relatively insensitive to FID detection.Measurements of acetonitrile have also been reported using various reduced gas and nitrogen-specific detectors.  Many recently reported atmospheric measurements of acetonitrile have been made using PTR-MS or AP-CIMS (Atmospheric Pressure Chemical Ionisation Mass Spectrometry).  The stability of acetonitrile in grab samples is highly uncertain so that garb sampling is not acceptable in the framework of GAW and measurements may be limited to a few selected comprehensive measurement sites.

Dimethyl sulphide (DMS) can be measured by GC/FID, GC/FPD (Flame Photometric Detector), GC/MS and PTR-MS. However, as DMS concentrations can be measured routinely as part of a standard non-methane hydrocarbon (NMHC) analysis, GC/FID analysis of the whole air samples would be the simplest choice of measurement strategy.  There is evidence in the literature that DMS is stable in some flasks so its measurement as a component of flask network is quite feasible.  It is also desirable to make in-situ measurements of DMS at least in the early stage of operation of the flask network to ensure methods compatibility.

# 16.2.3.4. Nitrogen Oxides (NOx)

The sum of nitric oxide (NO) and nitrogen dioxide (NO2) has traditionally been called NOx.  The sum of all nitrogen oxides with oxidation number greater than 1 is called NOy. Their measurement in the global atmosphere is very important since NO has a large influence on both ozone and on the hydroxyl radical (OH).  NO2 is now being measured globally from satellites and these measurements suggest that substantial concentrations of this gas are present over most of the continents. A large reservoir of fixed nitrogen is present in the atmosphere as NOy. The influence of the deposition of this reservoir on the biosphere is not well known at present but could be substantial.  There are efficient in situ measurement techniques for NO and NO2, whereas the reliability of NOy measurement techniques still needs to be improved. The widely used CLD technique with molybdenum converters gives a signal between NO2 and NOy and should be named NO2(Mo) or NO2+ (see below).

Detailed measurement guidelines for reactive nitrogen measurements are currently being developed in collaboration with ACTRIS network. The focus here is on NO and NO2 mostly because their measurements are presently more extensive and robust and allow for implementations of a fully quality assurance system. Recommendations on NO and NO2 measurements can be found in GAW Report No. 195.

Nitrogen oxides (NO and NO2) measurements can be done by passive, active and remote sensing techniques. The active techniques can be divided into integrating and in-situ techniques, with the integrating techniques consisting of a sampling step usually involving liquid-phase sample collection and off-line analysis, whereas in-situ (continuous) measurements directly analyse the sample air. Passive methods are always integrating. Active integrating methods comprise the Saltzman method and related methods like the Griess or Sodium Iodide method. The latter is being used for example in the European Monitoring and Evaluation Programme (EMEP, www.emep.int) network. Due to the high reactivity of NOx, flask sampling is impossible.

Ozone-induced chemiluminescence detection (CLD) is most widely used among the in-situ techniques. These instruments are typically very sensitive to NO, however they cannot measure NO2. Thus, NO2 must be converted to NO before detection. The instrument makes measurements in an NO mode and then an NO+NO2 mode. The difference, when conversion efficiency is determined carefully, gives the NO2 mixing ratio. Thus, a high time resolution (<10 min) is recommended to ensure sampling of the same air mass during subsequent NO and NOx measurements. The conversion of NO2 to NO is achieved by photolysis of NO2 at wavelengths 320<(<420 nm using a photolytic converter (PLC) with an arc lamp or a blue light converter (BLC) with light-emitting diodes (LEDs). Advantages of LEDs are the substantially longer lifetime and nearly constant conversion efficiencies, the mechanical simplicity, and the simple on/off characteristic of the LED (no additional valves/dead volumes). Disadvantage is the small conversion efficiency. However, new LED-based converters provide efficiencies equal to or even greater than traditional arc-lamp systems.  The use of UV-LED converters is thus recommended for GAW NO2 measurements.

The use of molybdenum converters for NO2 to NO conversion is strictly discouraged as this conversion technique is not selective to NO2 but also converts other oxidized nitrogen species in different quantities. Already existing measurements with Mo converter should be marked as NO2(Mo) or NO2+.

The Luminol-CLD measures NO2 directly and NO indirectly after oxidation. Since the sensitivity depends strongly on quality of luminol solution which decreases during use caused by aging, frequent re-calibration is needed.

In addition to these methods optical absorption techniques for NO2 detection have been developed including tuneable diode laser absorption spectroscopy (TDLAS), differential optical absorption spectroscopy (DOAS), laser induced fluorescence (LIF), Fourier-transformed infrared absorption spectroscopy (FTIR) and Cavity Ring-Down Spectroscopy (CRDS). They all measure NO2 directly. Recent developments in CRDS for the measurement of NO2 and of NO as NO2 after oxidation by ozone show some promise, but they still suffer from uncertainties in the zero level.

Recently, the suitability of research-type quantum cascade laser QCL instrumentation for continuous and direct measurements of NO and NO2 was shown (Tuzson et al., AMT, 2013). This technique may become an alternative standard method in the future.

Also recently a CAPS (Cavity Attenuated Phase Shift) monitor has become commercially available. Test at ACTRIS side-by-side intercomparison experiment showed excellent results. However, the lower detection limit (LDL) is some 50 ppt. Therefore, the instrument is very good for rural or anthropogenic influenced sites but not suitable for remote ones with typical NO2 mole fractions below 50 ppt.

At present, there is no mature technique that can compete with the ozone-induced chemiluminescence NO detection (CLD) measurement of NO at remote locations. Passive and active integrating methods are not accepted in the GAW Programme for their poor selectivity and time resolution.

# 16.2.3.5. Sulphur Dioxide (SO2)

Measurement guidelines for SO2 observations are so far not available in the GAW Programme. General recommendations are given in GAW Report No. 143. However, the GAW Scientific Advisory Group on Reactive Gases is planning to establish same quality assurance system for SO2 (including Measurement Guidelines and Central Facilities) as it is doing at present for nitrogen oxides after that item is solved.

There are various measurement techniques for determination atmospheric SO2. EMEP is using integrating techniques like using an alkaline impregnated filter (pack) or coated annular denuder; both followed by ion chromatography in central lab. These methods yield potentially more accurate results, but as typical for integrative techniques, with lower time resolution of typically one sample per day. Additionally they require frequent attention, and personnel costs for filter analysis are high.

In the group of in situ measurements the TCM (photometry after reaction of SO2 with Tetrachloromercurate) and the pulse fluorescence are widely used. The first one has a high accuracy but high LDL and the handling of mercury in the lab could be harmful.  Even though the response of the pulsed-fluorescence sensor is slower, its ease of calibration, dependability, accuracy, and SO2 specificity make it preferable.  More sensitive gas chromatographic techniques are also available.  However, they require significant technical expertise and regular attention. In order to enhance the sensitivity some fluorescence analysers are equipped with more selective excitation filters, e.g. two mirror assemblies are connected in series and specially selected PMTs are employed. Numerical corrections of interfering substances could be done, however, this is not necessary in rural or remote areas. Typical LDL which can be reached with these provisions is some 50 ppt. Maybe a further enhancement of sensitivity is possible by means of a 2nd channel in which only SO2 is removed, leading to a highly specific readout after subtraction of both channels.

Since SO2 has a short atmospheric lifetime, understanding the sulphur cycle requires knowledge of the source and sink terms.  This is best accomplished with sampling frequencies of less than 1 hour.  Therefore, the best technique for long-term monitoring of SO2 today is a combination of the pulsed-fluorescence analyser and filter sampling.  Filter samples would be exposed at intervals, but often enough to act as a quality control for the continuous analyser.

# More detailed recommendations can be found in the EMEP Manual for Sampling and Chemical Analysis, Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 1/95), available at http://www.nilu.no/projects/ccc/manual/index.html.

# 16.2.3.6. Molecular Hydrogen (H2)

The detailed Measurement Guidelines for molecular hydrogen are currently not available in the GAW Programme.

Molecular hydrogen is reported as dry mole fraction on the most recent scale (GAW Report No. 206). Measurements of H2 are possible both in-situ and by flask collection with subsequent analysis in the laboratory. An example of the measurement system set up at the GAW Global stations in given in Grant et al. (2010).

Molecular hydrogen measurements are performed with gas chromatography followed by hot mercuric oxide reduction/UV absorption detection (GC/HgO). An alternative GC set-up with pulsed-discharge detectors (PDD) has a more linear detector response and provides better repeatability for molecular hydrogen measurements.

Problems with instability of H2 in reference gases have frequently been experienced. Therefore, recommendations on calibration and quality assurance of H2 measurements available in GAW Report No. 206 should be consulted.

# 16.2.4. Atmospheric Wet Deposition

Atmospheric wet deposition refers to the gases and particles deposited by precipitation to the earth’s surface. These gases and particles have a wide variety of sources and compositions and generally are present in trace amounts in the atmosphere and in precipitation. These trace materials are captured by precipitation as it forms in the atmosphere and falls to the earth. These deposited materials constitute an important contribution to the mass balance of pollutants associated with long-range transport. These materials not only affect the chemistry of precipitation but also can affect the chemistry of the terrestrial and aquatic surfaces on which they are deposited. The effects can be harmful or beneficial and they can be direct or indirect. For example, acidic wet deposition is an environmental problem that results from combustion of fossil fuels. It occurs when oxides of sulphur and nitrogen, emitted during combustion, are transformed in the atmosphere and become acidic sulphate and nitrate in precipitation. Other trace materials that occur in wet deposition include sea salt, nutrients, chemicals found in soil particles, toxic organic and inorganic chemicals, organic acids, etc. Research has shown that some wet-deposited chemicals can stimulate marine biotic production, potentially linking atmospheric wet deposition to the carbon cycle and climate change.

Measuring the chemistry of precipitation tells us what trace materials are present in wet deposition and in what amounts. This information can be used to evaluate air quality and to identify and track changes in gaseous and particulate emissions to the atmosphere. In short, precipitation chemistry measurements provide information on the exchange of trace materials between the atmosphere and the land/oceans, and hence are important in furthering our understanding of the chemical cycles of these materials, especially those that can result in damage to terrestrial and aquatic systems or affect our climate.

Special care is required when planning precipitation chemistry measurements to ensure that the measurements are representative. Though these measurements are made at a particular location, on average they should represent measurements in the surrounding region. In general, the sample collection site should be characteristic of the land use in the region. For example, the site in an area dominated by agricultural activities should have an agricultural setting. This quality of spatial representativeness should extend across seasons and even over years. Ideally, a site would be both spatially and temporally representative. Contamination of a localized nature from agricultural, industrial, or other human activities must be avoided, as must the local impact of natural sources, such as oceanic shores, volcanoes, fumaroles, etc. Sample collection should not be impacted by trees or other vegetation and the on-site topography should be level and the exposure relatively unaffected by wind patterns that may result in an unrepresentative catch of rain and snow. Human contact with the sample or contact with anything that might change the sample chemistry must be avoided, as well.  Ensuring representative precipitation chemistry measurements entails strict adherence to requirements for site location, site conditions, equipment installation, and site operational protocols and maintenance. These requirements are documented in the Manual for the GAW Precipitation Chemistry Programme, (Guidelines, Data Quality Objectives and Standard Operating Procedures) (GAW Report No. 160).

Precipitation chemistry monitoring can be divided into activities associated with sample collection, and activities associated with chemical analysis.

# 16.2.4.1. Sample Collection

The primary goal of the GAW Precipitation Chemistry Programme is to collect wet-only deposition samples. This means that the samplers are exposed only during the occurrence of precipitation and trace materials in the samples are deposited only by precipitation. The trace materials from dust or fine particles or gases that are deposited during dry weather are excluded. This makes possible the study of precipitation chemistry without contamination from dry deposition. More importantly, the equipment and methods for collecting a representative wet deposition sample are inappropriate for collecting a representative dry deposition sample. The physical and chemical processes affecting wet and dry deposition are distinctly different from one another.

The best way to ensure collection of a wet-only sample is to employ an automated sampler that is open only during precipitation. A typical automated, wet-only deposition sampler has the following components: a precipitation sample container (funnel-and-bottle, bucket, etc.), a lid that opens and closes over the sample container orifice, a precipitation sensor, a motorized drive mechanism with associated electronic controls, and a support structure to house the components.  Precipitation sample containers should have sufficient volume to hold all precipitation collected during the sampling period.  A system that can be activated manually for testing, cleaning, and routine maintenance is recommended.  A modular design that allows removal of individual components, such as the sensor, facilitates rapid repair with a minimum of tools and expertise. An alternative to using an automated sampler is to collect samples by manually exposing a sample container at the very onset of precipitation and closing the container as soon as precipitation ceases to occur. This requires diligent round-the-clock observers alert to weather conditions; and as a consequence, manual sample collection is very labor-intensive.

Complementing the collection of wet-only deposition samples, the GAW Precipitation Chemistry Programme requires every site to measure precipitation depths using the NMHS-designated standard precipitation gauge or its equivalent (See chapter 6 of CIMO Guide). Manual gauges are preferred.  Precipitation depths are used to calculate the mass of a chemical deposited by precipitation on an area of earth surface (i.e., the wet deposition flux or loading).  Standard precipitation gauges are designed to be the most accurate and representative means of measuring precipitation depths.  Thus, each site must operate a precipitation gauge in parallel with its precipitation chemistry sampler.  Precipitation chemistry sampler volumes are used to calculate wet deposition fluxes only when the standard gauge fails or is temporarily out of service.  The data record should document such cases.

The highest priority of the GAW Precipitation Chemistry Programme is to collect a wet-only sample on a daily (24-hour) basis with sample removal set at a fixed time each day, preferably 0900 local time. Should the resources be inadequate to collect and analyse daily samples, multi-day sampling periods up to one week is the next highest priority. Alternative sampling protocols are described in the Manual for the GAW Precipitation Chemistry Programme (GAW Report No. 160).  Collecting samples daily reduces the potential for the degradation of labile chemical species and for other physical and chemical changes in the sample while it is held in the field sampler. Not only is the sample integrity less likely to be compromised by a daily sampling protocol but the data have greater value, as well. Storm trajectory analyses and source-receptor models are much less complicated when precipitation is more likely to have come from a single event or storm. Multi-day and one-week samples are much more likely to contain precipitation from several storms, each occurring under different meteorological settings. Further, daily data can be integrated mathematically to determine weekly or longer term averages, but weekly data cannot be differentiated into its daily components without making substantial assumptions.

Containers used to collect, store, and ship samples should be unbreakable and sealable against leakage of liquids or gases.  High density polyethylene (HDPE) containers are recommended. All sample containers must be cleaned with deionised water of known and assured quality. The Manual for the GAW Precipitation Chemistry Programme (GAW Report No. 160) contains detailed descriptions of the procedures for cleaning containers and ensuring that cleanliness standards are maintained throughout the collection, storage, and shipment of samples.

# 16.2.4.2 Chemical Analysis

The following chemical parameters are recommended for analysis in GAW precipitation samples: pH, conductivity, sulphate, nitrate, chloride, ammonium, sodium, potassium, magnesium and calcium.  Analyses for formate and acetate are recommended for areas suspected of having high organic acid concentrations.  Nitrite, phosphate and fluoride concentrations also may be important in certain areas, although their analyses are not required by GAW at this time.  Preferred analytical methods are given in Table 2.

Table 2. Chemical parameters required for analysis in the GAW precipitation chemistry programme with recommended analytical methods.

| Analyte | Status | Preferred Methods1 |
| --- | --- | --- |
| pH | Required | Glass electrode |
| Conductivity | Required | Conductivity cell |
| Alkalinity | Optional | Titration |
| Cl-, NO3-, SO42- | Required | IC |
| NH4+ | Required | IC, AC |
| Na+, K+, Ca2+, Mg2+ Organic Acids2 NO2-, F- PO43- | Required Optional Optional Optional | IC, ICP, AAS/AES IEC, IC IC IC, AC |

1	IC = ion chromatography; AC = automated colorimetry; ICP = inductively coupled plasma spectrometry, IEC = ion exclusion chromatography; AAS = atomic absorption spectrometry; AES = atomic emission spectrometry.

2 	For areas with high organic acid concentrations, formate and acetate analyses are recommended.

Past experience from regional networks and laboratory intercomparisons has shown that measuring pH in precipitation is difficult due mainly to the low ionic strength of the samples.  Samples may also degrade due to biological activities and should therefore be kept refrigerated until the time of analysis, when they are brought to room temperature.  The pH measurements should be carried out within two days of sample arrival in the laboratory.

Commercial pH meters are available with different specifications and options.  A pH meter should have both an intercept and slope adjustment and should be capable of measuring to within ±0.01 pH unit.  Combination electrodes containing both measuring and reference functions are often preferred since they require smaller amounts of sample, but a set of two electrodes may also be used with the pH meter.  The measuring glass electrode is sensitive to hydrogen ions and the reference electrode can be calomel, or silver/silver chloride.  Low ionic strength electrodes are now available commercially.  Other reference electrodes can also be used as long as they have a constant potential.  When selecting any electrode, confirm its ability to measure low ionic strength solutions by measuring a certified reference material.  Response time should be less than one minute and addition of KCl should not be needed.

The conductivity of a solution is the reciprocal value of its specific resistance and can be directly measured using a conductivity bridge with a measuring cell.  Conductivity varies with the temperature of the solution and is proportional to the concentration and the species of free ions present in the solution.  Since the conductivity also depends on the electrode area and its spacing, the measuring apparatus has to be calibrated to obtain the cell constant or to adjust the meter.  A KCl solution of known concentration and conductivity is used for calibration.  Conductivity is measured and expressed in units of microsiemens per centimetre (µS cm-1), corrected to 25°C.  The conductivity range of precipitation samples is 5 to 1000 µS cm-1. In case of small sample volumes, the aliquot that is used for conductivity measurement can be used for pH determination.  If this is done, the conductivity should be measured first to avoid any possible error due to salt contamination from the pH electrode.

The apparatus for conductivity measurements consists of:

- Conductivity meter (operating range of 0.1 to 1000 µS cm-1; better 0.01 to 1000 µS cm-1).  Precision has to be within 0.5% of the range and accuracy at 1% of the range.

- Conductivity cell (Use special conductivity cells, with a low cell constant, if the values in precipitation samples are expected to be mainly very low (<20 µS cm-1).

- Thermometer (0 to 40°C / 0.1°C).

- Water bath 25°C.

- Polyethylene or glass vessel corresponding to the diameter of the cell used.

Ion chromatography has been widely used in recent years to analyze major anions and cations in precipitation, mainly in combination with electrochemical detection.

Sulphate, nitrate, chloride, as well as other anions in precipitation are separated on an ion exchange column because of their different affinities for the exchange material.  The material commonly used for anion separation is a polymer coated with quaternary ammonium active sites.  After separation, the anions pass through a suppressor that exchanges all cations for H+ ions.  Instead of strong acid cation exchange columns, today micro membrane and self-regenerating suppressors with chemical or electrochemical regeneration are used.  As the result of the suppression reaction, corresponding acids of the eluent ions and of chloride, nitrate and sulphate will reach the conductivity detector.  A decreased basic conductivity and higher analytical signals now allow the detection of anions also in the lower µg L-1 range.

There are several anion exchangers with different properties available on the market.  The time for one analysis and the quality of separation of single signals are dependent on the type of the column and the eluent, the concentration and the flow rate of the eluent.

Any anions with a retention time similar to that of the main anions in solution can cause an interference.  For example, when NO2- is present, it elutes just after Cl-, which can cause the peak to be asymmetric.  In rare cases, when the concentration of Cl- is very high compared with NO3-, it can also influence the determination of NO3-.  The manual should be consulted to see how different integration programmes handle this problem.

With care, up to several thousand analyses can be performed with the same anion separator column.  The most effective method of protecting the separator column is the use of a pre-column in front of it.  Details are provided by the manufacturers in the manuals for the columns.

The principle of cation measurements is the same as that of anion determination except that different column materials are used and the suppressor column is often omitted.  The material commonly used for cation separation is a cation exchange resin with active surface groups.  Sodium, ammonium, potassium, calcium and magnesium ions are detected by a conductivity detector, without changing the eluent when certain columns are used.  In other columns monovalent cations (Na+, NH4+, K+) are determined using one eluent and divalent cations (Mg2+ and Ca2+) with another eluent (because of their higher affinity to the resin).

Any cation with a retention time similar to that of the main cations may cause interference.  For example, in samples with high concentrations of Na+, the peak of NH4+ becomes asymmetrical and often causes significant error.  In this case, measurement using more dilute eluent could improve the separation of peaks.

Sodium, potassium, magnesium, and calcium in precipitation are often analysed by atomic spectroscopic methods.  Both flame (Atomic Absorption Spectrometry, AAS and Atomic Emission Spectrometry, AES) and plasma (inductively coupled plasma - atomic emission spectrometry, ICP-AES and inductively coupled plasma - mass spectrometry, ICP-MS) based methods can be used.  For these ions, ion chromatography has no special advantage concerning sensitivity, precision and accuracy over the spectroscopic methods, although analysis of all ions in one sample run is not possible with flame AAS or AES (single element methods).

The ions in the sample solution are transformed to neutral atoms in an air/acetylene flame.  Light from a hollow cathode or an electrode-less discharge lamp (EDL) is passed through the flame.  In the AAS-mode, light absorption of the atoms in the flame is measured by a detector following a monochromator set at the appropriate wavelength.  Light absorption is proportional to the ion concentration in the sample.  In the AES-mode, the light emitted from the atoms excited in the flame is measured.  Most commercial instruments can be run in both modes.  AES is the preferred mode for sodium measurements.

In atomic absorption spectrometry, both ionization and chemical interferences may occur.  These interferences are caused by other ions in the sample, which reduce the number of neutral atoms in the flame.  Ionization interference is avoided by adding a relatively high amount of an easily ionized element to the samples and calibration solutions.  For the determination of sodium and potassium, cesium is added.  For the elimination of chemical interferences from aluminum and phosphate, lanthanum can be added to the samples and calibration solutions for calcium and magnesium.

Formic and acetic acids (HCOOH and CH3COOH, respectively) are major chemical constituents of precipitation in both continental and marine regions.  Available evidence suggests that these compounds originate primarily from natural biogenic sources; both direct emissions (over continents) and emissions of precursor compounds appear to be important.  Biomass and fossil fuel combustion also result in the emission of carboxylic acids and/or their precursors to the atmosphere.

Carboxylic acids in precipitation are very unstable and rapidly disappear from unpreserved samples.  To generate reliable data, precipitation must be sampled on a daily or event basis and immediately preserved with addition of a biocide such as chloroform (CHCl3).  Typically, 250 mL aliquots of sample (or less for low volume events) are treated with 0.5 mL of CHCl3. Samples are then tightly sealed and refrigerated until analysis.

Carboxylic species can be analyzed by both ion chromatography (IC) (using a dilute eluent) and ion exclusion chromatography (ICE).  However, acetate and propionate typically co-elute when analyzed by IC and are thus impossible to resolve quantitatively.  ICE exhibits fewer interferences associated with co-eluting species and is thus preferred for analysis of precipitation samples.

For analysis by ICE, samples are added to an HCl eluent which then flows through a separator column, a suppressor column and a detector.  Resin in the separator column partitions anions using the principle of Donnian exclusion; anions are retained and sequentially separated based on their respective pKas and Van der Waals interactions with the resin.  Anions of stronger acids with lower pKas such as H2SO4, HNO3 and HCl are effectively excluded and co-elute early in the chromatogram; those of weaker acids with higher pKas such as HCOOH and CH3COOH elute later in the chromatogram.  The suppressor column incorporates a cation exchange resin with silver added to the exchange sites; H+ exchanges with the silver; the released silver subsequently reacts with Cl- in the eluent to form AgCl, which precipitates within the column.  Acid analytes exit the suppressor in a stream of deionized water.  Detection is by conductivity.

16.2.5. Aerosols

Atmospheric aerosols are important for a diverse range of issues including global climate change, acidification, regional and local scale air quality, and human health.  The climate impact of aerosols is a result of both direct radiative effects and indirect effects on cloud properties.  Regional problems include potential impacts on human health and mortality and environmental impact such as visibility impairment. Major sources of aerosols include urban/industrial emissions, smoke from biomass burning, secondary formation from gaseous aerosol precursors, sea salt and dust. Outstanding problems include determining the natural sources of aerosols, and the organic fraction.

Table 3 provides a list of recommended aerosol parameters recommended for measurements in the GAW Programme. Comprehensive Measurement Guidelines for aerosol measurements are provided in GAW Reports No. 153 and 200.  Much of the information in the remainder of this section is taken directly from GAW Report No. 153, and may become obsolete when a major update of that report is released in 2014.

Table 3. 	List of comprehensive aerosol measurements that are recommended for long-term measurements in the global network (GAW Report No. 197).

| Variable | Frequency of observation |
| --- | --- |
| Multiwavelength aerosol optical depth | Continuous |
| Mass concentration in two size fractions (fine, coarse) |  |
| Mass concentration of major chemical components in two size fractions |  |
| Light absorption coefficient at various wavelengths |  |
| Light scattering and hemispheric backscattering coefficient at various wavelengths |  |
| Aerosol number concentration |  |
| Aerosol number size distribution |  |
| Cloud condensation nuclei number concentration at various super-saturations |  |
| Vertical distribution of aerosol backscattering and extinction |  |
| Detailed size fractionated chemical composition | Intermittent |
| Dependence of aerosol variables on relative humidity, especially aerosol number size distribution and light scattering coefficient. |  |

# 16.2.5.1. Aerosol Chemical Measurements

At present, filter collection of ambient aerosol, followed by laboratory analyses still remains the most commonly used and cost-efficient method for the determination of aerosol chemical composition despite well-documented artifacts. Artifacts are very often linked to presence of semi-volatile species that can either condense upon sampling (positive artifact) or evaporate from the filter media after sampling (negative artifact). A number of methods have been proposed to limit these artifacts but none of them will satisfactorily apply to all chemical substances present in aerosol.

Optimal set-up for characterization of chemical properties of aerosol would be composed of a series of denuders in order to remove condensable species present in the gas-phase (limiting positive artifact) and a filter pack collecting both particles and condensable species re-emitted from the first filter (accounting for negative artifact). Filter packs have been developed, consisting of a sandwich of filters and collection media of various types in series, to collect aerosols and selectively trap gases and aerosol volatilization products. Ideally, sampling for inorganic and carbonaceous species is performed with 2 different sampling lines since different kind of denuders and filter media are required for analyses of EC/OC and inorganic species. A third line could be implemented for sampling and analysis of elemental aerosol composition.

Clearly, methods for artifact limitation can be rather impractical and are relatively expensive. For GAW purposes, considering remoteness and available resources in a number of sites use of denuders should be considered as a simple recommendation, and not a requirement. This means, however, that sampling artifacts exist for a number of semi-volatile species, in particular when temperatures in the sampling system exceed 20°C. This should be accounted for in reporting data to World Data Center for aerosols.

High volume and low-volume sampling lines are accepted in the context of GAW. For simplicity, it is suggested that a differencing technique be utilized to separate the coarse fraction from the fine fraction. Specifically, one filter should be run behind the 10 (m aerodynamic diameter cut inlet. A parallel filter would be run behind the inlet suitable for the fine fraction (i.e. 2.5 (m aerodynamic diameter at ambient RH or 1.0 (m in dry air).  While the second filter will yield the fine fraction, the difference between the two filters will then yield the coarse fraction. For high-volume sampling, use of dichotomous samplers is an interesting alternative to differentiate fine and coarse aerosol fractions.

Low-volume samplers are more easily implemented than high-volume samplers. For the routine long-term aerosol measurements at GAW stations, it is recommended that up to three sets of 47 mm diameter filters should be collected in parallel by low-volume samplers. If financial constraints are limiting, the priorities for filter sampling are (a) Teflon filters for gravimetric and ionic analyses; (b) quartz-fiber filters for carbonaceous aerosol analyses; (c) Teflon filters for elemental analyses. Each set would consist, ideally, of two filters, one for total mass below 10 (m diameter and one for the fine fraction. The separation would be achieved by running the filters behind the size-selective inlets. High-volume samplers are usually more expensive and running more than one set of sampler in parallel is often unpractical. For that reason, it is recommended that high-volume samplers are operated with quartz fiber filter for both inorganic and EC/OC analyses. It should be noted that denuders for high-volume samplers are not commercially available.

There is no recommendation for sampling time integration as this will highly depend upon sites. In general, information provided by short sampling time (24 to 48h) are more easily used in models and should be preferred over week-long sampling, even if discontinuous. Filters should be removed from the sampling unit shortly after collection and stored between 0 and 5°C if analyses cannot be performed immediately. Regular blanks should be performed in order to control for contamination. Ideally, blanks are prepared by mounting filters into the sampling unit with the pump off. We recommend that 1 blank is performed every 10 sample.

For each GAW aerosol station, a list of core aerosol chemical measurements is strongly recommended: (i) mass; (ii) major ionic species; (iii) carbonaceous components, and (iv) dust aerosols.

- The mass concentration of atmospheric aerosols is clearly a fundamental parameter in the GAW measurement programme.  It is recommended that this be done gravimetrically on Teflon filters.  It is expressed in units of µg m-3 where the volume is related to standard temperature and pressure. Updated measurement guidelines on the mass concentration measurements using gravimetric analysis of Teflon filters are provided in GAW Report No. 200.

The TEOM (Tapered Element Oscillating Microbalance) has been widely used for aerosol mass measurements.  The instrument provides continuous measurements and can produce high time resolution data.  The original TEOM has a recommended temperature setting of 50oC.  There is evidence that semi-volatile components of aerosols are lost from the TEOM measurements at this temperature. The most likely explanation is that semi-volatile organic compounds, nitrates, and water are lost from the aerosols during sampling. In newer TEOM models, the operating temperature has been lowered to 30oC to reduce losses of the semi-volatile components.  In addition, sudden changes in ambient relative humidity may cause negative mass readings from the instrument. New modifications have been made to dry the aerosol particles before measurement to reduce the effect of ambient relative humidity.

A different type of instrument for continuous mass measurements, the β-meter, has also been commercialized. It operates on the principle of β -ray attenuation by a layer of aerosol.  The β -ray source is usually 14C or 85Kr decay and the attenuation can be calibrated with a known mass.  Sampling can be done with individual filters or filter tapes, and the β -ray that passes through the filter is continually monitored.  The β -meters have the same inherent difficulties concerning volatilization as the TEOM. However, comparison with gravimetric methods usually produces reasonable agreement. Updated guidelines on the mass concentration measurements with Beta Attenuation (with Met One Instruments BAM-1020) are provided in GAW Report No. 200.

-  The concentration of major inorganic species is one of the core information recommended for the GAW stations.  Major ionic species include sulphate, nitrate, chloride, sodium, ammonium, potassium, magnesium, and calcium. This selection of analyses is based on the fact that analytical procedures for them have become well established.  More importantly, under most atmospheric conditions, this set of ions is expected to account for a major part of the aerosol mass, and the measurements here are an important step toward mass closure of the aerosols. As mentioned above, quantitative measurements of nitrate with filter technique remain problematic and high uncertainties are associated to this determination.

GAW recommends that analyses are done using ion chromatography (IC) for the most cost-effective approach.  The IC technique has the advantage of chemical speciation and relatively low cost per analysis, and has matured to the degree that the sensitivities for each ionic species, cost, and maintenance are all reasonably well known.  If IC is set up properly, all the recommended ionic species can be analyzed in one single sample injection. Alternative analytical techniques exit but their use may introduce systematic differences amongst GAW stations. It is part of each laboratory’s responsibility to document the equivalence of these alternative techniques with IC whenever other techniques are used as for example, atomic absorption spectroscopy (AAS) or inductively-coupled plasma mass spectroscopy (ICP-MS).  Calibration of IC instruments is an integral part of every laboratory’s standard operating procedures and each laboratory must implement quality control procedures that guarantee the accuracy of calibrations. Recommendation for GAW is similar to those reported in GAW Report No. 160 for precipitation chemistry. In addition, protocols for filter extraction should be well documented.

Instruments are emerging on the scene for continuous and semi-continuous measurement of sulphate, nitrate, and organic carbon in aerosols. In particular, progresses in aerosol mass spectrometry (AMS) have led to the development of instruments that can be used for monitoring purposes, providing quantitative measurements of the total mass and size distribution of non-refractory chemical composition in the submicron size range. For example, the Aerosol Chemical Speciation Monitor (ACSM) is a simple version of AMS that can provide chemically speciated mass loadings and aerosol mass spectra with 1-hour time resolution. No standard operating procedure exist yet for ACSM or other AMS instruments but its suitability for long-term monitoring is clear and will undoubtedly bring revision of aerosol chemical speciation procedures in the future.

-  The concentration of carbonaceous species (with both Elemental and Organic fractions) is also one of the core information recommended at GAW stations. Carbonaceous species are still the least understood and most difficult to characterize of all aerosol chemical components. Total aerosol carbon mass (TC) can be divided into three fractions: inorganic carbonates, organic carbon (OC), and a third fraction ambiguously called elemental carbon (EC), black carbon (BC), soot, or refractory carbon in the scientific literature with no clear definition of the terms. Recommendations for proper use of terminology for BC-related species have been proposed by Petzold et al., (2013) to clarify the terms used for BC in atmospheric research; the recommendations state that the term black carbon be used only in a qualitative sense, and that terms related to the measurement technique be used when reporting quantitative results.  According to this terminology, thermo-optical methods can be used to derive Total Carbon (TC) and OC/EC fraction in atmospheric aerosol filters.  When using optical methods, the light-absorbing component is called equivalent black carbon, even though the optical method is not specific for carbon.

It is recommended that TC, OC and EC be measured in the GAW program, leaving out the relatively minor and difficult inorganic carbon component and the more complicated issues of organic carbon speciation.  Sampling of aerosol carbonaceous materials is recommended using quartz filters, pre-fired at 350-400oC for 2 hours, and at the same sampling frequency as the Teflon filters. The quartz filter can be analyzed for TC using the thermal evolution technique. The mass concentration of the total carbon is obtained by thermal oxidation of the carbon, usually at 750oC in the presence of a catalyst, to measurable carbon dioxide.  Detection of the evolved carbon dioxide is done in one of two ways, either by reduction to methane in the presence of a catalyst, and flame ionization detector (FID), or by direct detection by non-dispersive infrared (NDIR) detectors.

The measurement of the TC components OC and EC is more difficult than the measurement of TC (Schmid et al., 2001).  The distinction between the fractions is made by temperature-controlled volatilization/pyrolysis. This is followed by catalyzed oxidation to CO2 and detection by NDIR, or in some instruments, further catalyzed reduction of the CO2 to CH4 and final detection by FID.  There are different temperature-control programs in use. At present, GAW recommends the use of one of three alternative thermo-optical techniques:  IMPROVE (Interagency Monitoring of Protected Visual Environments) protocol (Chow et al., 1993, 2005), NIOSH (National Institute of Occupational Safety and Health) protocol (Birch and Cary, 1996) and EUSAAR-2 (European supersite for Atmospheric Aerosol Research) protocol (Cavalli et al., 2010). Relatively good agreement is obtained between NIOSH, IMPROVE and EUSAAR-2 protocols for TC determination while they strongly differ for EC determination (Chow et al., 2001, Cavalli et al., unpublished data). Because EC represents a relatively small fraction of TC, OC determination by the 3 protocols is also comparable. It is accepted that IMPROVE and EUSAAR-2 protocols are most suited for non-urban background sites while the NIOSH is applied to samples from urban sites in USA-EPA’s. The use of IMPROVE or EUSAAR-2 may therefore be preferred over NIOSH for global remote GAW stations.  Use of EUSAAR-2 and IMPROVE protocols for EC determination may lead to different results but should be preferred over any other alternative techniques. Whenever an alternative approach is used for OC measurements, it is recommended that a periodic determination of OC using one of the thermo-optical methods is conducted so that results can be compared. This is also true for AMS techniques before a standardized protocol is well established.

Use of optical methods for estimating equivalent BC (EBC) involves measuring the change in optical transmission of a deposit of particles on a filter (absorption) and applying a site-specific and instrument-type- specific mass absorption efficiency to derive EBC. Two key assumptions are required to derive the equivalent black carbon mass concentration from light absorption measurements: (i) black carbon is the only species responsible for the aerosol light absorption and (ii) the sampled black carbon has the same mass absorption efficiency as the standards used in laboratory calibrations of the absorption instrument.  These assumptions can be evaluated by experimentally determining the mass absorption efficiency by simultaneously making light absorption measurements and EC measurements as described above.  At sites where EC concentrations are not routinely determined on quartz-fibre filters, less frequent filter collections can be used to derive site- and season-specific values of the mass absorption efficiency.  Thus, for GAW measurements of EBC experimentally derived values of the mass absorption efficiency at a site are essential when estimating black carbon mass concentration from light absorption measurements.

Use of incandescent methods, such as Single Particle Soot Photometers, or volatility techniques such as Volatility Scanning Mobility Particle Sizer (SMPS) can provide information on refractory material present in aerosol but, their use in monitoring activities at GAW stations remains problematic given the lack of standardized protocols and of consistent intercomparison with thermo-optical techniques.

- Dust aerosols can be sampled relatively easily without the problems posed by more semi-volatile aerosol components such as organics and ammonium nitrate.  For GAW stations, it is recommended that a multi-elemental analysis approach should be used to determine the mineral dust component. Teflon filters should be analyzed for at least four of the major crustal elements Al, Si, Fe, Ti, Sc and the related elements Na, Mg, K, Ca.  No specific analytical technique is recommended as there is a good selection available including Proton Induced X-ray Emission (PIXE), Instrumental Neutron Activation Analysis (INAA), X-Ray Fluorescence (XRF), AAS and ICP-MS. These techniques usually have high sensitivities for the crustal elements.  Not all techniques can provide all the required elements, and depending on availability, a combination of two or more techniques may be necessary.

# 16.2.5.2. In-Situ Measurements of Aerosol Radiative Properties

The following aerosol radiative properties are needed for climate studies, all at multiple wavelengths across the visible spectrum:

1) aerosol light extinction coefficient and its two components (scattering and absorption).

2) δ - Aerosol optical depth (see 16.2.5.5), defined as the integral over the vertical column of the aerosol light extinction coefficient.

3) o - The aerosol single-scattering albedo, defined as sp/(ap+sp), describes the relative contributions of scattering and absorption to the total light extinction.  Purely scattering aerosols (e.g., sulphuric acid) have values of 1, while very strong absorbers (e.g., black carbon) have values of around 0.3.

4) g, Radiative transfer models commonly require one of two integral properties of the angular distribution of scattered light (phase function):  the asymmetry factor g or the upscatter fraction .  The asymmetry factor is the cosine-weighted average of the phase function, ranging from a value of -1 for entirely backscattered light to +1 for entirely forward-scattered light.  The upscatter fraction  gives the fraction of sunlight scattered in the upwards direction (back to space), which depends on the solar zenith angle as well as the size distribution and chemical composition of the particles.  can be estimated from hemispheric backscatter fraction b.

5) si - The mass scattering efficiency for species i, often calculated as the slope of the linear regression line relating sp and the mass concentration of the chemical species (though multiple linear regression is preferred, to deal with co-variance of some chemical species), is used in chemical transport models to evaluate the radiative effects of each chemical species prognosed by the model.  This parameter has units of m2 g-1.

6) ai - The mass absorption efficiency, for species i, often calculated as the slope of the linear regression line relating ap and the mass concentration of the chemical species (though multiple linear regression is preferred, to deal with co-variance of some chemical species), is used in chemical transport models to evaluate the radiative effects of each chemical species prognosed by the model.  This parameter has units of m2 g-1.

7) f(RH)- The functional dependence of components of the aerosol light extinction coefficient (ep, sp, ap) on relative humidity, expressed as a multiple of the value at a low reference RH (typically <40%).

Aerosol light scattering coefficient, sp, is measured with an integrating nephelometer.  Integrating nephelometers have been operated at baseline monitoring stations since the deployment of a four-wavelength instrument at NOAA's Mauna Loa Observatory in 1974.  At present, there are about four dozen sites monitoring sp routinely around the globe as part of the GAW global network.  A few of these are operating single-wavelength units, but most are measuring sp at three wavelengths.  The multi-wavelength integrating nephelometer TSI model 3563 operates at wavelengths of 450, 550, and 700 nm, and has the added feature of being able to measure sp over two angular ranges: total scattering (7-170° degrees) and hemispheric backscattering (90-170°, denoted as bsp); the Aurora 3000 integrating nephelometer manufactured by Ecotech makes comparable measurements.  While instruments do not exist for direct determination of the asymmetry parameter, g, or upscatter fraction, , the ratio b = bsp/sp can be used to estimate either of these parameters (updated measurement guidelines are available as GAW Report No. 200). Simpler, less expensive, less sensitive, one-wavelength instruments are also commercially available.  These instruments can provide useful information on aerosol light scattering coefficient at regional sites where aerosol loadings allow the use of a less sensitive instrument.

Instruments capable of high time-resolution determination of aerosol light absorption coefficient are commercially available, based on the rate of change of transmission through a fibre filter as particles are deposited on the filter.  Calibration of these filter-based methods is difficult, but is required because the relationship between change in light transmission and aerosol absorption optical depth on the filter depends on many factors, including the particular filter medium and the light-scattering nature of the particles.

- One instrument in common use is the Aethalometer. Originally, this instrument was calibrated in terms of an equivalent mass of black carbon rather than the fundamental property that provides the instrumental response: aerosol light absorption.  Early models of the Aethalometer have a very broad wavelength response, while newer versions offer narrow-band measurements at multiple wavelengths.

- Another commercial, filter-based instrument for determining ap is the Particle/Soot Absorption Photometer (PSAP) to laboratory aerosols with different single-scattering albedos, using a calibration standard based on the difference between ep, measured with a long-path extinction cell, and sp, measured with an integrating nephelometer. Updated measurement guidelines for PSAP instruments are provided in GAW Report No. 200.

- Yet another filter-based instrument is the Multi-Angle Absorption Photometer (MAAP).  The MAAP uses a different optical configuration than the Aethalometer and PSAP, with measurements of the filter reflectivity at two different angles in addition to the filter transmission measurement.  The two reflectivity measurements allow correction for multiple scattering processes involving the deposited particles and the filter matrix.  This approach eliminates the need for a correction scheme based on independent measurements of aerosol light scattering coefficient.  The MAAP operates at a wavelength of 670 nm. Updated measurement guidelines for MAAP instruments are provided in GAW Report No. 200.

Recent improvements in a different approach to determination of aerosol light absorption coefficient, photoacoustic spectroscopy, offer a promising alternative to filter-based methods.  Although not as sensitive, the photoacoustic method allows determination of aerosol light absorption coefficient while the particles are suspended in air, eliminating the artifacts introduced by depositing the particles on a filter.  The photoacoustic method can be used in regions where light absorption levels are moderately high, and as a calibration standard for filter-based instruments.

# 16.2.5.3. Particle Number Concentration and Size Distribution

Condensation nuclei can be detected after the condensation of water, or other condensable vapour (often an alcohol such as butanol) from a supersaturated atmosphere onto the particle.  The super saturation in CN counters (CNCs), which are also known as condensation particle counters (CPCs), is typically quite large, around 150%, allowing detection of particles as small as a few nanometers diameter.

Optical methods are usually used to detect the resulting droplets. Early counters relied on manual counting, either in-situ or after photography.  The Nolan-Pollak counter and its derivatives (such as the Gardner counter) rely on determination of the optical extinction of the resulting cloud.  Photometric observation of light scattered by the droplets is used in some counters. The most satisfactory method is individual counting of particles condensed in a continuous flow, as employed in the family of CNCs that have developed from the design of Bricard et al. (1976).  This procedure is employed in most modern commercial CNCs.  It has the advantage of giving a “direct“ determination of CN concentrations up to about 105 cm-3 (although this limit is not generally reached by simpler, lower-cost counters).

Because of the complexity of the particle number size distribution measurement, it is recommended that number size distribution determinations are undertaken only at GAW stations with more highly-developed aerosol programmes. Many commercial instruments are available to do this. They utilize a wide range of physical principles to classify particles according to size.  Some of the better known approaches utilize the electrical mobility of particles, aerodynamic size, or optical size determined by light scattering.  This latter class includes a number of relatively small, low-cost, instruments utilizing laser diodes. Whilst any of these measurement approaches has the potential to add useful data to the overall GAW aerosol measurement programme, it is assumed that stations operating such instrumentation will have a highly developed aerosol programme that includes a documented, calibration and quality control programme.

16.2.5.4. Cloud Condensation Nuclei

Cloud Condensation Nuclei (CCN) measurements are made to determine the concentration and establish climatologies of those particles that have the potential to produce cloud droplets at super saturations typical of natural clouds, that is, less than about 1%. Because of the complexity of the measurement, it is recommended that CCN determinations be undertaken at GAW stations with more highly-developed aerosol programmes. Past CCN measurements in the GAW Programme have been made predominantly using static thermal-gradient chambers, which are well-suited to relatively low frequency sampling and low-resolution (differential) CCN spectrum determination.  Instruments utilizing continuous flow offer another approach, and a commercial instrument using this approach is being deployed at a growing number of GAW stations. Measurement methods that have available droplet growth times comparable to real clouds are preferable.

# 16.2.5.5. Aerosol Optical Depth

Aerosol optical depth (AOD) is retrieved from observations of atmospheric spectral transmission. The solar spectral irradiance I at a given wavelength can be expressed as I=I0 exp(-mδ) with I0 the extraterrestrial (top of the atmosphere) irradiance of the sun, m the optical air mass and δ the total optical depth.  The optical air mass equals 1 for a vertical path through the atmosphere and is roughly proportional to 1/cos z with z the zenith angle of the sun during the observation.  The total optical depth δ at a given wavelength is composed of several components such as scattering by gas molecules δR (Rayleigh scattering), extinction by aerosol particles δA, absorption of trace gases δG (ozone, nitrogen dioxide, ..), and possible cloud contamination. Thus, the AOD can be obtained from the total optical depth by subtracting modeled estimates of the other components δA=δ-δR-δG.

Because AOD is essentially a difference between two larger numbers, it is sensitive to small calibration errors and to a minor degree also to the methods chosen to model the other components. A traceable calibration uncertainty of 1.5%, corresponding to an uncertainty of 0.015 optical depths at unit optical air mass, should be maintained for AOD observations (GAW Report No. 162).

Wavelengths and band passes specifically for AOD that are largely free of variable extinction components (water vapor and NOx) and strong ozone extinction have been recommended in (GAW Report No. 146). The BSRN and GAWPFR networks are using four AOD channels at 368, 412, 500 and 862 nm. While some other networks selected different wavelengths based on their specific needs (validation of satellite sensors, modeling efforts), measurements at 500±3 nm and 865±5 nm are typically available in most networks.

Measurements of the solar spectral irradiance are traditionally taken by sun-pointing radiometers (sun photometers) mounted on a 2-axis solar tracker with a sampling rate of once every minute in order to allow for objective quality control and cloud filtering algorithms. Homogeneous quality control is more difficult to achieve with hand-held sun photometers.

Rotating shadow band filter radiometers measure global and diffuse spectral irradiance at several wavelength bands. Direct normal irradiance obtained as difference between global and diffuse radiation, normalized by the solar zenith angle, can be used to retrieve AOD the same way as with sun photometers.

More advanced instruments like sky-scanning radiometers can be used to infer additional column aerosol optical properties, including size distribution, single-scattering albedo, or phase function, through sophisticated mathematical inversion models.

Sun photometers, shadow band, and sky-scanning radiometers are commercially available from several manufacturers. Centralized data evaluation and calibration services are offered for standardized instruments by global networks e.g. AERONET (AErosol RObotic NETwork), GAW-PFR, or SKYNET (GAW Report No. 162).

16.2.5.6. GAW Aerosol LIDAR

The basic lidar principle is the following: a laser pulse is transmitted into the atmosphere where it encounters gas molecules and particles; a small amount of this energy is backscattered in the direction of the receiver system, typically a telescope, and transferred to a photodetector as a photo-multiplier. The resulting electrical signal is proportional to the optical power received, which depends on the presence, range and concentration of atmospheric scatterers and absorbers in the light path volume.  Lidar techniques are able to characterize atmospheric aerosols in terms of vertical profiles of extinction and backscatter coefficients, lidar ratio, optical depth and microphysical properties such as shape, refractive index and size distribution, on time scales as short as minutes and vertical scales as short as meters.

Lidar observations are much more powerful if used in coordinated networks. Lidar networks are fundamental to studying aerosols on large spatial scale and to investigate transport and modification phenomena. There are several research lidar networks that contribute to GAW: Asian Dust Network (AD-Net), Latin American LIdar Network (LALINET or ALINE), Atmosphere aerosol and ozone monitoring in CIS regions through lidar stations network (CIS-LiNet),  European Aerosol Research Lidar Network (EARLINET), Micro-Pulse Lidar Network (MPLNET), Network for the Detection of Atmospheric Composition Change (NDACC), the NOAA Cooperative Remote Sencing Science and Technology (CREST) Lidar network (CLN). These networks are coordinated within GALION, the GAW Aerosol Lidar Observation Network (GAW Report No. 178).

Several different lidar techniques exist, depending on the specific instrument design and mainly on the specific laser-atmosphere scattering process.

- Elastic backscatter lidar

This is the simplest type of aerosol lidar: the backscattered wavelength is identical to the transmitted wavelength, and the magnitude of the received signal at a given range depends on the backscatter cross section of scatterers along the path to that range. Typical operating wavelengths are 355, 532, 1064 nm. The typical product of a backscatter lidar is the vertical profile of the aerosol backscatter coefficient obtained assuming a lidar ratio, i.e. the extinction-to-backscatter ratio, mostly constant throughout the profile and usually derived from existing climatology obtained with measurements from Raman lidar, described later. In this sense it is necessary to underline that without an a priori assumption about the lidar ratio, this kind of lidar systems can not provide quantitative aerosol backscatter data.

- Depolarization lidars

These are elastic backscatter lidars equipped with channels for the detection of the two parallel and cross-polarized components of the backscattered radiation. This provides quantitative information about particle shape, strongly contributing to aerosol typing as well as to the identification of thin clouds contaminating the profiles. A depolarization channel allows discrimination of volcanic ash and other aerosol particles.  Typical operating wavelengths are 355 and 532 nm. Depolarization lidar systems need accurate calibration.

- Raman lidar

The Raman lidar technique operates by measuring the inelastic Raman scattering by a specific gas. The Raman backscattered radiation from molecular nitrogen (or oxygen) is typically used for retrieving the vertical profile of aerosol extinction coefficient that, coupled with the elastic scattering collected at the same emission wavelength provides also the vertical profile of the aerosol backscatter coefficient without assuming a lidar ratio. Typical operating wavelengths are 355 and 532 nm. Most of the existing Raman lidar instruments are also equipped with a depolarization channel providing data on particle linear depolarization ratio. Advanced multi-wavelength Raman aerosol lidar techniques have been demonstrated to be the only technique capable of providing range-resolved aerosol microphysical properties. Moreover, rotational Raman lidar systems can be designed for optimizing extinction measurements in daytime conditions.

- High Spectral Resolution Lidar (HSRL)

The HSRL technique provides calibrated measurements of aerosol optical depth, extinction and backscatter. Measurements are computed from ratios of the particulate scattering to the measured molecular scattering. This provides absolute calibration and makes the calibration insensitive to dirt or precipitation on the output window. A very narrow, angular field-of-view reduces contamination from spurious sources, like multiple scattering contributions. The small field-of-view, coupled with a narrow optical bandwidth, nearly eliminates noise due to scattered sunlight improving the single-to-noise ratio also during daytime operations.

- Ceilometers

Ceilometers are basically elastic backscatter lidars that employ a diode laser source emitting at infrared wavelengths (typically 905 or 1064 nm) using a low energy but a high repetition rate (order of μJ of energy per pulse and kHz of rate) and detect the elastic backscattered radiation by clouds and precipitation. Ceilometers are a self-contained, 'turn-key', ground-based, active, remote-sensing device designed to measure cloud-base height and potentially the backscatter signals by aerosols. Ceilometers can provide qualitative information about aerosol vertical distribution.  Generally, older and typically less powerful instruments are barely able to detect aerosol layers in the atmosphere while newer instruments are quite useful for volcanic ash/dust detection and ash/dust plume tracking.

All these lidar techniques can provide data products suitable for monitoring the spatial and temporal distribution of aerosol up to the upper troposphere / lower stratosphere region and can characterize them from a dynamical and microphysical point of view. The main lidar limitation is related to the presence of rain, dense fog and thick clouds (optical depths larger than 2-3) that do not allow monitoring of the atmosphere above the cloud base region. The altitude range covered by lidars is limited at the bottom by the overlap height (altitude where there is a full overlap between the transmitter and the receiver) that it is typically about 250-500 meters above the ground level but could be also up to 2 km above the ground depending on the specific design. The maximum altitude range strongly depends on the laser power and optical design, but can be up to 25-30 km for high-power systems.  It is difficult to provide general estimates of the accuracy of the different lidar products because these strongly depend on the specific system and also on the meteorological conditions. On average, uncertainties for extinction and backscatter coefficient are about 20% (in case of Raman lidar or HSRL). Retrieval of microphysical properties is possible only if optical data have uncertainties lower than 20-30%.

Aerosol Lidar products (see Table 4 for more details):

- Geometrical properties

- layer identification (top, bottom and centre of mass)

- Optical properties profiles:

- Extensive optical parameters: aerosol backscatter coefficient (a), aerosol extinction coefficient (a)

- Intensive optical parameters: Lidar ratio (S), particle linear depolarization ratio (a), Ångström backscatter related exponent (å, Ångström extinction related exponent (å

- Optical properties in the identified layer

- Integrated backscatter, AOD

- Mean intensive optical parameters (Lidar ratio, particle linear depolarization ratio, Ångström backscatter related exponent, Ångström extinction related exponent)

- Aerosol typing classification

- Mass concentration estimate

- Microphysical properties retrieved

# Table 4 Lidar products related to specific ground-based lidar techniques

| Ground- based lidar techniques | Geometrical properties | βa | αa | S* | AOD | åβ | åα | type** | Microphysical properties |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ceilometer*** | ( | (x |  |  |  |  |  |  |  |
| Ceilo+sun photo. | ( | ( | ((d)1 |  | ((d) |  |  |  |  |
| Ceilo+sun photo.+ depol | ( | ( | ((d)1 |  | ((d) |  |  | ( (limited) |  |
| Backscatter lidar 1-λ | ( | ( |  |  |  |  |  |  |  |
| Backscatter lidar 1-λ + sun photo. | ( | ( | ((d)1 |  | ((d) |  |  |  |  |
| Backscatter lidar 1-λ + sun photo.+ depol | ( | ( | ((d)1 |  | ((d) |  |  | ((d) (limited) |  |
| Backscatter lidar m-λ# | ( | ( |  |  |  |  | ( |  |  |
| Backscatter lidar m-λ# + sun photo. | ( | ( | ((d)1 |  | ((d) | ((d)1 | ( |  | ((d)1 |
| Backscatter lidar m-λ# + sun photo.+ depol | ( | ( | ((d)1 |  | ((d) | ((d)1 | ( | ( | ((d)1 |
| 1-λ Raman lidar/HSRL | ( | ( | (2 | (2 | (2 |  |  | (2 (limited) |  |
| 1-λ Raman lidar/HSRL +sun photo. | ( | ( | (2 | (2 | (2 | ((d)1 | ((d)1 | ( (limited) | ((d)1 |
| 1-λ Raman lidar/HSRL +sun photo. + depol | ( | ( | (2 | (2 | (2 | ((d)1 | ((d)1 | ( | ((d)1 |
| m-λ# Raman lidar | ( | ( | (2 | (2 | (2 | (2 | ( | (2 | (2 |
| m-λ# Raman lidar +sun photo. | ( | ( | (2 | (2 | (2 | (2 | ( | ( | (2 |
| m-λ# Raman lidar +sun photo.+ depol | ( | ( | (2 | (2 | (2 | (2 | ( | ( | (2 |

( From two independent measurements

** Identification of scattering type (aerosol particles, cloud droplets, ice crystals, some aerosol type information)

*** A ceilometer is a single-wavelength, low-power lidar, with lower S/N ratio

x if calibrated

# m > 2

Estimate only

2 Most Raman lidar systems operate during night-time. 24h Raman lidar systems exist and their operability has been proved, however, few systems nowadays operate Raman channels also during daytime; HSRL is independent of daytime.

(d)= only daytime, (n)= only night-time

# 16.2.6. Natural Radioactivity

The global distributions of the source/sink terms of the naturally occurring radionuclides (7Be, 10Be, 210Pb, and 222Rn) and the anthropogenic radionuclides (85Kr) are reasonably well known.  7Be and 10Be are produced by cosmic-ray interactions in the upper troposphere and the lower stratosphere.  222Rn is exhaled from the Earth’s land surface as a result of uranium decay in soil.  210Pb is produced in the lower troposphere from the decay of 222Rn.  Most of the 85Kr in the atmosphere is released during nuclear fuel reprocessing. Atoms of 7Be, 10Be and 210Pb attach themselves to submicron-size aerosol particles, and therefore, act as aerosol-borne tracers in the atmosphere.  222Rn and 85Kr, chemically and physically inert, act as noble gases in the atmosphere.

Measurements of radionuclides are not a priority area within the GAW Programme. Some general recommendations can be found in GAW Reports No. 143 and 155.

# Referenced GAW REPORTS:

- Global Atmosphere Watch Measurements Guide (WMO TD No. 1073), 2001.

146. 	Quality Assurance in monitoring solar ultraviolet radiation: the state of the art. (WMO TD No. 1180), 2003.

153. 	WMO/GAW Aerosol Measurement Procedures: Guidelines and Recommendations. (WMO TD No.1178), 2003.

155. 	1st International Expert Meeting on Sources and Measurements of Natural Radionuclides Applied to Climate and Air Quality Studies (Gif sur Yvette, France, 3-5 June 2003) (WMO TD No. 1201), 2003.

160. 	Manual for the GAW Precipitation Chemistry Programme (Guidelines, Data Quality Objectives and Standard Operating Procedures) (WMO TD No. 1251), 186 pp, November 2004.

162.	WMO/GAW Experts Workshop on a Global Surface-Based Network for Long Term Observations of Column Aerosol Optical Properties, Davos, Switzerland, 8-10 March 2004 (edited by U. Baltensperger, L. Barrie and C. Wehrli) (WMO TD No.  1287), 153 pp, November 2005.

171.  A WMO/GAW Expert Workshop on Global Long-Term Measurements of Volatile Organic Compounds (VOCs) (Geneva, Switzerland, 30 January – 1 February 2006) (WMO TD No. 1373), 36 pp, February 2007.

172. WMO Global Atmosphere Watch (GAW) Strategic Plan: 2008 – 2015 (WMO TD No. 1384), 108 pp, August 2008.

176. The Tenth Biennial WMO Consultation on Brewer Ozone and UV Spectrophotometer Operation, Calibration and Data Reporting (Northwich, United Kingdom, 4-8 June 2007) (WMO TD No. 1420), 61 pp, March 2008.

178.	Plan for the implementation of the GAW Aerosol Lidar Observation Network GALION, (Hamburg, Germany, 27 - 29 March 2007) (WMO TD No. 1443), 52 pp, November 2008.

183. Operations Handbook – Ozone Observations with a Dobson Spectrophotometer (WMO TD No. 1469), 91 pp, March 2009

185. Guidelines for the Measurement of Methane and Nitrous Oxide and their Quality Assurance (WMO TD No. 1478), 49 pp, September 2009.

192. Guidelines for the Measurement of Atmospheric Carbon Monoxide (WMO TD No. 1551), 49 pp, July 2010.

195. WMO/GAW Expert Workshop on Global Long-term Measurements of Nitrogen Oxides and Recommendations for GAW Nitrogen Oxides Network (Hohenpeissenberg, Germany, 8-9 October 2009) (WMO TD No. 1570), 45 pp, February 2011.

197. Addendum for the Period 2012 – 2015 to the WMO Global Atmosphere Watch (GAW) Strategic Plan 2008 – 2015, 57 pp, May 2011.

200. WMO/GAW Standard Operating Procedures for In-Situ Measurements of Aerosol Mass Concentration, Light Scattering and Light Absorption (Edited by John A. Ogren), 134 pp, October 2011

201.	Quality Assurance and Quality Control for Ozonesonde Measurements in GAW (Prepared by Herman Smit and ASOPOS Panel), 95 pp. January 2013

204.  	Standard Operating Procedures (SOPs) for Air Sampling in Stainless Steel Canisters for Non-Methane Hydrocarbons Analysis (Prepared by Rainer Steinbrecher and Elisabeth Weiß), 25 pp. September 2012

206.  	16th WMO/IAEA Meeting of Experts on Carbon Dioxide, Other Greenhouse Gases and Related Tracers Measurement Techniques (GGMT-2011), Wellington, New Zealand, 25-28 October 2011, 67 pp, October 2012.

209. 	Guidelines for Continuous Measurements of Ozone in the Troposphere (WMO-No. 1110, ISBN: 978-92-63-11110-4), 76 pp, February 2013

# Additional references

Birch, M.E. and Cary, R.A., Elemental carbon-based method for monitoring occupational exposure to particulate diesel exhaust. Aerosol Sci. Technol., 25, 221-241, 1996.

Bricard, J. Delattre, P., Madelaine, G., Pourprix, M., Detection of ultra-fine particles by means of a continuous flux condensation nuclei counter. In Fine Particles, Liu, B.Y.H.m, Editor., Academic Press, New York, p. 566-580,1976.

Cavalli, F., Viana, M., Yttri, K. E., Genberg, J., and Putaud, J.-P., Toward a standardised thermal-optical protocol for measuring atmospheric organic and elemental carbon: the EUSAAR protocol, Atmos. Meas. Tech., 3, 79-89, doi:10.5194/amt-3-79-2010, 2010.

Chow, J.C., Watson J.G., Chen, L.-W.A., Paredes-Miranda, G., Chang M.-C.O., Trimble, D., Fung, K.K., Zhang, H. and Yu, J.Z., Refining temperature measures in thermal/optical carbon analysis, Atmos. Chem. Phys., 5, 2961-2972, 2005.

Chow, J.C., Watson, J.G., Crow, D., Lowenthal, D.H., and Merrifield, T., Comparison of Improve and NIOSH carbon measurements, Aerosol. Sci. Technol., 34, 23-34, 2001.

Chow, J.C., Watson, J.G., Prithett, L.C., Pierson, W.R., Frazier, C.A. and Purcell, R.G., The DRI Thermal/Optical Reflectance carbon analysis system: Description, evaluation, and applications in U.S. air quality studies, Atmos. Environ., 27A, 1185-1201, 1993.

Grant, A., Witham, C. S., Simmonds, P. G., Manning, A. J., and O’Doherty, S.: A 15 year record of high-frequency, in situ measurements of hydrogen at Mace Head, Ireland, Atmos. Chem. Phys., 10, 1203–1214, doi:10.5194/acp-10-1203-2010, 2010 (http://www.atmos-chem-phys.net/10/1203/2010/acp-10-1203-2010.pdf).

Heath, D.F., A.J. Krueger and H. Park, The solar backscatter ultraviolet (SBUV) and total ozone mapping spectrometer (TOMS) experiment. In: C.R. Madrid, The NIMBUS-7 User’s Guide. NASA Goddard Space Flight Center, Greenbelt, Maryland, pp. 175–211, 1978.

Herndon, S.C., Zahniser, M.S., Nelson Jr., D.D., Shorter, J., McManus, J.B., Jiménez, R., Warneke C., de Gouw, J.A., Airborne measurements of HCHO and HCOOH during the New England Air Quality Study 2004 using a pulsed quantum cascade laser spectrometer. Journal of Geophysical Research, 112, D10S03, doi:10.1029/2006JD007600, 2007.

## Jaroslawski, J., Improvement of the Umkehr ozone profile by the neural network method: analysis of the Belsk (51.80°N, 20.80°E) Umkehr data, International Journal of Remote Sensing, Volume 34, Issue 15, 2013, 5541-5550.

NIOSH (1999) Elemental Carbon (Diesel Particulate): Method 5040. www.cdc.gov./niosh/nmam/pdfs/5040f3.pdf.

Petropavlovskikh,I., Bhartia, P.K., De Luisi, J., New Umkehr ozone profile retrieval algorithm optimized for climatological studies, Geophys. Res. Lett., 32, L16808, doi:10.1029/2005GL023323, 2005.

Petzold, A., Ogren, J. A., Fiebig, M., Laj, P., Li, S.-M., Baltensperger, U., Holzer-Popp, T., Kinne, S., Pappalardo, G., Sugimoto, N., Wehrli, C., Wiedensohler, A., and Zhang, X.-Y., Recommendations for reporting "black carbon" measurements, Atmos. Chem. Phys., 13, 8365-8379, doi:10.5194/acpd-13-8365-2013, 2013.

Schmid, H., Laskus, L., Abraham, H.J., Baltensperger, U., Levenchy, V., Bizjak, M., Burba, P., Cachier, H., Crow, D., Chow, J., Gnauk, V., Even, V., ten Brink, H.M., Giesen, K.P., Hitzenberger, R., Hueglin, C., Maenhaut, W., Pio, C., Carvalho, V., Putaud, J.P., Toom-Sauntry, D. and Puxbaum, H., Results of the "carbon conference" international carbon round robin test stage I, Atmos. Environ., v. 35, p. 2111–2121, 2001.

Sofieva, V. F., N. Rahpoe, J. Tamminen, E. Kyrölä, N. Kalakoski, M. Weber, A. Laeng, T. von Clarmann, G. Stiller, S. Lossow, D. Degenstein, A. Bourassa, C.  Adams, C. Roth, N. Lloyd, P. Bernath, R. J. Hargreaves, J. Urban, D. Murtagh, A. Hauchecorne, M. Van Roozendael, N. Kalb,  and C. Zehner, Harmonized dataset of ozone profiles from satellite limb and occultation measurements, Earth Syst. Sci. Data Discuss., 6, 189-222, doi:10.5194/essdd-6-189-2013, 2013.

Tegtmeier, S., Hegglin, M. I., Anderson, J., Bourassa, A., Brohede, S., Degenstein, D., Froidevaux, L., Fuller, R., Funke, B., Gille, J., Jones, A., Kasai, Y., Kyrölä, E., Lingenfelser, G., Lumpe, J., Nardi, B., Neu, J., Pendlebury, D., Remsberg, E., Rozanov, A., Smith, L., Toohey, M., Urban, J., von Clarmann, T., Walker, K. A., and Wang, R., A comparison of ozone climatologies from international limb satellite sounders, J. Geophys. Res., submitted, 2013.

Tuzson, B.,  Zeyer, K., Steinbacher, M.,   McManus, J. B., Nelson, D. D., Zahniser, M. S.,  and L. Emmenegger, Selective measurements of NO, NO2 and NOy in the free troposphere using quantum cascade laser spectroscopy,  Atmos. Meas. Tech., 6, 927-936, 2013.

van der A, R. J.,  Allaart, M. A. F. and Eskes, H. J., Multi sensor reanalysis of total ozone, Atmos. Chem. Phys., 10, 11277-11294, 2010.

Warneke, C., Veres, P., Holloway, J. S., Stutz, J., Tsai, C., Alvarez, S., Rappenglueck, B., Fehsenfeld, F. C., Graus, M., Gilman, J. B., de Gouw, J. A., Airborne formaldehyde measurements using PTR-MS: calibration, humidity dependence, inter-comparison and initial results. Atmos. Meas. Tech., 4, 2345-2358, 2011.

Watson, J.G., Chow, J.C. , Chen, L.-W. A., Frank, N.H., Methods to Assess Carbonaceous Aerosol Sampling Artifacts for IMPROVE and Other Long-Term Networks, Journal of the Air & Waste Management Association, 59:8, 898-911, doi: 10.3155/1047-3289.59.8.898, 2009.

Wisthaler, A, Apel, E.C., Bossmeyer, J., Hansel, A., Junkermann, W., Koppmann, R., Meier, R., Müller, K., Solomon, S.J., Steinbrecher, R., Tillmann, R., Brauers, T., Intercomparison of formaldehyde measurements at the atmosphere simulation chamber SAPHIR. Atmos. Chem. Phys., 8, 2189–2200, 2008.

# ANNEX 16.1

List of GAW Central Facilities and host institutions as of September 2013. The World Central Facilities have assumed global responsibilities, unless indicated (Am: Americas; E/A: Europe and Africa; A/O: Asia and the South-West Pacific).

| Variable | QA/SAC | Central Calibration Laboratory (CCL) (Host of Primary Standard) | World Calibration Centre (WCC) | Regional Calibration Centre (RCC) | World Data Centre (WDC) |
| --- | --- | --- | --- | --- | --- |
| CO2 | JMA (A/O) | NOAA-ESRL | NOAA-ESRL (round robin) Empa (audits) |  | JMA |
| carbon isotopes |  | MPI-BGC |  |  | JMA |
| CH4 | Empa (Am, E/A) JMA (A/O) | NOAA-ESRL | Empa (Am, E/A) JMA (A/O) |  | JMA |
| N2O | UBA | NOAA-ESRL | IMK-IFU |  | JMA |
| CFCs, HCFCs, HFCs |  |  |  |  | JMA |
| SF6 |  | NOAA-ESRL | KMA |  | JMA |
| H2 |  | MPI-BGC |  |  | JMA |
| Total Ozone | JMA (A/O) | NOAA-ESRL1, EC2 | NOAA-ESRL1, EC2 | BoM1, NOAA-ESRL1, IZO2 ,JMA1, MOHp1, MGO3, OCBA1, SAWS1, SOO-HK1 | EC5, DLR6 |
| Ozone Sondes | IEK-8 | IEK-8 | IEK-8 |  | EC |
| Surface Ozone | Empa | NIST | Empa | OCBA | JMA |
| Precipitation Chemistry | NOAA-ARL | ISWS | ISWS |  | NOAA-ARL |
| CO | Empa | NOAA-ESRL | Empa |  | JMA |
| VOC | UBA | NPL, NIST | IMK-IFU |  | JMA |
| SO2 |  |  |  |  | JMA |
| NOx | UBA |  | IEK-8 (NOx)* |  | JMA |
| Aerosol | UBA (physical properties) |  | IfT (physical properties) |  | NILU5, DLR6 |
| Optical Depth |  | PMOD/WRC4 | PMOD/WRC |  | NILU |
| UV Radiation |  |  | PMOD/WRC | NOAA-ESRL (Am), EUVC/PMOD (E) | EC |
| Solar Radiation |  | PMOD/WRC | PMOD/WRC |  | MGO |

1Dobson, 2Brewer, 3Filter instruments, 4Precision Filter Radiometers (PFR), 5ground-based, 6satellite-based

* proposal is currently under evaluation

Host Institutions

| BoM | Bureau of Meteorology, Melbourne, Australia |
| --- | --- |
| BSRN | Baseline Surface Radiation Network, Federal Institute of Technology (ETH), Zürich, Switzerland |
| DLR | German Aerospace Centre, Oberpfaffenhofen, Wessling, Germany |
| EC | Environment Canada, Toronto, Canada |
| Empa | Swiss Federal Laboratories for Materials Testing and Research, Dübendorf, Switzerland |
| EUVC | European Ultraviolet Calibration Center, cf. PMOD/WRC |
| IEK-8 | Institute for Energy and Climate Research: Troposphere (IEK-8), Research Center Jülich GmbH, Jülich, Germany |
| IfT | Institute for Tropospheric Research, Leipzig, Germany |
| IMK-IFU | Karlsruhe Institute of Technology (KIT), Institute for Meteorology and Climate Research, IMK-IFU, Garmisch-Partenkirchen, Germany |
| ISWS | Illinois State Water Survey, Champaign IL, USA |
| IZO | Izaña Observatory, Tenerife, Spain |
| JMA | Japan Meteorological Agency, Tokyo, Japan |
| KMA | Korea Meteorological Administration, Seoul, the Republic of Korea |
| MGO | A.I. Voeikov Main Geophysical Observatory, Russian Federal Service for Hydrometeorology and Environmental, St. Petersburg, Russia |
| MPI-BGC | Max-Planck Institute (MPI) for Biogeochemistry, Jena, Germany |
| MOHp | Meteorologisches Observatorium Hohenpeissenberg, Hohenpeissenberg, Germany |
| NOAA-ARL | U.S. Department of Commerce, National Oceanic and Atmospheric Administration, Air Resources Laboratory, Silver Spring MD, USA |
| NOAA-ESRL | U.S. Department of Commerce, National Oceanic and Atmospheric Administration, Earth System Research Laboratory, Global Monitoring Division, Boulder, CO, USA |
| NILU | Norwegian Institute for Air Research, Kjeller, Norway |
| NIST | National Institute for Standards and Technology, Gaithersburg MD, USA |
| NPL | National Physical Laboratory, Teddington, Middlesex, United Kingdom |
| OCBA | Observatorio Central Buenos Aires, Argentina |
| PMOD/WRC | Physikalisch-Meteorologisches Observatorium Davos/World Radiation Centre, Davos, Switzerland |
| SAWS | South African Weather Service, Pretoria, South Africa |
| SOO-HK | Solar and Ozone Observatory, Hradec Kralove, Czech Republic |
| UBA | German Environmental Protection Agency, Berlin, Germany |

# Footnotes

CO2, for example, mostly consists of 12C16O16O, while the smaller abundance higher-mass isotopologues from mass 45 up to mass 49 fractions of (13C16O16O, 14C16O16O, or 12C18O16O, the corresponding 17O siblings and the mixed-isotope species)  are also found present in the atmosphere.

This is for example common practice in gas-chromatography measurements

# Comments

Chapter on sat observations