Appendices: I

Draft: as of 7 NOV 2017

Tentative ILC protocol

INSTRUCTION FOR THE PARTICIPANTS IN THE INTERLABORATORY COMAPRISON

Title: Intercomparison in the field of temperature, humidity and pressure

Date of approval of the protocol:

XX.XX.2017

Items:

- Two Pt-100 resistance thermometers ELPRO type 2210 4700/X in combination with

Keysight/Agilent/Hewlett Packard 34420A

- Capacitive hygrometer Vaisala HMP155 A2GB11A0A1A1A0A

- Barometer Vaisala PTB220 ACA2A3A1AB

TABLE OF CONTENT

1 Introduction	3

1.1 Coordinator/Reference laboratory for pressure, temperature and humidity	3

1.2 Data analysis coordinator	3

1.3 Participants	3

1.4 Time schedule and deadlines	5

1.5 Transportation of the equipment	5

2 Description of the equipment	7

2.1 General	7

2.2 Environmental conditions	8

2.3 Handling	8

2.3.1 Packing and unpacking	8

2.3.2 Mounting	8

2.3.3 Precautions	8

3 Calibration/Test method	9

3.1 Start-up and initial inspection	9

3.2 Measuring points	9

3.3 Reporting of results	10

3.4 Measurement uncertainties	11

4 Appendix A: Report form	12

5 Appendix B: User's Guide for Keysight/Agilent/Hewlett Packard 34420A	15

6 Appendix D: User's Guide for Vaisala HMP155	17

7 Appendix C: User's Guide for Vaisala PTB220	18

8 Appendix D: Packaging instructions	20

9 Appendix E: Instrument Check List	21

# 1 Introduction

Interlaboratory comparison (ILC) serves as a tool for comparison of measurement results carried out by accredited or non-accredited calibration laboratories in the relevant field of measurement. ILC represents very effective means to demonstrate technical competence of the participant and also serves as a technical base for accreditation. Furthermore, it is the most important element for monitoring of quality of measurement results as required by ISO/IEC 17025:2005 standard for laboratories in part 5.9.

This ILC will be organized by the Regional Instrument Centre Tsukuba (Japan Meteorological Agency) in cooperation with University of Ljubljana and Slovenian Environmental Agency based on ILC protocol prepared within MeteoMet project.

It is recommended that the participants use their standard procedure during the temperature, humidity and pressure calibration and if possible avoid making extra time-consuming measurements. The proficiency test is carried out in accordance with ISO/IEC 17043:2010.

## 1.1 Coordinator/Reference laboratory for pressure, temperature and humidity

The reference value will be determined by RIC Tsukuba (Japan Meteorological Agency).

RIC Tsukuba (Japan Meteorological Agency)

1-2 Nagamine Tsukuba Ibaraki 305-0052, Japan

Kouichi NAKASHMA (Mr)

Phone: +81 298 51 4123, Fax: +81 298 51 1670

Email: kouichi.nakashimamet.kishou.go.jp

## 1.2 Data analysis coordinator

Contact persons in case of technical questiopns:

RIC Tsukuba (Japan Meteorological Agency)

1-2 Nagamine Tsukuba Ibaraki 305-0052, Japan

Kouichi NAKASHIMA (Mr)

Phone: +81 298 51 4123, Fax: +81 298 51 1670

Email: kouichi.nakashimamet.kishou.go.jp

## 1.3 Participants

Following are the participants to this intercomparison. Contact details are as listed alphabetically:

Participating laboratory: Slovenian Environmental Agency (Slovenia) (RA-VI)

Contact person: Mr. Drago Groselj

Address: Vojkova 1b

Zip: 1000

City: Ljubljana

Country: Slovenia

Phone: +386 1 478 4100

Fax: +386 1 478 4054

GSM: +386 31 655 216

Email: drago.groseljgov.si

Participating laboratory: University of Ljubljana, Faculty of Electrical Engineering

(Slovenia) (only temperature and humidity) (RA-VI)

Contact person: Mr. Gaber BEGES

Address: Tržaška cesta 25

Zip: 1000

City: Ljubljana

Country: Slovenia

Phone: +386 1 4768 224

Fax: +386 1 4264 633

Email: gaber.begesfe.uni-lj.si or infolmk.fe.uni-lj.si

Participating laboratory: RIC Beijing (RA-II)

Contact person: XXXX

Address: XXXX

Zip: XXXX

City: XXXX

Country: XXXX

Phone: XXXX

Fax: XXXX

Email: XXXX

Participating laboratory: RIC Tsukuba (Japan Meteorological Agency, RA-II)

Contact person: Kouichi NAKASHIMA

Address:1-2 Nagamine Tsukuba Ibaraki 305-0052, Japan

Zip: 305-0052

City: Tsukuba

Country: Japan

Phone: +81 298 51 4123

Fax: +81 298 51 1670

Email: kouichi.nakashimamet.kishou.go.jp

Participating laboratory: RIC Manila (RA-V)

Contact person: Ferdinand BARCENAS

Address: PAGASA, Science Garden, Agham Road, Diliman

Zip: 1100

City: Quezon City

Country: Philippines

Phone: (632) 9292121

Fax: (632) 9292121

Email: fybarce8yahoo.com

Participating laboratory: RIC Melbourne (RA-V)

Contact person: XXXX

Address: XXXX

Zip: XXXX

City: XXXX

Country: XXXX

Phone: XXXX

Fax: XXXX

Email: XXXX

## 1.4 Time schedule and deadlines

All participating laboratories have one and half months for calibration including transport to the next laboratory.

If a participant anticipates difficulties in keeping the deadlines, the coordinator must be contacted immediately. In such a case the other participants will be contacted as soon as possible and be informed about eventual changes.

Deadline for reporting the results is 4 weeks after the equipment has left the laboratory. It is important that the deadline is met since the results are being analyzed continuously by the reference laboratory. If there are any problems or doubt regarding the results of the participant laboratory, the laboratory will be contacted immediately. Any suspicion that the equipment is defective or drifted, will lead to return of the equipment to the reference laboratory, which then will make an extra check and take an appropriate action.

If deadlines are respected, the participants will receive the report of the intercomparison before XX.XX.2018.

The first measurement in the first laboratory in schedule is expected to start on XX.XX.2017.

## 1.5 Transportation of the equipment

As soon as the equipment is delivered/sent, the coordinator shall be informed (e.g. by e-mail). The equipment is then unpacked, and an inspection carried out. If the equipment has any visible damage due to transportation, this must be reported to the coordinator before the calibration begins.

Costs related to transport of the equipment from Europe to Japan and back to Europe and insurance costs of the equipment is covered by RIC Tsukuba.

Equipment will be accompanied with ATA CARNET forms for non-EU countries. In case that your country is not signatory of ATA CARNET convention, please perform temporary import/export procedure.

The participants in RA II and V should bear the costs of customs and transportation inside the country, and pay for shipping the instruments to the next destination (airport).

The participants must ensure that the equipment is covered by insurance within the responsibility of its transportation.

(This scheme referred to the Methane Reference Gas Intercomparison by WMO/GAW World Calibration Centre (WCC) for Methane in Asia and the South-West Pacific.)

(1) RIC Tsukuba sends the equipment to the Melbourne Airport.

(2) RIC Melbourne receives the equipment at the Melbourne Airport and passes its customs. After calibration, RIC Melbourne sends the equipment to the Ninoy Aquino International Airport.

(3) RIC Manila receives the equipment at the Ninoy Aquino International Airport and passes its customs. After calibration, RIC Manila sends the equipment to the Beijing Capital International Airport.

(4) RIC Beijing receives the equipment at the Beijing Capital International Airport and passes its customs. After calibration, RIC Beijing sends the equipment to the Narita International Airport.

# 2 Description of the equipment

## 2.1 General

The instrument’s owner: UL/FE-LMK and ARSO.

For transportation purposes the measuring instruments will be placed in a protecting case.

In a case any of the above-mentioned equipment is missing at the receipt, the coordinator must be contacted.

## 2.2 Environmental conditions

Calibration is carried out at an ambient temperature of (20 ± 5) °C and relative humidity of (30 ÷ 60) %. The ambient conditions shall be reported.

Barometer stabilization (warm-up) before measurements should have been made according to the common laboratory practice, but at least for 2 hours in the laboratory environment.

## 2.3 Handling

### 2.3.1 Packing and unpacking

Procedure for unpacking is as follows:

1. Inspect the transportation boxes for damage. If the boxes are damaged, the coordinator shall be contacted before continuing.

2. Unpack the equipment and check that all equipment mentioned in the section “Description of equipment” is present. Please, fill up “Instrument Check List” – Appendix E and send it to kouichi.nakashima@met.kishou.go.jp

3. If any equipment is missing, the coordinator shall be contacted.

4. Inspect the equipment. If any of the equipment shows visible signs of damage, the coordinator shall be contacted.

The packing procedure is as follows:

1. Before packing, slowly cool down the equipment (thermometers) to room temperature and clean them with pure alcohol. Clean also other calibrated items if necessary.

2. Check that all equipment mentioned in the section “Description of equipment” is packed before the equipment is transported to the next participant. For details of packing, please see Appendix D: Packaging instructions.

### 2.3.2 Mounting

1. The Pt100 are cleaned before use with pure alcohol. Other instruments should be cleaned only if necessary.

2. Thermometers are carefully placed in the calibration media (bath, climatic chamber). Relative humidity sensor is carefully placed in the calibration media (climatic chamber or humidity generator). Barometer is carefully connected to the barometric pressure system.

### 2.3.3 Precautions

- Pt100 are very sensitive device to vibration and mechanical shock.

- When not in use, it should be stored in a safe place in the provided transport boxes.

- Check that the thermometers are completely clean and dry before placing them in the calibration media.

- Ensure that the thermometers are cooled down and cleaned with pure alcohol before placing them in the transportation box.

- Don’t expose relative humidity sensor to temperatures higher than 25 °C and lower than 15 °C.

- Don’t expose barometer to pressures other than barometric!

Contact the coordinator in a case of doubt about the above-mentioned precautions.

# 3 Calibration/Test method

It is recommended that the participants use their standard procedure during temperature, humidity and pressure calibration and avoid making extra time-consuming measurements, if possible. For accredited laboratories it will be advantageous to apply the accredited procedures in preparation for later use of the report in relation with documentation to the accreditation body.

Details about the applied procedure have to be stated in the report form.

## 3.1 Start-up and initial inspection

The Keysight/Agilent/Hewlett Packard 34420A instrument can measure 1 Pt 100 at a time. Connect first Pt100 to the Keysight/Agilent/Hewlett Packard 34420A instrument as described on the page 21 of Appendix B. Before start of the measurement, you can setup automatic acquisition of temperature measurement for Keysight/Agilent/Hewlett Packard 34420A instrument using GPIB communication bus. Needed commands are available in the User's Guide (see pages 71 and 72).

BE CAREFUL NOT TO CHANGE ANY OF PRESET COEFFICIENTS!

After setting up the equipment for the calibration, take the first measurements as follows:

1. For temperature, after the stabilization take the first measurements in the triple point of water, or the ice-point, or in the thermal bath at 0 °C. Note the readings of the both thermometers as well as laboratory’s reference thermometer and report all these values to the coordinator/reference laboratory, immediately.

2. For relative humidity, after the stabilization take the first measurements at relative humidity of about 55 %. Note the readings of the hygrometer as well as laboratory’s reference hygrometer and report all these values to coordinator/reference laboratory, immediately.

3. For atmospheric pressure, after the stabilization take the first measurements at atmospheric pressure of about 950 hPa. Note the readings of the barometer as well as laboratory’s reference barometer and report all these values to coordinator/reference laboratory, immediately.

Wait for the instruction from the coordinator/reference laboratory.

## 3.2 Measuring points

- Temperature

The subject of the ILC is the calibration of two Pt100 in combination with Keysight/Agilent/Hewlett Packard 34420A. The calibration shall be performed in the following measurement points within tolerances ±0.2 °C using standard laboratory procedures:

- Relative humidity:

The subject of the ILC is the calibration of the capacitive hygrometer Vaisala HMP155. The calibration shall be made in the following measurement points within tolerances ±3 %r.h. at temperature of 20°C using standard laboratory procedures:

In case of salt solution calibration other calibration points can be used covering as wide measuring range as possible.

- Air pressure:

The subject of the ILC is the calibration of the digital barometer Vaisala PTB220 ACA2A3A1AB. The calibration shall start at a minimum calibration point followed by increasing pressure and return steps by decrease of pressure. The calibration shall be made in the following measurement points within tolerances 20 hPa using standard laboratory procedures:

| Warning: | For the purposes of this ILC it is strictly forbidden to make any adjustments to the measuring instrument – test item, especially to activate the calibration regime! The only actions permitted are the movement among functions and operations described in the enclosed instruction manual. |
| --- | --- |

## 3.3 Reporting of results

The results are reported electronically in the forwarded Excel spreadsheet. The green fields of the spreadsheet should be filled in, if possible. Send your results only to the data analysis coordinator (kouichi.nakashima@met.kishou.go.jp).

The results shall be sent to the coordinator no later than 4 weeks after having finalized the calibration. Electronic reporting by e-mail is preferred.

As a final result of the ILC send to the coordinator is fulfilled Excel spreadsheet form and Calibration certificate for each quantity.

Outline of statistical analysis

The assigned values are to be determined as the arithmetic mean of measurements made by the coordinator. Any outliers are detected by Cochran’s test and Grubbs’ test (ISO 5725-2, 7.3.3 and 7.3.4).

Information to be returned to participants

Final results of the participants will be anonymized (each laboratory with different code). Coding system will be known to the data analysis coordinator only. The participants will receive summary of all measurements, assigned values and uncertainties of assigned values, and evaluation of the performance.

The evaluation of measurement results will be made on the basis of En number:

En=xlab-xrefUlab2+Uref2

where xlab is the participant’s result, xref is the assigned value, Ulab is the expanded (k=2) uncertainty of a participant’s result and Uref is the expanded (k=2) uncertainty of the assigned value.

The assigned value xref will be calculated as mean of reference laboratories. The uncertainty of the assigned value Uref will be calculated as uncertainty of mean, with uncertainties of reference laboratories at each calibration point.

Criteria for performance evaluation will be based on statistical determination for En number:

En≤1=satisfactory

En>1=unsatisfactory

Complaints and appeals

Participants may file a complaint or appeal to the data analysis coordinator in 60 days after they receive the final PT report.

## 3.4 Measurement uncertainties

The measurement results should be stated with their associated uncertainties, the evaluation of uncertainties should be done according to the document EA 4/02 M rev01 (September 2013), ILAC P14:01/2013 ILAC Policy for Uncertainty in Calibration. The interlaboratory comparison should be made according to the best laboratory measurement practice.

# 4 Appendix A: Report form

Standard forms for reporting of results and uncertainties for temperature, relative humidity and pressure. The form will be forwarded electronically.

# 5 Appendix B: User's Guide for Keysight/Agilent/Hewlett Packard 34420A

The original User's Guide for Keysight/Agilent/Hewlett Packard 34420A is appended to the ILC Protocol and provided in the ILC kit. Only one thermometer Pt100 can be measured by the Keysight/Agilent/Hewlett Packard 34420A at a time.

The ILC kit for temperature measurements consists of:

- Keysight/Agilent/Hewlett Packard 34420A

- Two identical thermometers Pt100

- Connection unit

- Connection cable (34420A – connection unit)

The Pt100 is connected to the connection unit as shown in the following picture.

CONNECTION UNIT

The thermometer is connected (4 wire) to the connection unit as follows:

The connection unit is connected to the front connection of 34420A. Measured values can be read on display or via GIPB interface bus. When thermometer is connected and 34420A switched on, the temperature of the thermometer can be observed by pressing SHIFT and TEMP key. The display shows temperature of Pt100 in degrees Celsius.

SHIFT KEY

TEMP KEY

If thermometers reading via GPIS is preferred, following setup for 34420A is used:

GPIB address XX

Command string is <read?> to read value.

| Warning: | For the purposes of this ILC it is strictly forbidden to make any adjustments to the measuring instrument. The only actions permitted are the movement among functions. |
| --- | --- |

# 6 Appendix D: User's Guide for Vaisala HMP155

The original User's Guide for Vaisala HMP155 is appended to the ILC Protocol and provided in the ILC kit. The Vaisala HMP155 type A2GB11A0A1A1A0A is subject for the ILC. The instrument has analog outputs (voltage) for relative humidity and air temperature:

- relative humidity: 0 – 1V corresponds 0 to 100 %r.h. Multiplying by 100 is needed to calculate relative humidity.

- air temperature: 0 – 1V corresponds -40 to +60°C. Multiplying by linear function y=100.x-40 is needed to calculate air temperature.

Power supply: although power supply in the range from 7 to 28 VDC may be used, please, use 12V ± 2V in order to minimise power supply impact.

Vaisala HMP155 connections and wiring diagram:

Note: Please, remove sensor protection cap before use and return it back after use.

# 7 Appendix C: User's Guide for Vaisala PTB220

The original User's Guide for Vaisala PTB220 is appended to the ILC Protocol and provided in the ILC kit. The Vaisala PTB220 type ACA2A3A1AE is subject for the ILC. The Vaisala PTB220 type ACA2A3A1AE has three pressure transducers (upper left – UL, upper right- UR, lower left – LL) and fourth value is average value (lower right - LR). The pressure connector is barbed fitting 1/8".

Please, don’t expose barometer to pressures other than barometric!

LOWER RIGHT -LR

UPPER RIGHT -UR

UPPER LEFT -UL

LOWER LEFT –LL

L

R

The barometers are supplied with RS232 communication cable combined with power supply. Please, use supplied power supply for the purpose of ILC.

Measured values can be read on display or via RS232C. When barometer is switched on, the pressure can be observed on display. The display shows pressure in hecto Pascal.

If reading via RS232C is preferred, following setup for barometers is used:

| Baud rate | 9600 |
| --- | --- |
| Data bits | 7 |
| Parity | Even |
| Stop bits | 1 |
| Duplex | Full duplex |

Command string is <send> to read value. The read string holds all four pressures.

Example:

>send <cr>

1020.31 1020.32 1020.33 1020.32 hPa

| Warning: | For the purposes of this ILC it is strictly forbidden to make any adjustments to the measuring instrument. The only actions permitted are the movement among functions. |
| --- | --- |

# 8 Appendix D: Packaging instructions

Dimensions of the shipment package: width/length/height: 58 cm / 38 cm / 41 cm.

Weight of the shipment package (including instruments, case, connecting wires, manuals, ..): 15 kg.

# 9 Appendix E: Instrument Check List

Instrument Check List from Laboratory: ________

NOTE: Put in the table "OK" or write a comment if NOT OK.

Scan this document and send it to the coordinator: kouichi.nakashima@met.kishou.go.jp