| WORLD METEOROLOGICAL ORGANIZATION __________________ COMMISSION FOR INSTRUMENTS AND METHODS OF OBSERVATION CIMO EXPERT TEAM ON STANDARDIZATION First Session Geneva, Switzerland 26 – 29 November 2012 |  | CIMO/ET-Stand-1/Doc. 5 (26.XI.2012) _______ ITEM: 5 Original: ENGLISH |
| --- | --- | --- |

Maintenance classification

(Submitted by M. Leroy, Météo-France)

Summary and purpose of document This document provides information on the maintenance classification. Changes compared to the report of the Meeting of the CIMO Ad-hoc Working Group on the WIGOS Pilot Project, Third Session are indicated in revision mode. They are based on the definition of the maintenance classification used by Meteo-France

Action proposed

The Meeting is invited to note the information provided in this document and to take into consideration for its further deliberations

________________

| Parameter | Class A | Class B | Class C | Class D |
| --- | --- | --- | --- | --- |
| Air temperature | 0.2°C (achievable measurement uncertainty). Temperature probe with uncertainty below or equal 0.05 °C (in laboratory conditions, over the measuring range). Uncertainty of the acquisition system < 0.02 °C. High performance artificially ventilated screen. Laboratory calibration of the temperature probe every year. | 0.5 °C Temperature probe with uncertainty below 0.25°C (corresponds of class A of IEC 751 standard, Pt100 platinum probe). Acquisition uncertainty < 0.1°C. Radiation screen with known characteristics and over-estimation of Tx (daily max. temperature) < 0.25°C in 95% of cases. Laboratory calibration of the temperature probe every 5 years. | 1.0°C Temperature probe with uncertainty < 0.4°C, with good stability, such as Pt100. Acquisition uncertainty < 0.2°C. Radiation screen with known characteristics and over-estimation of Tx < 0.7°C in 95% of cases. | > 1°C Temperature probe and/or acquisition system uncertainty lower than for class C. Unknown radiation screen or with “unacceptable” characteristics (for example, over-estimation of Tx > 0.7°C in 5% of cases). |
| Relative humidity | 3% (achievable measurement uncertainty). Performance verified over the full range of humidity and a temperature range typical for the location of the station. Acquisition uncertainty < 0.2%. Calibration every 6 months, in an accredited laboratory. | 6% Sensor specified for ( 6%, over a temperature range typical for the location of the station. Acquisition uncertainty < 1%. Calibration every year, in an accredited laboratory. | 10% Sensor specified for ( 10%, over a temperature range typical for the location of the station. Acquisition uncertainty < 1%. Calibration every two years in an accredited laboratory. | > 10% Sensor with performances or specifications worst than ( 10% over the common temperature conditions or Calibration not organized. |
| Atmospheric pressure | 0.3 hPa (achievable measurement uncertainty). Sensor with a numeric output. Influence of dynamic pressure due to wind reduced by a static head. Yearly calibration in an accredited laboratory. | 0.5 hPa Sensor with a numeric output. Sensor specified for ( 0.5 hPa, including possible drift between calibrations. Two-year calibration in an accredited laboratory. | 1 hPa Sensor specified for ( 1 hPa, including possible drift between calibrations. Calibration organized for this uncertainty. | > 1 hPa Specifications lower than for class C or no regular calibration organized. |

| Wind | Wind speed: 10% (or 0.5 m/s) Starting threshold (for wind speed) ( 0.5 m/s wind direction: 5° Calculation of wind parameters following WMO recommendations: 4 Hz samples, gust over a 3 seconds period. Yearly control of bearings, for rotating anemometers. Yearly calibration. Note : wind speed uncertainty could be reduced to 5% for wind energy. To be changed if 5% is introduced in the CIMO guide. | Wind speed: 10% (or 0.5 m/s) Starting threshold (for wind speed) ( 1 m/s wind direction: 10° Calculation of wind parameters following WMO recommendations, with the possible difference concerning gust calculation: min. 1 Hz sampling, gust calculated over a period  3 s. Yearly control of bearings, for rotating anemometers. | Wind speed: 15% (or 0.5 m/s) Starting threshold (for wind speed) ( 2 m/s wind direction: 10° Two-year control/maintenance of the mechanical status of sensors. | Wind speed: > 15% (or 1 m/s) Wind Direction: > 20° Starting threshold (for wind speed) > 2 m/s. Or no regular maintenance organized. |
| --- | --- | --- | --- | --- |
| Precipitation (liquid). Classification still to be defined for solid precipitation. To be checked with the standardization proposed in Item 6 (Luca G. Lanza) | The larger of 5% and 0.1 mm. (achievable measurement uncertainty). Reported resolution better than or equal to 0.1 mm. If any, error related to precipitation intensity corrected. Use of a wind shield. Daily control of the collecting cone for rain gauges using a cone. 6 months calibration for tipping bucket rain gauges. | The larger of 5% and 0.2 mm. Reported resolution better than or equal to 0.2 mm. If any, error related to precipitation intensity corrected or at least known. 6 months calibration for tipping bucket rain gauges. Weekly control of the collecting cone for rain gauges using a cone. | The larger of 10% and 0.5 mm. Unknown error related to precipitation intensity. Calibration period of tipping bucket rain gauges lower than 18 months. A preventive maintenance is defined and applied. | > 10% or no control and adjustment methods defined or no regular maintenance organized. |

| Global solar radiation | Pyranometer of ISO class 1. 5% for daily total. Ventilated sensor. Calibration every two years. Regular cleaning of the sensor (at least weekly and daily in case of lithometeor deposition). | Pyranometer of ISO class 1. No ventilation. Calibration every two years. Regular cleaning of the sensor (at least weekly). | Pyranometer of ISO class 2. No ventilation. Calibration every five years. No regular cleaning of the sensor. | Uncertainty > 10% for daily total or sensor not using a thermopile. Or Calibration not organized |
| --- | --- | --- | --- | --- |
| Visibility (MOR) | 50 m below 600 m, 10% between 600 and 1500 m, 20% above 1500 m. All, in 95% of cases in homogenous visibility conditions (ratio of standard deviation to mean value over 10 minutes < 0.1). 3 months calibration (or periodicity recommended by the manufacturer, if lower). At least, weekly cleaning of the optics. | The larger of 20% and 50 m, up to 10000 m. In 90% of cases in homogenous visibility conditions. 6-months calibration (or periodicity recommended by the manufacturer). For forward scatter meters, full control of the calibration chain: reference transmissometer, transfer control forward scatter meter, calibration plates. Use of internal warning from the sensor to clean the optics. | The larger of 40% and 100 m, up to 10000 m. Yearly calibration. Defined calibration chain (and applied !). | Specifications lower than for class C or No control and adjustment methods defined or No regular maintenance organized. |
| Temperature above or below ground level. | Not specified by WMO. 0.5°C Laboratory calibration of the temperature probe every 2 years, for temperature above ground. | 1°C Temperature probe with uncertainty < 0.25°C (corresponds to class A of IEC 751 standard, Pt100 platinum probe). Acquisition uncertainty < 0.1°C. Laboratory calibration of the temperature probe every 5 years, for temperature above ground.. | 1.5°C Temperature probe with uncertainty < 0.4°C. Acquisition uncertainty < 0.2°C. | Specifications lower than for class C or Height (or depth) of measurement unknown. |