| WORLD METEOROLOGICAL ORGANIZATION __________________ COMMISSION FOR INSTRUMENTS AND METHODS OF OBSERVATION INTERNATIONAL ORGANIZING COMMITTEE (IOC) FOR THE WMO SOLID PRECIPITATION INTERCOMPARISON EXPERIMENT (SPICE) Fourth Session Davos, Switzerland 17 – 21 June 2013 |  | CIMO/SPICE-IOC-4/Doc. 3(5) REV. 1 (21.VI.2013) _______ ITEM: 3 Original: ENGLISH |
| --- | --- | --- |

SITE REPORT FOR HAUKELISETER (NORWAY)

(Submitted by Mareile Wolff)

Summary and purpose of document This document provides site report of Haukeliseter (Norway) for the 2012/2013 winter.

Action proposed

The Meeting is invited to take this information into consideration when deciding on necessary modifications and clarifications on the overall set-up of the experiment and procedures to be followed to ensure best quality observations are collected on all sites and appropriate coordination and data transfer mechanisms are implemented.

________________

Site Report v.1.2

# Site: Haukeliseter, Norway

Date: June 21, 2013

## Site layout

## Configuration of references:

Covering: gauge used, heating (hardware, algorithm), sampling strategy, physical configuration (height, shields, etc):

| Reference type | Gauge | Heating | shield | Data sampling interval | Sampling strategy | Output interval | Height of the rim |
| --- | --- | --- | --- | --- | --- | --- | --- |
| R1 | - |  |  |  |  |  |  |
| R2 | Geonor, 1000mm, 3 transducers | WMO-SPICE heaters and algorithm | Alter windshield, mounted at same height as Geonor opening, to be changed in Summer 2013 | 6s | Period averaging | 1 min (6s data to be stored locally, no real time access) | 4.5m |
| R3 | Geonor, 1000mm, 3 transducers | WMO-SPICE heaters and algorithm | Shielded /unshielded pair; shielded Geonor has Alter windshield, mounted at same height as Geonor, to be changed in Summer 2013 | 6s | Period averaging | 1min (6s data to be stored locally, no real time access) | 4.5m |

## Changes made during the season 12/13, if any.

List changes made to the reference configuration from Dec 01, 2012 to April 30, 2013, indicating the reason and the impact.

- Installation of capacitive precipitation detector at the test-site, 18.January 2013.

- We had not enough antifreeze solution to provide all gauges with sufficient antifreeze before the 18. January 2013 – thereafter issue solved

- Installation of unshielded Geonor in mid-December 2012

- Several changes of Geonor-transducers because of strong signal variations

## Issues: heating, data quality, vibrations, capping;

Issues noted during the operating of the site related to:

- Heating: The WMO-SPICE heater does not reach positive temperatures for strong wind situations, see figure:

- Heating: Minimum threshold (-5°C) too high for “heating off”, see figure

- Capping: maybe, we have seen some sudden bucket-content raises. These cases are not yet evaluated. No webcam available at that time.

- Vibrations of the mast: We chose very stable gittermasts to minimize vibrations. But we do experience vibrations on the link between the gittermast and the Geonor pedestal, which can be seen in the Geonorsignal

- Vibration of the gauge: we do see wind-induced vibrations in the Geonor-signal. Also at the reference in the DFIR, see figure (data from 8.3.2013, T_Air between -7°C--12°C, FF between 0-3m/s:

- Orientation: Different orientation of various precipitation detectors seem to influence the data, see figure under precipitation detectors (10).

- Grounding: Grounding could not be completed because of most of the work was done during winter, to be completed in Summer 2013

- Length and type of cables and connections used: We set up a little cabin (container) at the test site to house the dataloggers and the local back-up machine. That required the use of long cables for all systems. Not all of them were tested with the chosen cable-lengths. We have issues of unknown reason for a humidity sensor and two of 18 Geonor transducers.

- Materials:-

- Data sampling – 6s data are not recorded before April 2013 – there is no real time access, data are not yet evaluated

- Data output – there were difficulties to setup a working datastream from the test site to the long-time site-archive. Automatic data transfer started in the end of May, we are currently in the process of transferring the complete data set of the first winter to our long-time site archive. Due to the amount of data and the now working automatic hourly datastream does this work go very slowly. Expected to be finished in July 2013.

- Data quality – We still have two Geonor-transducers (one of them in the shielded reference) which show enormous variations due to unknown reasons; several changes of the transducers haven’t solved the issue. We are working on improving the grounding during the summer season.

- Instrument specific issues: The site manager suggests using the Geonor X3 as the shielded R3-reference instead of Geonor X2 (which has a unsolved transducer issue). The X3-Geonor had the Geonor-heater installed during the 2012/2013 winter. That can be changed to the WMO-SPICE heating system during the summer period

## Heating report:

    - Summary of configurations throughout the winter of 12/13:

        - Hardware:

          - how was implemented; We hired the company Geonor to install the WMO-SPICE heaters onto the instruments

          - power applied; We use 24 V DC, heater elements in serial in order to have 50W in each (similar to 12V and parallel coupling)

        - software:

          - upper limit temperature: Instrument temperature =2deg C

          - lower limit temperature: T Air= -5 deg C

          - heating interval

          - other factors considered.

    - Summary of changes made during the season: no changes yet

    - Effects noted: e.g. not sufficient melting of snow, too much or not enough heat, evaporation, degrading of gauge performance, increased level of noise; need to extend heating to additional surfaces, etc.: Not sufficient (does not reach 0 degrees) of heating under wind conditions. We experience “strong”

    - Capping: noted, suspected; signature? – suspected, try to find examples

    - Changes needed for the future? – Try with parallel coupling of heating elements, which would give 100W heating from each element

    - Provide pictures! (currently searching for pictures)

    - Any other topic of challenge with impact on data, specific to your experience.

## What has worked well;

Receiving of the test instruments, contact with Campbell before the installation of the PWS100, contact with MPS System went well after some “starting difficulties”.

## What has not worked that well: lessons learned;

## Data available:

    - # of days of data collected for each sensor on site:

      - 3 Reference Geonors (R2, R3)  130 days

      - Geonor 130 days

      - Pluvio2 130 days

      - TRWS 405 130 days

      - Several wind sensors at gauge height 130 days

      - Wind sensor 10 m 130 days

      - PT 100 130 days

      - Vaisala HMP155 130 days

      - Campbell PWS100 130 days

      - Ott Parsivel 80 days

      - Thies LPM 130 days

      - Vaisala PWD 21 90 days

      - 6 optical precipitation detectors (Thies) 130 days

      - Capacitive precipitation detector (RQA6.1) 130 days

      - Laser Snow Height sensor 130 days

      - Webcamera 60 days

    - Data transmitted to NCAR – NONE, see issues (data output)

    - Data QC’d – only manual QC, not regularly, not enough manpower to do daily checkups

    - Issues in data – humidity doesn’t work, two Geonor-transducers have large variations, Ott Pluvio has issues which needs to be checked

## Instruments under test: list, issues

- Have you had problems with the instruments under test? No obvious issues, but needs to be controlled by instrument provider

- Have all instruments allocated to your site from Instrument Providers, been installed? YES

- Has the data been shared with the Instrument Providers? NO –  that is planned

- Have them visited the site? NO

## Information on the Precipitation Detector(s) used;

experience regarding their ability to detect precipitation (if applicable) –

- RQA6.1 (MPS System) as capacitive precipitation detector, not yet analysed.

- Thies optical precipitation detectors seems to be highly dependent on their orientation (wind dependency)

- Recommendation: Is it possible to install the reference precipitation detector inside the DFIR?

## Commissioning:

- date; incomplete commissioning in February

- configuration at commissioning,

- availability of report.  In progress

## Results to date:

- summary of preliminary results, if available.

- Plans for data analysis: focus, interests, timelines, plans, etc….

## Interaction Site manager and the IOC and Project team

- interactions,

- information sharing,

- feedback, etc…

## Small things, big impacts?

What can we do differently?