BETOLOUM BEINDE                                                INSTRUMENT INTER-

DIRECTORY WATER RESOURCES                        COMPARISONS AND TESTBED

AND METEOROLOGY                                               EXPERIMENTS

BP: 429

Office phone: +235 252 30 81

Handy: +235 66 48 29 15

E.mail: betololoumbeinde@yahoo.fr                                 ABSTRACT

sacdrem@intnet.td

For measuring the hydro-meteorological parameters (River water level, air temperature, humidity, the dry and the wet bulb…), we use classical instruments and the new concept of automatic in the member countries of world Meteorological Organization (WMO).

In Chad, modern screen for measuring level water were installed in 1996.

The results of measurement vary significantly.

The level water in automatic instrument (DCP) and traditional (gauge) is ranging from 2cm to 5cm.

This difference of measurements cannot give a good data series for analysis or not having fixed plan

INTRODUCTION

There are many changes at measurement level with the new automatic network: data acquisition, frequency, instruments characteristics, installation features….Many impact can be produced on data series by these changes, which many appear at post-processing level or climatologically analysis.

The purpose of this work is not to find data series inhomogentics but it is to try to reassure some of the mains changes that we can expect on the new data series that  result from new methodological errors. This are mainly due to the different time response and measuring frequency of sensors witch limit the atmospheric scales to be observed.

CHS (classical hydrological systems) values and AHS (automatics hydrological systems) located in MOUNDOU/Chad were calculed with month collected recorded. These calculate is the average.

The wide use of AHS in Chad network is a major change in the methods of CHS because the data in long climatological series have been recorded with those methods. It is pertinent to ask if we can continue olds series with new data coming from AHS. With the new automatic network there are many changes at measurement level. For example data acquisition, frequency, instruments characteristics, installation, features…, like inhomogentics in data series, which many appear on climatological analysis or not having fixed plan. This work is about getting knowledge on the difference between results obtained with classical and automatic systems for the daily water level of rivers value using statical methods for the comparison of data series.

Data from one station in Chad have been using:

Moundou: This site had good regular functioning of classical (gauge) and automatic (DCP) systems for more than one month producing two daily AHS series.

Classical hydrological is measured by qualified personnel. One of the other hand AHS use electrical or battery sensors to record the level of rivers values every minutes or hour (sampling interval complies with WMO recommendation all day long).

The level water measured in automatic instrument with the DCP (Data collect plat form) has been used in several station of hydrological. The comparative measure with the reading in the station gauge (limnimeter) and the DCP is ranging to one centimetre to ten (Fig).

The inconvenient of this instrument is the battery which we can recharge all five months of use and the duration of charge three or four hours. But if the tension comes down to 6VDC it is very difficult to recharge or to use it.

The evaluation of this parameter summarized is that:

The level water river unexpected results were nowhere found. For both stations the monthly differences fluctuated between one to ten centimetres.

This difference of measurements cannot give a good data series for analysis or not having fixed plan.

River: Logone station: Moundou instrument: Gauge+ DCP- month: July

| Date | Gauge 7H | DCP | RANGING |
| --- | --- | --- | --- |
| 1 | 210 | - | - |
| 2 | 228 | 222 | 06 |
| 3 | 238 | - | - |
| 4 | 236 | - | - |
| 5 | 227 | 220 | 16 |
| 6 | 243 | 237 | 10 |
| 7 | 261 | 259 | 16 |
| 8 | 254 | 253 | 01 |
| 9 | 249 | 242 | 07 |
| 10 | 244 | 236 | 03 |
| 11 | 258 | 251 | 09 |
| 12 | 272 | 263 | 05 |
| 13 | 277 | 274 | 03 |
| 14 | 267 | 261 | 06 |
| 15 | 260 | 256 | 04 |
| 16 | 254 | 247 | 07 |
| 17 | 256 | 244 | 12 |
| 18 | 260 | 252 | 08 |
| 19 | 268 | 263 | 05 |
| 20 | 261 | 253 | 08 |
| 21 | 261 | 256 | 05 |
| 22 | 273 | 268 | 05 |
| 23 | 266 | 259 | 07 |
| 24 | 274 | 264 | 10 |
| 25 | 282 | 273 | 11 |
| 26 | 281 | 274 | 07 |
| 27 | 267 | 261 | 04 |
| 28 | 264 | 253 | 07 |
| 29 | 359 | 329 | 30 |
| 30 | 359 | 352 | 07 |
| 31 | 338 | 331 | 07 |