Intercomparison between Calibration Laboratories

of Central and Eastern European Countries

Ján Danč, Oľga Novanská

Slovak Hydrometeorological Institute, Bratislava, Slovak Republic

Microstep-MIS, Bratislava, Slovak Republic

Very effective form of the verification of calibration qualities is Interlaboratory Comparisons (ILC). Following WMO recommendation, ILC was organized by Calibration Laboratory of SHMU - RIC Bratislava as Pilot laboratory in the field of temperature, barometric pressure and relative humidity for Calibration laboratories of meteo-services of Central and Eastern European Countries. ILC was realized in accordance with the EAL-P7. Five Calibration laboratories took a part in this action. All participant laboratories have the numerical value of En < 1 for all quantities and measuring points.

The Calibration Laboratory of Slovak Hydrometeorological Institute is accredited since June 25th 2002 in accordance with the International Standard ISO/IEC 17025 by the assessment of the accreditation body in Slovakia SNAS (Slovak National Accreditation Service, accreditation number: K 017) for the calibration of the following instruments: classic liquid-in-glass thermometers, PRTs platinum resistance thermometers, thermographs, mechanical and electronic hygrometers, hygrographs, electronic barometers and anemometers. Since June 17th 2005, the Laboratory has accreditation in calibration of automatic rain gauges and for calibration of air pollution analysers of sulphur dioxide, carbon monoxide, nitrogen monoxide, nitrogen dioxide and ozone. Non-accreditation activities of the Calibration Laboratory are as follow: the field calibration of rain gauges and barometers and checking the composition of calibration gas mixtures.

Resolution 4.2/5 adopted two Regional Instrument Centres for RA VI. - Calibration Laboratory of the Slovak Hydrometeorological Institute and Calibration Laboratory of the Environmental Agency of the Republic of Slovenia, providing services predominantly in Central and Eastern and South-eastern parts of Europe, respectively, with the following functions:

- To assist Members of the Region in calibrating their national standard meteorological instruments,

- To organize instrument evaluations and comparisons;

- To assist Members of the Region in building Calibration Laboratory (CL)

- Quality assurance procedures

- to assist Members of the Region in the case of need accreditation in accordance with the International Standard ISO/IEC 17025

Laboratories continually strive to ensure the quality of instruments, procedures, and capabilities. No matter how carefully they conduct these internal checks, systematic deviations often go unnoticed. It is therefore important that they regularly compare their measurement results with those of other laboratories.

Participation in interlaboratory comparisons is a key criterion for accreditation too, but is also a vital investment in reliability, operational certitude, and staff confidence. Therefore proficiency testing schemes (PTS) are regularly organized interlaboratory comparisons to determine the technical competence of laboratories.

ILC was organized by Calibration Laboratory of SHMU - RIC Bratislava as Pilot laboratory in the field of temperature, barometric pressure and relative humidity for Calibration laboratories of meteo-services of Central and Eastern European Countries. ILC was realized in accordance with the EAL-P7.

The results of the interlaboratory comparisons are strictly confidential. The results will be presented to the other participants in an anonymous fashion. Each laboratory individually decides who else is allowed to learn about their performance.

Interlaboratory comparisons offer a broad assessment of measurement accuracy. This not only extends to a laboratory’s measurement standards, but also to underlying procedures. Incorrect procedures can cause systematic deviations in the measurements.

An En value is then calculated for each laboratory, based on comparison of the lab’s measurements with reference values (Y and Yref), in combination with the reported uncertainties for these two values (U and Uref).

If the absolute value of En is less than or equal to 1, the lab’s results are in agreement with the reference values. However, if the absolute value of En is greater than 1, the laboratory must determine the cause of the systematic deviations, which may result from improper standards, instruments or procedures.

To establish good organization of ILC, it was sent the questionnaire on calibration possibilities of participants. On base of answers the Pilot Laboratory supplies this action by suitable instruments.

These instruments will be calibrated in concerned CLs, using their standard procedures. After calibration, each participant sent instruments to next CL according time table (schedule) and sent their calibration results (certificates, protocols) to Pilot laboratory. After finishing ILC the Pilot laboratory will perform data evaluations in the Report.

Instrumentation for temperature:

indicator testo 400, No. 01426399/710

sensor PRT, testo 0614.0240 , No. 01394969/710

Instrumentation for relative humidity:

indicator testo 400, No. 01426399/710

sensor testo 0636.9715, No. 710

Instrumentation for barometric pressure:

Vaisala, PTB220A, No. C4050004

Participants:

- Czech Hydrometeorological Institute, Prague, Czech Republic,

- IMGW, Warsaw, Poland,

- Estonian Meteorological and Hydrological Institute, Tallin, Estonia,

- Latvias Vides Geologijasun meteorologijas agentura, Riga, Latvia,

- Lithuanian Hydrometeorological Service, Vilnius, Lithuania.

Time table of action:

    - calibration of RH sensor in DKD-testo, Germany,

14.11.2007	calibration of all instruments in Pilot Laboratory,

22.11.2007	Czech Hydrometeorological Institute,

return,

    - calibration of thermometer and barometer in Slovak Institute of Metrology,

break - structure reorganization of Slovak Hydrometeorological Institute,

    - calibration of all instruments in Pilot Laboratory,

    - Latvias Vides Geologijasun Meteorologijas Agentura, Riga,

    - Estonian Meteorological and Hydrological Institute, Tallin,

    - Lithuanian Hydrometeorological Service, Vilnius,

    - IMGW, Warsaw, Poland.

return

    - calibration of all instruments in Pilot Laboratory.

We would like to thank to Latvias Vides Geologijasun Meteorologijas Agentura, Riga, for significant help in organization of this ILC.

Evaluation:

The following table shows the results from the participating laboratories with number 1 to 5. Each laboratory was given a number which is used in the table for each measuring point.

| Quantity | laboratory |
| --- | --- |
|  | Pilot | 1 | 2 | 3 | 4 | 5 |
| Temperature | U | U | En | U | En | U | En | U | En | U | En |
| -20 °C | 0,05 | 0,023 | 0,11 | 0,25 | 0,04 | 0,077 | 0,07 | 0,22 | 0,09 | 0,024 | 0,11 |
| 0 °C |  | 0,023 | 0,00 | 0,24 | 0,08 | 0,077 | 0,56 | 0,11 | 0,08 | 0,024 | 0,00 |
| 20 °C |  | 0,023 | 0,07 | 0,25 | 0,12 | 0,084 | 0,10 | 0,09 | 0,19 | 0,024 | 0,02 |
| Pressure | U | U | En | U | En | U | En | U | En | U | En |
| 850 hPa | 0,12 | 0,18 | 0,11 | 0,58 | 0,05 | 0,086 | 0,07 | 0,16 | 0,35 | 0,13 | 0,28 |
| 900 hPa |  | 0,18 | 0,32 | 0,58 | 0,15 | 0,086 | 0,15 | 0,16 | 0,29 | 0,13 | 0,23 |
| 950 hPa |  | 0,18 | 0,21 | 0,58 | 0,15 | 0,085 | 0,07 | 0,16 | 0,35 | 0,13 | 0,23 |
| 1000 hPa |  | 0,18 | 0,16 | 0,58 | 0,12 | 0,090 | 0,01 | 0,16 | 0,35 | 0,13 | 0,23 |
| 1050 hPa |  | 0,18 | 0,42 | 0,58 | 0,17 | 0,094 | 0,10 | 0,16 | 0,35 | 0,13 | 0,28 |
| Humidity | U | U | En | U | En | U | En | U | En | U | En |
| 25 % | 2,5 | 0,28 | 0,14 | 1,43 | 0,03 | 2,0 | 0,03 | 1,30 | 0,18 | 1,8 | 0,22 |
| 50 % | 2,5 | 0,29 | 0,43 | 1,43 | 0,14 | 2,2 | 0,24 | 1,30 | 0,25 | 1,8 | 0,27 |
| 75 % | 2,5 | 0,41 | 0,38 | 1,43 | 0,28 | 2,1 | 0,12 | 1,30 | 0,35 | 1,8 | 0,16 |
| 90 % | 2,5 | 0,50 | 0,30 | 1,43 | 0,45 | 2,3 | 0,03 | 1,30 | 0,21 | 1,8 | 0,26 |

A result is considered satisfactory, when the numerical value of En < 1.

A following table shows that all laboratories for all quantities and measuring points have the numerical value of En < 1.