| WORLD METEOROLOGICAL ORGANIZATION COMMISSION FOR BASIC SYSTEMS ---------------------------- SECOND MEETING OF INTER-PROGRAMME EXPERT TEAM ON DATA REPRESENTATION AND CODES BRASILIA, BRAZIL, 31 AUGUST-3 SEPTEMBER 2010 |  | IPET-DRC-II / Doc.3.3(8) (6.VIII.2010) ------------------------- ITEM 3.3 ENGLISH ONLY |
| --- | --- | --- |

BUFR/CREX descriptors related to AWS

Submitted by Eva Červená (Czech Republic)

________________________________________________________________

Summary and Purpose of Document

The document contains BUFR/CREX descriptors that have been proposed or modified since the IPET-DRC I meeting, and a revised version of Functional Specifications for AWS (Rec. 6.1/2 (CBS-XIV)).

________________________________________________________________

ACTION PROPOSED

The meeting is invited to discuss the content of this document and recommend the included BUFR/CREX descriptors for validation.

References:

[1]  Manual on Codes, WMO-No. 306, Volume I.2.

[2]  Draft Recommendation 6.1/2 (CBS-XIV), Commission for Basic Systems, Fourteenth

Session, Dubrovnik, Croatia, 2009.

[3] Joint Meeting of the CT on Migration to Table Driven Code Forms and of the ET on

Data Representation and Codes. Final Report, Geneva, 1 - 5 September 2008.

[4] First meeting of the Inter-Programme Expert Team on Data Representation and

Codes. Final Report, Geneva, 15 - 18 September 2009.

5] Sixth meeting of Expert Team on Requirements and Implementation AWS Platforms.

Final Report, Geneva, 22 to 25 June 2010.

5 - 8  December 2005ntion on International Civil Aviation

1. Background

14th Session of the Commission for Basic Systems (Dubrovnik, Croatia, 25 March to 2 April 2009) reviewed the Functional Specifications for Automatic Weather Stations and adopted draft Recommendation 6.1/2 (CBS-XIV).

The author presented a document to ET-AWS-6 (Geneva, 22 to 25 June 2010), containing detailed analysis of the requirements for representation of the listed variables, a revised version of the Functional Specifications for AWS and a proposal for several new BUFR/CREX descriptors. These descriptors had beed recommended for validation by the IPET-DRC meeting (Geneva, 15 – 18 September 2009).  The ET-AWS-6 meeting examined carefully all proposals in the document ET-AWS-6/Doc. 10.1, agreed with the proposed new BUFR/CREX descriptors and procedures for representation of the listed variables, and suggested some modifications of the requirements. The ET AWS members were encouraged to participate actively in the validation process of the proposed new descriptors.

The proposed or modified BUFR/CREX descriptors are included in paragraph 2 of this document. The problem of terminology for radiation is discussed in paragraph 3. The Revised Functional Specifications for Automatic Weather Stations of June 2010 are included in paragraph 4.

2. BUFR/CREX descriptors proposed or modified since IPET-DRC-I meeting

For UV-B radiation

A proposal for representation of UV radiation was developed in cooperation of Météo-France and the Czech Hydrometeorological Institute (CHMI). The proposed descriptor 0 14 072 was recommended for validation in September 2008 [3]. At the ET-AWS-6 meeting, Mr. Michel Leroy formulated new requirements for UV-B radiation; based on them descriptor 0 14 072 has been modified as shown below.

| Table reference |  | BUFR | CREX |
| --- | --- | --- | --- |
| F X Y | Element name | Unit | Scale | Reference value | Data width | Unit | Scale | Data width |
| 0 14 072 | Global UV irradiation | J m-2 | 0 | –4000000 | 23 | J m-2 | 0 | 7 |

Note:

- Global UV irradiation (0 14 072) is UV energy integrated over period specified for spectral band specified. 0 14 072 shall be preceded by a time period descriptor and by 0 02 071 (Spectrographic wavelength) and 0 02 072 (Spectrographic width). E.g. If 0 14 072 is used for Global UV-B irradiation, 0 02 071 and 0 02 072 shall specify spectral band 280 to 315 nm.

For Humidity

The ET-AWS-6 meeting recommended changing the name of 0 13 111 from “Soil moisture, volumetric or water potential” to “Soil moisture”.

| Table reference |  | BUFR | CREX |
| --- | --- | --- | --- |
| F X Y | Element name | Unit | Scale | Reference value | Data width | Unit | Scale | Data width |
| 0 13 111 | Soil moisture | g kg-1 | 0 | 0 | 10 | g kg-1 | 0 | 4 |

For Lightning

Detailed discussion about the unit of 0 20 126 (Lightning rate of discharge) has resulted into deleting of the note originally proposed to refer to this descriptor and to modification of the unit to become “h-1” preferred by the ET-AWS-6 meeting.

Additional discussion took place in July 2010, regarding the data width of 0 20 126. Ms Paula Taylor (UK) explained that the maximum value 100 000 requested by the ET AWS corresponded to sensors capable of detection up to 30 km, but not to detectors with the operation range of 200 km. Consequently, the data width of 0 20 126 has remained unchanged, but the data width of 0 20 127 (Lightning – distance from station) was modified to cover the range of 200 km. Moreover, also scale and data width of 0 20 128 (Lightning – direction from station) was modified as the existing scale 0 (the direction expressed in whole degrees) would lead to gross errors at longer detection ranges.

| Table reference |  | BUFR | CREX |
| --- | --- | --- | --- |
| F X Y | Element name | Unit | Scale | Reference value | Data width | Unit | Scale | Data width |
| 0 20 126 | Lightning rate of discharge | h-1 | 0 | 0 | 23 | h-1 | 0 | 7 |
| 0 20 127 | Lightning - distance from station | m | -3 | 0 | 8 | m | -3 | 3 |
| 0 20 128 | Lightning - direction from station | Degree true | 1 | 0 | 12 | Degree true | 1 | 4 |

For Depth of fresh snow

The existing BUFR/CREX descriptor 0 13 012 Depth of fresh snow (m, 2, -2, 12) is not capable to meet the requested demand for high precision (in 0.001 m). To avoid the necessity to apply all three operators 201Y, 202Y and 203Y, a new descriptor 0 13 118 was proposed for validation. This proposal had been discussed with IPET-DRC members in February 2010, and consequently supported by the ET-AWS-6 meeting.

| Table reference |  | BUFR | CREX |
| --- | --- | --- | --- |
| F X Y | Element name | Unit | Scale | Reference value | Data width | Unit | Scale | Data width |
| 0 13 118 | Depth of fresh snow (high accuracy) | m | 3 | 0 | 14 | m | 3 | 5 |

For Soil heat flux

At the end of the ET-AWS-6 meeting, the author of this document received preliminary specification for “Soil heat flux” from Mr. Francesco Sabatini (CAgM). Based on this requirement, a new BUFR/CREX descriptor 0 14 057 (Soil heat flux) is proposed for validation.

| Table reference |  | BUFR | CREX |
| --- | --- | --- | --- |
| F X Y | Element name | Unit | Scale | Reference value | Data width | Unit | Scale | Data width |
| 0 14 057 | Soil heat flux | J m-2 | –1 | –1000 | 11 | J m-2 | –1 | 4 |

3. Radiation terminology

At the end of June 2010, Mr. Karl Monnik, the Chair of ET-AWS, distributed a proposal for replacing the terms “radiation” and “irradiation” in the names of the descriptors from Class 14, currently in the stage of validation, by “radiat exposure".

According to radiation experts of CHMI, the following terms could be used (all of them correct, having identical meaning) :

“radiation integrated over period specified” or

"irradiation" or

“radiat exposure".

If the term “radiat exposure" were to be used in the names of BUFR/CREX descriptors, it would have to be applied not only in the descriptors proposed for validation, but also in the already operational entries of Class 14.

Note: Mr. Monnik spotted an editorial mistake in the name of 0 14 054 that should be corrected to “Photosynthetically active radiation, integrated over period specified”.

4. Revised Annex to draft Recommendation 6.1/2 (CBS-XIV) – JUNE 2010

Functional Specifications for Automatic Weather Stations

| VARIABLE 1) | Maximum Effective Range 2) | Minimum Reported Resolution 3) | Mode of Observation 4) | BUFR / CREX 5) | Status 5) |
| --- | --- | --- | --- | --- | --- |
| ATMOSPHERIC PRESSURE |  |  |  |  |  |
| Atmospheric Pressure | 500 – 1080 hPa | 10 Pa | I, V | 0 10 004 | OP |
| TEMPERATURE 9) |  |  |  |  |  |
| Ambient air temperature (over specified surface) 14) | -80 °C – +60 °C | 0.1 K | I, V | 0 12 101 | OP |
| Dew-point temperature 14) | -80 °C – +60 °C | 0.1 K | I, V | 0 12 103 | OP |
| Ground (surface) temperature (over specified surface) 14) | -80 °C – +80 °C | 0.1 K | I, V | 0 12 120 | VAL |
| Soil temperature14) | -50 °C – +50 °C | 0.1 K | I, V | 0 12 130 | OP |
| Snow temperature 14) | -80 °C – 0 °C | 0.1 K | I, V | 0 12 131 | VAL |
| Water temperature - river, lake, sea, well | -2 °C – +100 °C | 0.1 K | I, V | 0 13 082 or 0 22 043 | OP OP |
| HUMIDITY 9) |  |  |  |  |  |
| Relative humidity | 0 – 100% | 1% | I, V | 0 13 003 | OP |
| Mass mixing ratio | 0 – 100% | 1% | I, V | 0 13 110 | VAL |
| Soil moisture | 0 – 103 g kg-1 | 1 g kg-1 | I, V | 0 13 111 | VAL |
| Water vapour pressure | 0 – 100 hPa | 10 Pa | I, V | 0 13 004 | OP |
| Evaporation/evapotranspiration | 0 – 0.2 m | 0.1 kg m-2 0.0001 m | T | 0 13 033 | OP |
| Object wetness duration | 0 – 86 400 s | 1 s | T | 0 13 112 | VAL |
| WIND |  |  |  |  |  |
| Direction | 0 11,13); 1( – 360( | 1( | I, V | 0 11 001 | OP |
| Speed | 0 – 75 m s-1 | 0.1 m s-1 | I, V | 0 11 002 | OP |
| Gust Speed | 0 – 150 m s-1 | 0.1 m s-1 | I, V | 0 11 041 | OP |
| X,Y component of wind vector Z component of wind vector (horizontal and vertical profile) | -150 – 150 m s-1 -40 – 40 m s-1 | 0.1 m s-1 | I, V | 0 11 003 0 11 004 0 11 006 | OP OP OP |
| Turbulence type (Low levels and wake vortex) 16) | up to 15 types | BUFR Table Not specified yet | I, V | - | N |
| Turbulence intensity 16) | up to 15 types | BUFR Table Not specified yet | I, V | - | N |
| RADIATION 6) |  |  |  |  |  |
| Sunshine duration | 0 – 86 400 s | 60 s | T | 0 14 031 | OP |
| Background luminance | 0 – 1∙105 Cd m-2 | 1 Cd m-2 | I, V | 0 14 056 | VAL |
| Global downward solar radiation | 0 – 1∙108 J m-2 | 1. 102 J m-2 | I, T, V | 0 14 028 | OP |
| Global upward solar radiation | -1∙108 – 0 J m-2 | 1. 102 J m-2 | I, T, V | 0 14 052 | VAL |
| Diffuse solar radiation | 0 – 1∙108 J m-2 | 1. 102 J m-2 | I, T, V | 0 14 029 | OP |
| Direct solar radiation | 0 – 1∙108 J m-2 | 1. 102 J m-2 | I, T, V | 0 14 030 | OP |
| Downward long-wave radiation | 0 – 6∙107 J m-2 | 1. 103 J m-2 | I, T, V | 0 14 002 | OP |
| Upward long-wave radiation | -6∙107 – 0 J m-2 | 1. 103 J m-2 | I, T, V | 0 14 002 | OP |
| Net radiation | -1∙108 – 1∙108 J m-2 | 1. 103 J m-2 | I, T, V | 0 14 053 | VAL |
| UV-B radiation 8) | 0 – 26∙104 J m-2 | 1 J m-2 | I, T, V | 0 14 072 | VAL |
| Photosynthetically active radiation 22) | 0 – 6∙107 J m-2 | 1∙103 J m-2 | I, T, V | 0 14 054 | VAL |
| Surface albedo | 0 – 100% | 1% | I, V | 0 14 019 | OP |
| Soil heat Flux | -1∙104 – 1∙104 J m-2 | 1. 101 J m-2 | I, T, V | 0 14 057 | VAL |

| CLOUDS |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Cloud base height | 0 – 30 km | 10 m | I, V | 0 20 013 | OP |
| Cloud top height | 0 – 30 km | 10 m | I, V | 0 20 014 | OP |
| Cloud type, convective vs. other types | up to 30 classes | BUFR Table | I | 0 20 012 | OP |
| Cloud hydrometeor concentration | 1 – 700 hydrometeors dm-3 | 1 hydrometeor dm-3 | I, V | 0 20 130 | VAL |
| Effective radius of cloud hydrometeors | 2∙10-5 – 32∙10-5 m | 2∙10-5 m | I, V | 0 20 131 | VAL |
| Cloud liquid water content | 1∙10-5–1.4∙10-2 kg m-3 | 1∙10-5 kg m-3 | I, V | 0 20 132 | VAL |
| Optical depth within each layer | Not specified yet | Not specified yet | I, V | - | N |
| Optical depth of fog | Not specified yet | Not specified yet | I, V | - | N |
| Height of inversion | 0 – 1 000 m | 10 m | I, V | 0 20 093 | VAL |
| Cloud cover | 0 – 100% | 1% | I, V | 0 20 010 | OP |
| Cloud amount | 0 – 8/8 | 1/8 | I, V | 0 20 011 | OP |
| PRECIPITATION |  |  |  |  |  |
| Accumulation 7) | 0 – 1600 mm | 0.1 kg m-2, 0.0001 m | T | 0 13 011 | OP |
| Depth of fresh snowfall | 0 –1000 cm | 0.001 m | T | 0 13 118 | VAL |
| Duration | up to 86 400 s | 60 s | T | 0 26 020 | OP |
| Size of precipitating element 17) | 1∙10-3 – 0.25 m | 1∙10-3 m | I, V | 0 13 058 0 20 066 | OP OP |
| Intensity - quantitative | 0 – 2000 mm h-1 | 0.1 kg m-2 s-1, 0.1 mm h-1 | I, V | 0 13 155 | OP |
| Type | up to 30 types | BUFR Table | I, V | 0 20 021 | OP |
| Rate of ice accretion | 0 – 1 kg dm-2 h-1 | 1∙10-3 kg dm-2 h-1 | I, V | 0 13 114 | VAL |
| OBSCURATIONS |  |  |  |  |  |
| Obscuration type | up to 30 types | BUFR Table | I, V | 0 20 025 | OP |
| Hydrometeor type | up to 30 types | BUFR Table | I, V | 0 20 025 | OP |
| Lithometeor type | up to 30 types | BUFR Table | I, V | 0 20 025 | OP |
| Hydrometeor radius | 2∙10-5 – 32∙10-5 m | 2∙10-5 m | I, V | 0 20 133 | VAL |
| Extinction coefficient | 0 – 1 m-1 | 0.00001 m-1 | I, V | 0 15 029 | VAL |
| Meteorological Optical Range 10) | 1 – 100 000 m | 1 m | I, V | 0 15 051 | VAL |
| Runway visual range | 1 – 4 000 m | 1 m | I, V | 0 20 061 | OP |
| Other weather type | up to 18 types | BUFR Table | I, V | 0 20 023 | OP |
| LIGHTNING |  |  |  |  |  |
| Lightning rates of discharge | 0 – 100 000 | h-1 | I, V | 0 20 126 | VAL |
| Lightning discharge type (cloud to cloud, cloud to surface) | 3 types | BUFR Code Table | I, V | 0 20 023 | OP |
| Lightning discharge polarity | 2 types | BUFR Code Table | I, V | 0 20 119 | VAL |
| Lightning discharge energy | Not specified yet | Not specified yet | I, V | - | N |
| Lightning - distance from station | 0 – 3∙104 m | 103 m | I, V | 0 20 127 | VAL |
| Lightning - direction from station | 1( – 360( | 1 degree | I, V | 0 20 128 | VAL |

| HYDROLOGIC AND MARINE OBSERVATIONS |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Flow discharge – river | 0 – 2.5∙105 m3 s-1 | 0.1 m3 s-1 | I, V | 0 23 040 | VAL |
| Flow discharge – well | 0 – 50 m3 s-1 | 0.001 m3 s-1 | I, V | 0 23 041 | VAL |
| Ground water level | 0 – 1 800 m | 0.01 m | I, V | 0 13 074 | VAL |
| Ice surface temperature 14) | -80 °C – +0 °C | 0.5 K | I, V | 0 12 132 | VAL |
| Ice thickness - river, lake 15) | 0 – 50 m | 0.01 m | I, V | 0 08 029 0 13 115 | VAL |
| Ice thickness - glacier, sea 15) | 0 – 4 270 m | 1 m | I, V | 0 08 029 0 13 115 | VAL |
| Water level | 0 – 100 m | 0.01 m | I, V | 0 13 071 0 13 072 | OP OP |
| Wave height | 0 – 50 m | 0.1 m | V | 0 22 021 | OP |
| Wave period 18) | 0 – 100 s | 1 s | V | 2 01 129 0 22 011 2 01 000 | OP |
| Wave direction | 0 13); 1 – 360 degrees | 1 degrees | V | 0 22 001 | OP |
| 1D spectral wave energy density 18) | 0 – 5x105 m2Hz-1 | 10-3 m2Hz-1 | V, T | 2 01 135 0 22 069 2 01 000 | OP |
| 2D spectral wave energy density 18) | 0 – 5x105 m2Hz-1 | 10-3 m2Hz-1 | V, T | 2 01 135 0 22 069 2 01 000 | OP |
| Water practical salinity 18) | 0 – 400 psu12 | 10-3 psu | I, V | 2 01 130 0 22 064 2 01 000 | OP |
| Water conductivity 18) | 0 – 600 S m-1 | 10-6 S m-1 | I, V | 2 01 132 0 22 066 2 01 000 | OP |
| Water pressure 18) 19) | 0 – 11x107 Pa | 100 Pa | I, V | 2 07 001 0 22 065 2 07 000 | OP |
| Ice thickness 18) | 0 – 3 m | 0.015 m | T | 2 01 133 2 02 129 0 20 031 2 02 000 2 01 000 | OP |
| Ice mass | 0 – 50 kg m-1 | 0.5 kg m-1 (on 32 mm rod) | T | 0 20 135 | VAL |
| Snow density (liquid water content) | 100 – 700 kg m-3 | 1 kg m-3 | T | 0 13 117 | VAL |
| Tidal elevation with respect to local chart datum 18) | -10 – +30 m | 0.001 m | I, V | 2 01 129 0 22 038 2 01 000 | OP |
| Tidal elevation with respect to national land datum 18) | -10 – +30 m | 0.001 m | I, V | 2 01 129 0 22 037 2 01 000 | OP |
| Meteorological residual tidal elevation (surge or offset) 18) 20) | -10 – +16m | 0.001 m | I, V | 0 22 040 Note 20) | OP |
| Ocean Current - Direction | 0 13); 1( – 360( | 1( | I, V | 0 22 004 or 0 22 005 | OP OP |
| Ocean Current - Speed | 0 – 10 m s-1 | 0.01 m s-1 | I, V | 0 22 031 or 0 22 032 | OP OP |
| OTHER SURFACE VARIABLES |  |  |  |  |  |
| Runway conditions | up to 10 types | BUFR Table | I, V | 0 20 085 | OP |
| Braking action/friction coefficient | up to 7 types | BUFR Table | I, V | 0 20 089 | OP |
| State of ground | up to 30 types | BUFR Table | I, V | 0 20 062 | OP |
| Type of surface specified | up to 15 types | BUFR Table | I, V | 0 08 010 | OP |
| Snow depth | 0 – 25 m | 0.01 m | T | 0 13 013 | OP |
| OTHER |  |  |  |  |  |
| Gamma radiation dose rate 21) | 1 – 103 nSv h-1 | 0.1 nSv h-1 | I, T | 0 24 014 | VAL |
| Categories of stability | 9 types | BUFR Table | I, V | 0 13 041 | OP |

Notes:

- Name of variable, in line with WMO vocabulary and Technical Regulations.

- Maximum Effective Range - Maximum range of measuring capability; units traceable to SI.

- Minimum Reported Resolution – Lower resolution of reporting is not permitted.

- Mode of Observation – Type of data being reported:

I:  Instantaneous – 1-minute value (instantaneous as defined in WMO-No. 8, Part II, paragraph  1.3.2.4);

V:  Variability – Average (mean), Standard Deviation, Maximum, Minimum, Range, Median, etc. of samples – those reported depend upon meteorological variable;

T: Total – Integrated value over defined period; maximum 24 hours for all parameters except radiation which requires a maximum of one hour (exception, see note 6), and precipitation accumulation (6 hours maximum). The relevant element descriptor shall be preceded by a time period descriptor 0 04 024 (in hours) or 0 04 025 (in minutes).

A:  Average (mean) value.

- BUFR/CREX descriptors for representation of the listed variables;

OP:   Operational descriptors of BUFR/CREX Table B, Version Number 14.

VAL: Descriptors to be validated.

N:     Not yet specified requirements.

- Radiation energy amounts are given over a 24-hour period.

- Maximum interval: 6 hours.

- Definition of UV-B according to WMO-No. 8 (Vol. 1, Chapter on Radiation). Descriptor 0 14 072 (Global UV irradiation) was recommended for validation in September 2008, revised in July 2010.

- Humidity related variables (i.e. dew point) expressed as temperature are collected under temperature.

- MOR uniquely related to "extinction coefficient", (, by MOR = -ln(5%)/(.

- Direction to indicate 0 (zero) if speed = 0.

- Absolute Salinity (kg.kg-1) is now being used for ocean applications (IOC Res XXV-7). However, salinity that is reported to national oceanographic data centers remains the Practical Salinity (psu). Ocean water is about 35 psu. Lake Asal (Ethiopia) is the most saline body of water on earth with 348 psu salt concentration.

- Calm.

- Temperature data represented by 0 12 101, 0 12 103, 0 12 113, 0 12 120, 0 12 130, 0 12 131 and 0 12 132 shall be reported with precision in hundredths of a degree even if they are measured with the accuracy in tenths of a degree. This requirement is based on the fact that conversion from the Kelvin to the Celsius scale has often resulted into distortion of the data values. Temperature t (in degrees Celsius) shall be converted into temperature T (in degrees Kelvin) using equation: T = t + 273.15.

- Ice thickness 0 13 115 shall be preceded by 0 08 029 (Surface type) set to 11, 12, 13 or 14 to specify river, lake, sea or glacier, respectively.

- If the UNIT is specified as “BUFR Table”, the BUFR descriptor cannot be proposed without the content of the table being available.

- 0 13 058 (Size of precipitation element) is capable to express size of any precipitation element, apart from hailstones. Size of hailstones shall be represented by 0 20 066.

- These requirements being confirmed, it has to be noted that the selected descriptors are suitable for the normal operating conditions and shall be combined with appropriate operator descriptors to allow representation of the extreme values or the requested high precision.

- Operator 2 07 Y is recommended to be used with Water pressure 0 22 065 (Pa, -3, 0, 17), if the data are produced in BUFR, Edition 4. The same result, i.e. change to (Pa, -2, 0, 21) would be obtained by the combined use of the less sophisticated operators 201Y and 202Y:

2 01 132

2 02 129

0 22 065 Reported value of “Water pressure”

2 02 000

2 01 000

- The following sequence is to be used to change data width and reference value of 0 22 040 (m, 3, -5000, 14) to become (m, 3, -10000, 15):

2 01 129

2 03 015

0 22 040 New reference value = -10000

2 03 255

0 22 040 Reported value of “Meteorological residual tidal elevation”

2 01 000

2 03 000

- Gamma radiation dose rate 0 24 014 is intended to be used for reporting of this element under normal conditions, nuclear accidents excluded.

- Photosynthetically active radiation (PAR). Various forms of the electromagnetic energy flux in the 400 – 700 nm wavelength range, either as integrated spectra or using different weighting functions. For example converted to the photosynthetic photon flux (PPF) in quanta per second per square meter, or mole of quanta per second per square meter or microeinsteins per second per square meter.  Approximate conversion is 1 J m-2 s-1 equivalent to 5 μE m-2 s-1 based on a mean wavelength of 550 nm.

The introduced changes are written in blue or marked by side bars.