| WMO/IOC/UNEP/ICSU GLOBAL CLIMATE OBSERVING SYSTEM (GCOS) WMO COMMISSION FOR BASIC SYSTEMS (CBS) WMO COMMISSION FOR INSTRUMENTS AND METHODS OF OBSERVATIONS (CIMO) _________________ GCOS/CBS/CIMO WIGOS Pilot Project on GRUAN Observing Practices and Governance GENEVA, SWITZERLAND, 25 – 27 JANUARY 2012 |  | WIGOS-PP-GRUAN/Doc. 2.2(1) (17.I.2012) ________ ITEM: 2.2 |
| --- | --- | --- |

PROPOSED REVISION OF GRUAN MANUAL

(Submitted by Dr John Nash on behalf of CIMO)

###### SUMMARY

ISSUES TO BE DISCUSSED: This document provides a proposed revision of the GRUAN Manual that aligns the document with CBS and CIMO’s vision for GRUAN, takes into account a WMO desire for the introduction of an operational component to GRUAN, and shapes the document overall for easier adaptation of parts of it for use as input to WMO’s WIGOS Manual and Guide. DECISIONS/ACTIONS REQUIRED: The meeting is requested to consider this input as a CIMO contribution to revision of the GRUAN Manual. REFERENCES: N/A

THE GRUAN MANUAL

Version 1.0.0.7jn

## Purpose of this Manual

The GCOS Reference Upper Air Network (GRUAN) manual describes standardized operating protocols, data reduction, meta-data collection, and data dissemination which shall be employed among participating sites and the GRUAN Lead Centre in order for  the goals of GRUAN to be achieved. These goals are agreed between   GCOS and WMO. The primary goals of GRUAN are to provide vertical profiles of reference measurements suitable for reliably detecting changes in global and regional climate on decadal time scales, initially for temperature and water vapour, with the aim of expanding to other essential climate variables in liaison with other existing scientific networks such as in Global Atmosphere Watch. The measurements will provide a calibrated reference standard for global satellite-based measurements of atmospheric essential climate variables,  will be a reference standard for the measurements of the existing GCOS Upper Air network.

The GRUAN manual establishes mandatory practices and operating protocols (distinguished by use of the words ‘must’ or ‘shall’) having the status of requirements in a technical resolution, which it is necessary that sites within GRUAN follow or implement. It defines the requirements for GRUAN site operations, including requirements on uncertainty and long-term stability. In this way the manual establishes the operational philosophy under which GRUAN shall operate and informs current and future GRUAN sites of the expected modus operandi for GRUAN.

.Relevant information from this manual will be incorporated in the WMO Manual on the Gobal Observing System(WMO-No. 544) and the Guide on the Global Observing System(WMO-No.488). GRUAN Observing systems and operating practices will be referred to in the WMO Guide to Instruments and Methods of Observation (WMO-No.8) and this document will be linked to the relevant scientific documentation for the specialised scientific sounding systems. A GRUAN station may be a scientific observing site  outside of the WMO operational Global Observing System, but the long term observational procedures shall follow the guidelines laid down in this GRUAN manual whether it is a purely scientific site or whether it is already part of the Global Observing System, and the WMO Manuals shall reference  these practices.

The manual defines  the principles that are intended to direct the development of the methods, techniques and processes needed to achieve the stated goals of GRUAN. The initial essential climate variables to addressed in detail have been temperature and water vapour, and the scope of GRUAN will be extended as resources permit to cover a wider range of essential climate variables in liaison with existing scientific networks. Where possible, the document does provide more in-depth detail on specific methodologies , and these shall be  appropriate for incorporation into  or reference by the WMO Guide to Meteorological Instruments and Methods of Observation.

## Table of Contents

Purpose of this Manual	1

Table of Contents	2

Executive Summary	4

1.	INTRODUCTION	11

1.1.	GRUAN heritage	11

1.2.	The purpose of GRUAN	11

1.3.	Organisation and design of GRUAN	12

1.4.	Implementation of GRUAN	13

1.5.	Links to partner networks	14

1.6.	Link to satellite-based measurement programmes	19

2.	REFERENCE MEASUREMENTS	21

2.1.	Terminology	21

2.2.	The concept of a reference measurement	21

2.3.	Managing Change	22

3	MEASUREMENT UNCERTAINTY	27

3.1	Estimating measurement uncertainty	27

3.2	Reporting measurement uncertainty	29

3.3	Reducing measurement uncertainty	29

3.4	Reducing operational uncertainty	29

3.5	Validating measurement uncertainty	30

4	ESSENTIAL CLIMATE VARIABLES MEASURED IN GRUAN	31

4.1	Justification and context for Essential Climate Variables	31

4.2	Priority 1 ECVs	32

4.3	Moving beyond priority 1 variables	36

5	GRUAN SITES	37

5.1	Introduction	37

5.2	Mandatory Operating Protocols	38

5.3	Criteria for Assessing Added Value	39

5.4	The Assessment and Certification Process	42

5.5	Site Auditing	44

6	INSTRUMENTATION	45

6.1	Instrument selection	45

6.2	Measurement redundancy	46

6.3	Surface measurements	46

6.4	Upper-air measurements	47

6.5	Instrument co-location	47

6.6	Calibration, validation and maintenance	47

7	METHODS OF OBSERVATION	50

7.1	Guiding principles	50

7.2	Measurement scheduling	50

7.3	Operation and maintenance, quality standards	56

8	DATA MANAGEMENT	57

8.1	Overview of GRUAN data flow	57

8.2	GRUAN data policy	58

8.3	Collation of meta-data	60

8.4	Data format	60

8.5	Data submission	61

8.6	Data dissemination	61

8.7	Data archiving	62

8.8	Quality control at the instrument/site level	62

9	POST-PROCESSING ANALYSIS AND FEEDBACK	63

10	QUALITY MANAGEMENT	64

ACRONYMS	66

Appendix A – Expanded details on additional GRUAN Essential Climate Variables	67

REFERENCES	70

## Executive Summary

The executive summary as it currently stands is excessively long. It will be condensed in subsequent versions of the manual.

The development and current operation of the GRUAN network is described by a number of distinct but often overlapping documents, including GCOS-112, GCOS-121, GCOS-134, GCOS-140, web-based material, reports from GRUAN task teams and papers published in the international peer reviewed literature. The purpose of this manual is to provide comprehensive documentation of information essential  to the ongoing operation of GRUAN . The high level messages emerging from this GRUAN manual are summarized in this executive summary.

- The purpose of GRUAN

The purpose of GRUAN shall be to:

  - Provide vertical profiles of reference measurements of temperature and water vapour suitable for reliably detecting changes in global and regional climate, on multi-decadal time scales, for all climatically distinct regions of the globe, acting as a reference to the  stability of the observations provided by the operational GCOS Upper Air Network.

  - Provide a calibrated reference standard for global satellite-based measurements of atmospheric essential climate variables, starting with temperature and water vapour.

  - Measure sufficient additional atmospheric and surface variables, necessary to allow satisfactory simulation of the radiances observed by satellite sounding systems, at as many stations as possible..

  - Ensure that any interruptions in satellite-based measurement programmes do not invalidate the long-term climate data record.

  - .

- 2	Organisation and design of GRUAN

  - The GRUAN reference network will operate under the joint governance of GCOS and WMO.

  - Working oversight  of the network will be performed by the GCOS Working group on Atmospheric Reference Observations, incorporating representatives at working level from WMO.

  - A GRUAN Lead Centre agreed by GCOS and WMO, will be responsible for integrating best practices into GRUAN operations, managing the network systems, including data management.

  - GRUAN sites shall use  a designated system of methods, techniques and facilities, implemented for making and archiving  best quality  upper air observations on a global scale. At any site, this system will not be changed without permission from the GRUAN Lead Centre.

  - GRUAN network operations will incorporate an assurance programme to validate the stability and uncertainty of the measurements, agreed with WG-ARO, and managed in detail, by the GRUAN Lead centre

- GRUAN shall also be responsive to the latest technological and scientific progress in measurement techniques and   observational climate requirements. Development work can continue at a site until mature and validated, when it could be introduced into GRUAN operations with the agreement of the GRUAN Lead centre.

- 3.	Implementation of GRUAN

  - 3.1    The implementation of GRUAN shall be guided by the Working Group on Atmospheric Reference observations (WG-ARO)

  - 3.2     Specific issues in support of network design and other decisions shall be performed by GRUAN task teams as agreed by  WG-ARO .These will entrain operational and other relevant expertise in support of GRUAN.

  - A GRUAN Analysis Team for Network Design and Operations Research (GATNDOR) shall undertake focused, short-term research to address specific topics identified by the the WG-ARO. The work will be conducted in coordination with other  relevant GRUAN task teams.

  - The WG-ARO  shall agree on the appropriate method of establishing standard operational procedures for observing systems within GRUAN. This could be a new task team, an instrument mentor, or an existing instrument team within other associated WMO projects/ operational groups.

  - The GRUAN Lead Centre shall identify sites where instrument operators need training, and organise cost-efficient training courses for the network at appropriate locations, as advised by the appropriate task team/instrument mentor, to encourage uniformity of instrument operation between sites.

  - All activities associated with the implementation of GRUAN shall be the responsibility of the countries hosting GRUAN sites and should, as far as possible, be met through national funding.

- Partner networks

  - GRUAN shall not operate in isolation of existing networks but shall collaborate and liaise with existing networks to leverage skills and expertise available from those networks, avoid undue duplication of effort, and ensure that GRUAN data products are tailored to best meet the needs of partner networks.

  - 4.2	The WG-ARO shall work in close coordination with the governing bodies of partner networks, with respect to the development of GRUAN operations. the GRUAN sites where GRUAN operations coincide with other network operations. Specifically, the WG-ARO shall then  identify  methods of improving liaison  with partner networks to ensure close communication between GRUAN and those networks, particularly at expert team level.

  - 4.3	Where possible, GRUAN shall identify, adopt, and extend if necessary, tools and methodologies that have been developed in existing networks that can serve the needs of GRUAN. In particular, data QA/QC procedures developed in existing networks shall, where suitable, be adopted in GRUAN. Where networks are working towards QA/QC procedures, GRUAN should partner with these networks to develop systems that meet the operational requirements of both parties. Where GRUAN develops QA/QC techniques that are superior to those used elsewhere, those techniques shall be shared with partner networks.

  - The GRUAN Lead Centre shall identify methods of access to measurement database from partner networks to enable cross-calibration with GRUAN measurements and to quantitatively link GRUAN measurements to similar measurements made within partner networks.

  - The WG-ARO shall establish close connections to the relevant satellite technical programmes, such as GSICS. WG-ARO members shall be assigned to liaise with key clients within the satellite community to ensure that GRUAN data products are tailored, where possible, to best meet the needs of this community.

  - GRUAN shall be operated in such a way that homogeneity of measurements across the network will ensure that significant site specific differences between GRUAN data and those from satellite-based instruments result are  not from the GRUAN data products.

  - GRUAN shall provide a reference-standard that will serve as a common baseline for splicing satellite-based measurement time series to create CDRs.

  - GRUAN shall collaborate with the WMO SCOPE-CM programme to generate CDRs of upper air ECVs and in this way contribute to Action C10 defined in the GCOS implementation plan (GCOS-92) viz. 'Ensure continuity and over-lap of key satellite sensors ...undertaking reprocessing of all data relevant to climate for inclusion in integrated climate analyses and reanalyses'.

  - The WG-ARO shall maintain active links with the following partner networks:

  - 4.8.1	GUAN/WIGOS

    - GAW (Global Atmospheric Watch

    - NDACC (Network for the Detection of Atmospheric Composition Change)

    - SHADOZ (Southern Hemisphere Additional Ozonesondes)

    - ARM (Atmospheric Radiation Measurement) Programme

    - BSRN (Baseline Station Radiation Network)

    - AERONET (AErosol RObotic NETwork)

- Reference measurements and managing change

  - All GRUAN measurements systems shall make reference quality measurements,  i.e that, at a minimum, the observations are tied to an internationally accepted traceable standard, that the uncertainty on the measurement (including corrections) has been determined, and that the entire measurement procedure and set of processing algorithms are properly documented and accessible

  - For GRUAN measurement systems making vertically resolved measurements, measurement uncertainties shall also be vertically resolved such that each measurement in a profile shall be treated as a single measurement result requiring both the measurement and its uncertainty.

  - The methods by which the measurements are obtained and the data products derived shall be documented for future reference. Meta-data shall be archived that describe how the measurements were made, which corrections were applied, what changes occurred during the observation and post-observation periods to the instruments and the data reduction algorithms.

  - GRUAN shall operate in a way such that changes in instrumentation, changes in operating procedures, changes in data processing algorithms and changes in operators do not introduce unidentified  temporal or spatial discontinuities or biases into GRUAN data products.

  - When changing instrument types, operating procedures, or data processing algorithms, GRUAN shall develop the necessary procedures to fully characterize any systematic biases between the old and new measurement systems so that corrections can be made to the historical data to maintain a long-term homogeneous measurement series. The procedures used  shall be informed by robust scientific investigations including a detailed understanding of the error limitations of the instrumentation in use .

  - Data reduction processes and data archiving within GRUAN shall be designed such that reprocessing of historical data can be easily and quickly conducted i.e. that the original raw data can be efficiently reprocessed, as required, to form a single homogeneous time series.

  - Every reprocessing generating a new homogeneous time series shall be reflected in an increment in the data version and such updates shall be communicated to users who have accessed earlier versions of the data.

  - The GRUAN shall work with CIMO and CBS  to make manufacturers aware of GRUAN’s needs and to understand the constraints on instrument performance that manufactures face.

  - GRUAN shall work with CIMO in fostering  instrument intercomparisons to develop the in-depth understanding required to manage changes from one instrument to another and to inform decisions on the relative advantages and disadvantages of changing instrumentation.

- Measurement uncertainty

  - The three primary steps for managing measurement uncertainty in GRUAN shall be:

- Describe/Analyze all sources of measurement uncertainty, as far as possible.

- Quantify/Synthesize the contribution of each source of uncertainty to the total measurement uncertainty.

- Verify that the derived net uncertainty is a faithful representation of the true uncertainty.

  - All GRUAN observations must be reported in association with an uncertainty budget. The Lead Center shall liaise with the GRUAN site in developing this uncertainty budget, as agreed with relevant instrumentation experts.

  - The uncertainty in the geo-location and time coordinates associated with each measurement shall also be considered when identifying and describing sources of measurement uncertainty.

  - To reduce operational measurement uncertainties, optimal standard operating procedures shall be identified  by the GRUAN Lead Centre and disseminated to all sites making that particular measurement, and adopted where practical, with exceptions clearly documented and justified.

  - The Lead center shall be responsible for ensuring that  estimates of measurement uncertainty are validated through suitable procedures, by nominated instrument experts.

- Site assessment and Certification

  - Site assessment and certification shall fall within the mandate of the Working Group on Atmospheric Reference Observations (WG-ARO) together with the GRUAN Lead Centre .This shall be consistent with the guidance developed  and standards agreed with the WMO Commission for Instruments and Methods of Observation (CIMO; WMO-No. 8) and the WMO Commission for Basic Systems (CBS; WMO-No. 488) .

  - Sites shall propose specific measurement programmes for inclusion in GRUAN and it is these that will be required to conform to the operating protocols defined in this document, and will be agreed as appropriate for GRUAN by the WG-ARO..

  - The WG-ARO shall exercise its discretion in evaluating the proposed inclusion of a site against the criteria defined in this document.

  - To identify potential issues early,  the quality of the data received from the sites shall be reviewed annually.

  - Where site reassessments identify measurement programmes that consistently fall short of GRUAN operating standards, GRUAN certification of that programme shall be suspended.

  - If all measurement programmes at a site lose their GRUAN certification the site shall be suspended from the GRUAN network.

GRUAN is a heterogeneous network that includes sites from both the research community and the operational meteorological community. The mandatory requirements for sites reflect GRUAN’s primary goal of providing reference quality observations of the atmospheric column while accommodating the diverse capabilities of sites within the network. GRUAN sites shall:

  - Provide reference quality observations, including uncertainty estimates for each datum. Profile measurements require uncertainty estimates for each measurement point on the profile

  - Provide access to raw data and assure long-term storage of the raw data either at the site, another GRUAN facility, or at another internationally accessible archive.

  - Provide complete metadata for each measurement.

  - Perform regular traceable pre -launch ground checks for balloon borne systems and record the results. Other  instruments which provide vertical profiles extending from the surface require regular checks to assure correct operation..

  - Work with the GRUAN Lead Centre to perform checks as requested  to validate the derivation of the uncertainty on the primary measurement.

  - Provide annual reports summarizing GRUAN operations at the site.

  - Conduct measurement programmes with an operational philosophy of continually striving to sustain the measurement quality at a given level. If improvements to measurement accuracy can be obtained these need to be documented and the introduction agreed with the GRUAN Lead Centre.

  - Manage changes in instrumentation, operating procedures, and processing algorithms  pro-actively to avoid the introduction of spatial or temporal biases in GRUAN data products.

  - Actively communicate with the Lead Centre.

Once a site has committed to operating a set of measurement programmes under these protocols the added value that a site brings to the GRUAN network will be a function of:

  - The extent to which a site can fulfil the measurement programmes expected of a GRUAN site as defined above.

  - The extent to which the site measurement programmes provide measurements in regions, or of atmospheric phenomena, which were not previously sampled.

  - The extent to which a site brings unique observational and/or analysis capabilities to the network as a whole and the likelihood of being able to propagate those capabilities across other sites in the network.

  - The availability of historical measurements that conform to the GRUAN standard.

  - The extent to which a site can commit to a multi-decade programme of measurements.

  - The extent to which a site can provide redundant observations of the priority 1 variables.

  - The extent to which a site is capable of measuring other ECVs identified in GCOS-112 as being desired quantities.

  - The level of institutional support for the site and commitment to maintaining long-term reference quality measurement programmes.

  - The level of institutional support for the site (and any partner institutions) to undertake fundamental scientific research of the measurements from the site and other GRUAN sites.

Certification of GRUAN sites will not be a single event. Periodic (e.g. every 3-4 years) complete auditing of the measurement programmes included in the GRUAN certification for a site shall be conducted to ensure that the programmes continue to meet GRUAN standards.

- Instrumentation

  - . New equipment shall not be introduced to GRUAN operations  without the permission of the Lead Centre. If new instrumentation is proposed by a manufacturer/ scientific institution the performance of this system shall be reviewed by appropriate instrument experts and a case for incorporation in GRUAN  then submitted to WG-ARO/ Lead Centre/CBS/CIMO.

  - To the extent possible, the number of different instrument types employed across the network shall be minimized to encourage homogeneity of data products and to reduce administrative overhead. However, in the case of radiosonde consumables it is essential to use radiosonde from more than one manufacturer to minimise the fluctuation in the consumable production quality/ manufacturing changes  /weaknesses in the radiosondes.

  - A number of factors shall guide the selection of instruments for use in GRUAN including, but not limited to, instrument heritage, sustainability of measurement systems, robustness of uncertainty, information content, manufacturer support, and site location.

  - Redundancy in core measurement systems at sites shall be used to validate derived measurement uncertainties, and also to detect changes in the systematic bias between the observing systems. The cross-checking of redundant measurements for consistency shall be an essential part of the GRUAN quality assurance procedures, but may not be required at all sites, because of logistical limitations.

  - Independent measurements of the same (or related) variables shall be reported in a consistent way.

  - Where a cluster of instruments at different locations is operating as a single GRUAN site, robust scientific analysis of the resultant increase in apparent measurement uncertainty shall be undertaken to clearly document how the lack of co-location contributes to the net measurement uncertainty. It must be very clearly indicated which observations are made from each location, and the data should not be merged as though it were all from one location.

  - Operational procedures should be harmonised across the GRUAN network so that a measurement of a given parameter at one site is comparable as far as the stability of the equipment allows  to a measurement of the same variable at a different site. When two identical instruments are deployed at two different sites, they shall also use the same calibration procedures.  {Ed. Note: This does not ensure all the measurements are the same as the equipment may operate differently in different conditions. Where there are 300 sites using the same equipment around the world using one set of software , shouldn’t  GRUAN use the same if it is supposed to be the backbone of the networks???}

  - GRUAN sites shall maintain a “GRUAN site working standard” for each basis unit, e.g. a thermometer periodically calibrated to a National Metrology Institute or other accredited agency standard to ensure traceability to an SI standard.

  - The Lead Centre shall implement a mechanism to address the compatibility with the rest of the network of those systems not traceable to SI standards.

  - Each station shall maintain accurate meta-data records and provide these to the GRUAN archives. Copies of calibration certificates shall be submitted to the GRUAN meta-database.

- Measurement Scheduling

9.9.1	Core radiosonde measurements.

In order to sustain good quality, and provide a backbone to GUAN, a programme of regular high quality radiosonde measurements needs to be maintained throughout the year. The nature of this programme needs to be negotiated with the GRUAN Lead centre, and will depend on the availability of suitably skilled staff. The minimum programme should be one radiosonde observation per day at least 5 times per week. However, the preferred option is two radiosondes per day continuously. For temperature the most reliable measurements can be obtained in the dark, and this is not usually  a problem at most sites regularly performing radiosonde observations .The programme of operation of remote sensing system at a site  needs to be agreed with the Lead Centre and those sites which are included in GRUAN because of the deployment of special scientific systems( balloon borne, or remotely sensed), may be allowed to fly lower numbers  of radiosondes than those sites making core radiosonde measurements.

Selected  GRUAN sites are to fly one special scientific sounding system for measuring stratospheric water vapour once per month. Where several GRUAN sites are located close together in the same latitude band there is no point in all making this expensive measurement and the individual site flight schedules should take this into account

Schedules of twin radiosonde flights for assurance purposes are to be agreed with the GRUAN Lead centre. E.g. the schedules for a tropical site may be different for those at midlatitides and there will be a limit on the number of cross-checking flights required in mid-latitudes, given the need to conserve limited  financial resouces .

- Data Management

GRUAN data circulation shall be based on the use of the WMO WIS system. Similarly access to the GRUAN data base shall be through the WIS system.

- Post Processing Analysis and Feedback

This section of the executive summary has not yet been written.

- Quality Management

ARM quality radiosonde measurements shall be subject to the same quality evaluation procedures against numerical weather prediction fields as is applied to GUAN measurements.

This section of the executive summary has not yet been written.

# INTRODUCTION

## GRUAN heritage

The need for a reference upper-air network to better meet the needs of the international climate research community has long been recognized (Trenberth, 2003). In response to this need, the inception of the GCOS Reference Upper-Air Network (GRUAN; GCOS-112, GCOS-134) was formalized between 2005 and 2007 when a reference upper-air network, envisaged to eventually include 30-40 sites, worldwide was planned. In contrast to the GCOS Upper Air Network (GUAN), which is based on weather observing stations, GRUAN is specifically designed for climate research. Therefore, rather than being a purely operational network like GUAN, it is a network that serves the international climate community through a combination of research and operational sites, giving high quality operational network observations and elements of research and development for the future. GRUAN provides reference observations of upper-air essential climate variables (ECVs), through a combination of in situ measurements made from balloon-borne instruments and from ground-based remotely sensed observations. Furthermore, management decisions in GRUAN are driven by a variety of climate  requirements for long-term time series of measurements of assured measurement stability, but also by the need for good operational practices to ensure stability in the measurements.  So, on one hand GRUAN is partly  a research network constantly striving to improve measurement techniques, and quantify and reduce measurement uncertainties by improving precision and accuracy, but on the other hand the network measurements need to be made in a stable way over multi-decadal time scales to achieve data homogeneity in time and spatially between measurement stations..    These two aspects of GRUAN operations are not mutually exclusive, but do need to be carefully balanced. This dual-purpose nature of GRUAN has been accommodated in this manual.

## The purpose of GRUAN

The purpose of GRUAN shall be to:

- Provide vertical profiles of reference measurements of temperature and water [and additional essential climate variables], suitable for reliably detecting changes in global and regional climate, on multi-decadal time scales, for all climatically distinct regions of the globe. The uniformity and coherence of standard operating procedures at GRUAN stations and the resultant homogeneity of GRUAN climate data records not only provides a global reference standard for operational upper-air network stations, but improves the detection of changes in the climate of the troposphere and stratosphere.

- Provide a calibrated reference standard for global satellite-based measurements of atmospheric ECVs. This facilitates the creation of seamless, stable, and long-term databases of satellite-based measurements suitable for detection of trends and variability in climate in the upper troposphere and stratosphere on all time scales.

- Fully characterize the properties of the atmospheric column. This is necessary for process understanding and for radiative transfer modelling.???

- Ensure that any interruptions in satellite-based measurement programmes do not invalidate the long-term climate data record.

GRUAN shall also provide radiosonde observations in real-time for incorporation in meteorological analysis in order to fulfil the requirement of providing a reference to the operational observations.

In achieving the four goals stated above, GRUAN will address some of the current deficiencies of the GUAN network. The reliable detection of the vertical structure of changes in climate variables in the atmosphere requires high quality atmospheric observations with well characterised measurement uncertainties. GUAN provides upper air measurements over large regions of the globe, using radiosondes that are similar to those used in GRUAN. However GUAN sites seldom include additional systems to validate data stability, and rely on the assumption of stability in the radiosonde quality with time. If GRUAN can identify the changes that occur in production consumables, this will benefitthose using  GUAN measurements.

In the context of the other WMO observing systems, GRUAN will need to be the climate reference backbone of the existing global operational upper-air network. As noted in GCOS-112, GRUAN sites need not necessarily be current GUAN sites. Because GUAN sites often operate with different equipment, sensors, and operating protocols, the different requirements of GRUAN and GUAN operations may require careful management. The envisaged capabilities of a fully-implemented GRUAN are detailed in GCOS-112. The scientific justification and requirements for GRUAN are summarized in Section 3 of GCOS-112 and in Seidel et al. (2009) and are not repeated here. Continued implementation of GRUAN is specifically called for under Action A16 of the 2010 update to the implementation plan for GCOS (GCOS-138).

## Organisation and design of GRUAN{Ed Note: this needs to be replaced}

Figure 1: Schematic outline of the structure of GRUAN. GRUAN elements are shown in red while external support structures are shown in black. GATNDOR=GRUAN Analysis Team for Network Design and

The GRUAN reference network will operate under the joint governance of GCOS and WMO. A schematic outline of the GRUAN governance structure is given in Figure 1.

  - Working oversight  of the GRUAN network will be performed by the GCOS Working group on Atmospheric Reference Observations(WG-ARO), including representatives at working level from WMO. GCOS and WMO will select those groups ( e.g. GCOS/WCRP Atmospheric Observations Panel, or WMO Technical Commission working groups/experts) through which WG-ARO will report.

A GRUAN Lead Centre agreed by GCOS and WMO, will be responsible for integrating best practices into GRUAN operations and managing the GRUAN network systems, including data management. This Lead Centre is currently hosted by the German Weather Service at the Lindenberg Meteorological Observatory in Germany. The GRUAN Lead Centre acts as the interface between GRUAN and the community of users of GRUAN products. For example, data transfer to end-users is not made from GRUAN measurement sites but is first shared within the GRUAN community, subjected to the QA/QC procedures developed within GRUAN (Section 10), and then submitted by the GRUAN Lead Centre to the GRUAN data repository (NCDC, Section 8.6).

  - GRUAN sites shall use a designated system of methods, techniques and facilities, implemented for making and archiving best quality upper air observations on a global scale. At any site, this system will not be changed without permission from the GRUAN Lead Centre. GRUAN operations shall integrate where possible and when feasible with other international climate monitoring programmes.

  - GRUAN network operations will incorporate an assurance programme to validate the stability and uncertainty of the measurements, agreed with WG-ARO, and managed in detail, by the GRUAN Lead centre

GRUAN shall also be responsive to the latest technological and scientific progress in measurement techniques and   observational climate requirements. Development work can continue at a site until mature and validated, when it should be introduced into GRUAN operations with the agreement of the GRUAN Lead centre. The design of GRUAN shall recognise the heterogeneity of the network of sites, many of which will have primary responsibility to networks other than GRUAN.

## Implementation of GRUAN

The implementation of GRUAN shall be guided by the Working Group on Atmospheric Reference observations (WG-ARO).

Specific issues to be investigated in support of GRUAN implementation shall be performed by GRUAN task teams set up by  WG-ARO .These will entrain operational and other relevant expertise in support of GRUAN and will work in coordination with the GRUAN Lead Centre.

A GRUAN Analysis Team for Network Design and Operations Research (GATNDOR) shall undertake focused, short-term research to address specific topics identified by the WG-ARO. The work will be conducted in coordination with other  relevant GRUAN task teams. GATNDOR activities shall be coordinated with the GRUAN task teams and with national GCOS programmes when appropriate.

The WG-ARO  shall agree on the appropriate method of establishing standard operational procedures for all observing systems within GRUAN. This could be a new task team, including investigations at the GRUAN Lead centre, an instrument mentor, or an existing instrument team within other associated WMO projects/ operational groups. The task teams shall evaluate the appropriateness of uncertainty estimates, the usefulness of particular measurements and operational procedures, synthesize the available knowledge and develop recommendations to improve GRUAN measurements and operations. These task teams shall confer regularly to evaluate the current status of GRUAN observations, to identify weaknesses, and to incorporate new scientific understanding into GRUAN. The expertise of these teams shall also be used to support the Lead Centre in guiding individual stations through instrumental and operational changes without impacting long-term measurement time series.

  - The GRUAN Lead Centre shall identify sites where instrument operators need training, and organise cost-efficient training courses for the network at appropriate locations, as advised by the appropriate task team/instrument mentor, to encourage uniformity of instrument operation between sites.

All activities associated with the implementation of GRUAN are the responsibility of the countries hosting GRUAN sites and should, as far as possible, be met through national funding.

To best serve the needs of climate monitoring and research, it is essential that GRUAN be informed by a good understanding of the evolving science issues that drive the measurements and accuracy of the GRUAN data. Therefore, the establishment of an internal or external science advisory panel should be considered (GCOS-134)

The instrumentation deployed and the observing schedules may differ between sites , as agreed with WG-ARO, but the methods of observations used with the main observing systems, i.e. radiosondes and balloon borne Specialised Sounding Instruments shall be uniform between all the GRUAN sites.

## Links to partner networks

The purpose of this section is to provide, as early as possible in this document, a context for GRUAN in the broader community of climate monitoring networks. For instance, in the charter for GRUAN (GCOS-92) it is stated that ‘where feasible, these reference sites should be co-located and consolidated with other climate monitoring instrumentation’,

GRUAN shall not operate in isolation of existing networks and GRUAN is not intended to replace in any way existing networks. Many GRUAN initial and candidate sites already belong to existing networks such as GUAN, GAW, NDACC, BSRN and SHADOZ. One of the essential characteristics of a successful GRUAN is close coordination with the user community and many of these networks are also likely to be users of GRUAN data. Similarly, complementary measurements from these other networks should be collated in a database to enable cross-calibration and to quantitatively link GRUAN measurements to similar measurements made within other networks. As a result, close coordination between the governing bodies of these networks and with the WG-ARO is required on a continuous basis. This close coordination can be achieved by having members of the WG-ARO attend steering group meetings of partner networks and by inviting co-chairs from partner networks to attend WG-ARO meetings.

There is a wide range of tools and methodologies that have been developed in existing networks that GRUAN can adopt, extend if necessary, and learn from. Similarly, existing networks will have skills and expertise likely to be useful to GRUAN and its operations. As a result, contact with expert teams from existing networks shall be made by WG-ARO, GRUAN task teams, and GATNDOR to support GRUAN operations and to avoid duplication of effort by utilizing existing scientific knowledge.

A number of networks currently in operation make measurements which fall within the scope of GRUAN. For instance, those stations that make upper air measurements that are not part of the typical meteorological measurements of temperature, pressure and humidity. Many of these networks have developed systems for assuring the quality of their measurements, including GUAN. Where the systems currently in place are sufficient to meet the operational requirements of GRUAN, they should be used by GRUAN. Where networks are working towards QA/QC procedures, GRUAN should partner with these networks to develop systems that meet the operational requirements of both parties. In some cases, sites within these partner networks may also become GRUAN sites. This is encouraged since it facilitates a traceable link between GRUAN measurements and measurements made at all other sites within the partner network (assuming that the measurements within the partner network are cross-calibrated and can be quantitatively linked).

Existing networks and potential resources from within those networks likely to be of value to GRUAN are discussed below.

### GAW (Global Atmospheric Watch)

The GAW programme of WMO is a partnership involving 80 countries, providing reliable scientific data and information on the chemical composition of the atmosphere, and the natural and anthropogenic drivers of changes in chemical composition. In this way, GAW improves understanding of the interactions between the atmosphere, the oceans and the biosphere. GAW has strong linkages to GCOS and so is likely to have skills and resources that could be used to support GRUAN.

### NDACC (Network for the Detection of Atmospheric Composition Change)

NDACC reports to GAW. The NDACC comprises more than 70 remote-sensing research stations for observing and understanding the physical and chemical state of the stratosphere and upper troposphere and for assessing the impact of stratospheric changes on the underlying troposphere and on global climate. Primarily based on lidar measurement techniques. NDACC incorporates 5 water vapour measurement sites and maximum of 8 temperature measurement sites. There are a number of key differences between NDACC and GRUAN that require GRUAN to operate as a new and independent network, including:

- NDACC aims to observe and understand the chemical composition of the stratosphere and upper troposphere. For GRUAN the highest priority observations are the atmospheric state variables of temperature, pressure and humidity.

- The primary focus of NDACC is on ozone and the chemicals responsible for ozone depletion. The primary focus of GRUAN is on climate and the factors driving changes in climate.

- NDACC operates as a federation of independent measurement sites. NDACC does have in place stringent standards which must be met for measurement programmes to become part of the network. However, large numbers of  balloon borne measurements in GRUAN requires by a Lead Centre that aims to implement a minimum set of standard operating procedures across the network as a whole.

There are, however, a number of measurements and operational procedures common to both networks and every effort should be made to avoid duplication of effort and to ensure that the lessons learned within NDACC are assimilated into GRUAN. For example:

- The NDACC has established 'working groups' that are primarily centred on specific instruments used within the NDACC. GRUAN task teams currently include a mix of teams focussing on specific measurements systems (radiosondes and precipitable water from GNSS) and on network wide operational issues. As more measurement systems are incorporated into GRUAN operations, consideration should be given to later expanding the ‘Ancillary Measurements’ Task Team to include specific measurement systems in addition to the 'cross-cutting' task teams that focus on issues common to the network as a whole. This could be achieved through assigning ‘instrument mentors’ as recommended in GCOS-112 and utilized operationally in the ACRF programme. Task teams focussing on specific measurement systems or on specific ECVs would better link to advisory groups within partner networks e.g. the Scientific Advisory Groups within GAW (see Section 1.5.4). SCOPE-CM (see Section 1.6) intends to establish one or two centres to lead the generation and provision of fundamental climate data records for each ECV and so establishing task teams within GRUAN focussed on specific ECVs or groups of ECVs would mirror the structure within SCOPE-CM and thereby facilitate interactions with the satellite-based measurement community (one of the key clients of GRUAN).

- Measurements of vertical ozone and water vapour profiles made within NDACC will be common to measurements made within GRUAN. This includes both balloon-sonde and lidar measurements.

- Techniques have been developed within NDACC to manage changes in instrumentation. GRUAN should build off the expertise developed in this community over the past two decades e.g.

- The JOSIE ozonesonde intercomparisons (Smit et al., 2007).

- Regional ozone profile intercomparisons from multiple instruments (McDermid et al., 1998a; McDermid et al., 1998b).

- Intercomparisons of vertical water vapour profile measurements.

- Measurement redundancy in the NDACC network sites has been a strength of the network since it allows intercomparisons of supposedly identical measurements by different instruments which often highlight previously unknown deficiencies in the measurements (Brinksma et al., 2000). GRUAN will include similar measurement redundancy (see Section 6.2).

### BSRN (Baseline Station Radiation Network)

The BSRN provides a worldwide network to continuously measure radiative fluxes at the Earth's surface. The network comprises about 40 stations between 80°N and 90°S many of which began operation in 1992 and each year more stations are added to the network. These stations provide data for the calibration of measurements made within the GEWEX Surface Radiation Budget (SRB) Project and other satellite-based measurements of radiative fluxes. BSRN data are also used to validate radiative flux models. BSRN data are archived at the Alfred Wegener Institute (AWI) in Bremerhaven, Germany. In 2004, BSRN was designated as the global surface radiation network for the GCOS. The BSRN stations also contribute to GAW (see Section 1.5.4).

The primary goal of BSRN is to monitor the shortwave and longwave radiative components and their changes with the best methods currently available. Therefore the measurements of longwave and shortwave incoming and outgoing radiation within GRUAN will overlap with the measurements made within BSRN. Access to the BSRN calibration facilities at the Physikalisch-Meteorologisches Observatorium Davos (PMOD)/World Radiation Centre (WRC) would be highly advantageous to GRUAN. The BSRN includes a working group on measurement uncertainties that could be used to provide guidance for establishing the radiation measurement uncertainties within GRUAN.

### WOUDC (World Ozone and UV Data Centre)

The WOUDC is one of the World Data Centres which are part of the GAW (see Section 1.5.4) programme of WMO. The WOUDC, operated by the Experimental Studies Section of Environment Canada in Toronto, is not so much a network as an international repository for ozone and UV data. There are many practices employed within the ozone measurement community that are likely to be useful to GRUAN. For example, the management of the Dobson Spectrophotometer and Brewer Spectroradiometer networks, both of which provide data to the WOUDC, demonstrate many of the principles that form the foundation for GRUAN. These include:

- Undertaking regular regional intercomparisons of instruments which always include a travelling standard which facilitates standardization of instrument performance between regions.

- Archiving of raw data to permit later reprocessing should new improved ancillary data become available e.g. the shift to the Bass and Paur ozone absorption cross-sections in the late 1980s. A similar process is now underway to evaluate a possible change from the Bass and Paur cross-sections to e.g. the Daumont (Daumont et al., 1992) cross sections.

- Careful QA/QC of data before archiving and strict version control of data submitted to international archives.

These principles have resulted in ground-based total column ozone time series of sufficient quality to allow detection of the multi-decade decline in ozone until the end of the 20th century and the onset of ozone increases thereafter.

### SHADOZ (Southern Hemisphere Additional Ozonesondes)

The SHADOZ project was initiated to remedy the lack of consistent tropical ozonesonde observations. This was done by increasing the frequency, and improving the quality, of ozonesonde launches at selected tropical ozone observing stations (Thompson et al., 2003). Rather than establishing an entirely new network, SHADOZ aims to enhance ozonesonde launches at existing facilities on a cost-share basis with international partners. The geographical coverage of the network was specifically designed to address target research questions.

### AERONET

AERONET (AErosol RObotic NETwork) is a federation of ground-based remote sensing aerosol networks with contributions from national agencies, institutes, universities, individual scientists, and research partners. The programme provides a long-term, continuous and publically accessible database of aerosol optical, microphysical and radiative properties. The standardization of instruments, calibration procedures, and data processing and distribution is well aligned with the needs of GRUAN.

The AERONET programme provides globally distributed observations of spectral aerosol optical depth (AOD), products derived from the raw measurements, and precipitable water in diverse aerosol regimes. Aerosol optical depth data are computed for three data quality levels: Level 1.0 (unscreened), Level 1.5 (cloud-screened), and Level 2.0 (cloud-screened and quality-assured). It is primarily the level 2.0 data that are likely to be of interest to GRUAN since these data are quality-assured. Inversions, precipitable water, and other AOD-dependent products are derived from these levels and may implement additional quality checks.

### Atmospheric Radiation Measurement (ARM) Programme

The goal of the Department of Energy ARM programme is to study alterations in climate, land productivity, oceans or other water resources, atmospheric chemistry, and ecological systems that may alter the capacity of the Earth to sustain life. This includes improving the atmospheric data sets used in regional and global climate models. A primary objective of the ARM user facility is improved scientific understanding of the fundamental physics related to interactions between clouds and radiative feedback processes in the atmosphere.

Four of the 15 initial GRUAN sites are also ARM sites in part because the radiation measurements made at these sites satisfy many of the ECV measurement requirements within GRUAN. The dedicated Data Quality (DQ) Office which ARM established in July 2000 to coordinate and implement efforts to ensure the quality of the data collected by ARM field instrumentation will likely provide a number of tools which could be implemented across the GRUAN network to ensure the quality and network homogeneity of the radiation measurements. The DQ Office has the responsibility for ensuring that quality control results are communicated to data users so that they may make informed decisions when using the data, and to ARM's Site Operations and Engineers to facilitate improved instrument performance and thereby minimize the amount of unacceptable data collected. The ARM DQ Office has developed a suite of sophisticated data quality visualisation tools that are likely to be of interest to GRUAN.

Another ARM organizational structure that is likely to be relevant for GRUAN is the assignment of instrument mentors. Because GRUAN task teams are not structured by instrument (as is the case for NDACC where each working group focuses on one instrument), having ARM-type instrument mentors that advise on instrument operation, maintenance and calibration across the network as a whole may be beneficial. Instrument mentors have an excellent understanding of in situ and remote-sensing instrumentation theory and operation and have comprehensive knowledge of the scientific questions being addressed with the measurements made. They also possess the technical and analytical skills to develop new data retrievals that provide innovative approaches for creating research-quality data sets.

### Partnership with Meteorological agencies

Meteorological agencies producing global real-time analyses or historical reanalyses (e.g. NCEP/NCAR, ECMWF, JMA, Met Office, DWD  and NASA) are likely to be users of the high quality data produced by GRUAN. Well developed systems for monitoring the quality of operational observations ( whether it is the performance of individual radiosonde stations or the bias corrections required by current satellite observations In turn, diagnostics obtained from the 4D-Var assimilation schemes used in such activities will provide information on the consistency of the GRUAN measurements with other data used in the operational analyses (thereby facilitating easier comparisons of GRUAN measurements with e.g. satellite-based measurements) and on the representativeness of the uncertainty estimates on the GRUAN data. If the GRUAN data are to be used in the 4D-Var assimilation schemes used by the reanalysis centres, it is essential that the precise 4D (latitude, longitude, altitude and time) coordinates associated with any measurement are available (see Section 8.3). Reference sites will prove essential for helping to characterize observational biases and the impact of observing system changes, as well as to understand model errors, all of which are important aspects in creating high-quality reanalyses (Schubert et al., 2006).

## Link to satellite-based measurement programmes

GRUAN will provide data sets, not currently available, that will be useful to the satellite measurement community for calibrating and validating satellite-based sensors, for providing input to radiative transfer calculations used in satellite-based measurement retrievals, and for removing offsets and drifts between satellite-based data sets when creating merged data products. Because the GRUAN measurements are likely to serve a wide range of end-users within the satellite measurement community, WG-ARO members shall be assigned to liaise with key clients within the satellite community, and with other data providers (e.g. the Radiation Panel within GEWEX), to ensure that GRUAN data products are tailored, where possible, to best meet the needs of this community. Once GRUAN datasets are available, pilot studies on enhanced datasets using these reference measurements need to be undertaken.

### Forward modelling for satellite-based measurement retrievals

### {Ed Note: This retrieval process for temperature and humidity is not required by numerical models, so probably should not be used in any argument here.]

Satellite-based measurements of atmospheric parameters often rely on an optimal estimation approach (e.g. Rodgers 2000) to derive profile or slant column density information of these parameters. Optimal estimation employs a forward model that is used to simulate the radiance field that a satellite-based sensor would sample for a given state of the atmosphere. To determine a state vector (the true values of the atmospheric parameters of interest), together with uncertainties, from the observed satellite-based radiance measurement, typically in the form of a spectrum, the forward model calculations need to be inverted. Such an inversion is typically poorly constrained, i.e. does not have a unique solution, and, as a result, known a priori (background) information about the variables to be retrieved is usually required as input to the forward model. GRUAN measurements may provide such a priori information. Furthermore, GRUAN measurements of atmospheric state variables such as temperature, pressure and humidity that partially define the radiative transfer properties of the atmosphere, and which are required as input to the forward model, can significantly reduce the uncertainties on other retrieved atmospheric parameters.

### Calibration and validation of satellite-based sensors

Ground-based reference profile measurements may also provide an independent standard against which the satellite retrievals may be validated. For example, Vömel et al. (2007a) demonstrate how reference-quality in situ water vapour measurements can be used to validate satellite-based observations of stratospheric water vapour. In addition to validating retrieved products, satellite radiances require calibration against a ground truth to unambiguously remove biases (Ohring et al., 2005) in order to be useful for climate monitoring. GRUAN and the GSICS are complementary in meeting this need.

Global Positioning System Radio Occultation (GPS-RO) measurements are in use  for upper-tropospheric and lower-stratospheric temperature, and validated by comparison with numerical weather prediction fields. To this end GRUAN will also provide shorter-term quality assured measurements for the validation of satellite-based retrievals.

The need for inter-station homogeneity within GRUAN has special significance for validation of satellite-based measurements. Ground-based measurements made at all GRUAN stations shall be made in similar fashion, so that differences in the soundings of temperature and water vapourbetween GRUAN sites are as small as possible . If this is achieved, differences in collocation biases  between GRUAN measurement and satellite radiance will then primarily be a function of systematic bias in the satellite radiance or  caused by a difference between sites in other conditions, e.g. thin clouds in the satellite field of view, surface emissivity, etc..

The issue of measurement scheduling within GRUAN to accommodate satellite validation activities is discussed further in Section 7.2.

### Creating global homogeneous atmospheric climate data records

While satellite-based measurements have the advantage of providing global or near-global geographical coverage, the quality and usefulness of the measurements is compromised by an inability to conduct regular calibrations, limited vertical resolution, difficulties in continuity due to drifting orbits (which, for species showing strong diurnal variation can alias into apparent trends), and limited instrument lifetimes which require data series from multiple instruments to be spliced together to form long-term data records. Discontinuities between satellite-based measurements of climate variables, can be ruinous for detecting variability and long-term changes in climate. The reference measurements that GRUAN will produce can be used to remove offsets and drifts between these separate satellite-based measurement series within the limitations imposed by the uncertainties on the GRUAN measurements. In this way GRUAN shall provide a reference-standard that will serve as a common baseline when splicing satellite-based measurement time series. Specifically, differences between a given satellite-based data set and the GRUAN reference-standard can be analyzed using the algorithms detailed in Alexandersson et al. (1997) and Khaliq et al. (2007) to automatically detect steps and drifts in the differences. The underlying systematic structure in such differences can then be used to homogenize the satellite-based measurements with the GRUAN reference-standard. Similar approaches using the global ground-based Dosbon and Brewer spectrophotometer networks to create long-term global total column ozone records from multiple satellite-based measurements have been developed (Bodeker et al., 2001) as have techniques for homogenizing radiosonde temperature time series (Haimberger, 2007).

By contributing to the creation of global homogeneous trace gas data bases, GRUAN will connect to the WMO SCOPE-CM (Sustained, Co-Ordinated Processing of Environmental Satellite Data for Climate Monitoring) programme. The aim of SCOPE-CM is to establish a network of facilities ensuring continuous and sustained provision of high-quality satellite products related to ECVs, on a global scale, responding to the requirements of GCOS. GRUAN and SCOPE-CM shall collaboratively contribute to Action C10 defined in the GCOS implementation plan (GCOS-92) viz. 'Ensure continuity and over-lap of key satellite sensors ...undertaking reprocessing of all data relevant to climate for inclusion in integrated climate analyses and reanalyses' (Action C8 in the 2010 update of the GCOS implementation plan; GCOS-138).

# REFERENCE MEASUREMENTS

## Terminology

The following terminology is used throughout this manual to describe the uncertainty components of a reference measurement:

Measurement uncertainty: Every measurement has imperfections that give rise to some uncertainty in the result, hereafter referred to as the measurement uncertainty. Traditionally, a measurement uncertainty is viewed as having two components, a random component and a systematic component.

Random error: The random error component of any measurement is the result of stochastic variation in quantities that influence that measurement. While random errors cannot be designed out of a system, the random error on the mean of multiple measurements is reduced since, by definition, the expected value for the random error is zero. The term ‘random error’ is preferred over the term ‘precision’ since precision is often used to designate the number of bits or significant digits to which a value is specified.

Bias: The bias represents the difference between the measured value and the true value. It results from systematic errors that do not average to zero as the number of measurements increases. However, if biases can be identified and quantified, they can be corrected for. The term ‘bias’ is preferred over the term ‘accuracy’ since it denotes more clearly that the deviation is systematically in one direction.

Stability: Stability refers to the consistency of random errors and biases with time. Undetected trends in biases induce artificial trends in measurement time series.

## The concept of a reference measurement

As denoted by its title, the primary objective of GRUAN is to provide reference measurements for a range of upper-air climate variables. Reference quality atmospheric observations are based on key concepts in metrology (measurement science), in particular traceability. Metrological traceability is the process whereby a measurement result, i.e. a measurement and its error, can be related to a reference through a documented, unbroken chain of calibrations, each of which contributes to the measurement error.

A reference measurement does not refer to a measurement that is perfect, nor to a measurement that will never change. Rather it refers to our current best estimate of the value for some atmospheric parameter, as well as a best estimate for the level of confidence that is associated with this value, recognising that future improvements in measurement techniques and/or reprocessing following new knowledge may lead to refinements in that reference value. In most cases it will be the best technology available that will achieve the best estimate of the value for some target atmospheric parameter. Reference measurements accommodate the unavoidable sources of uncertainty in the compilation of the net measurement uncertainty while excluding those sources of uncertainty that can be avoided. For example, in the pre-deployment calibration of a sensor, there will be some unavoidable uncertainty in the accepted measurement standard and hence some unavoidable uncertainty in the calibration which must then be included in the net measurement uncertainty. However, contributions to measurement uncertainty from e.g. an improperly documented traceability chain, proprietary methods, appeal to physical principles without experimental verification, or the use of an improper calibration standard must be avoided. Similarly, when the instrument is later deployed, there will be numerous, unavoidable contributions to the total measurement uncertainty from e.g. uncertainty in the input data, data processing constants, the data retrieval algorithm, and in the physical/chemical model of the measurement system used to convert raw measurements into data. However, contributions to measurement uncertainty from the use of ‘black box’ software, undocumented or unvalidated measurement adjustments, or the disregard of known biases must be avoided.

A reference measurement may not necessarily be the outcome of a measurement using a single instrument but may be an average of measurements from one instrument or an average of results from multiple instruments. This highlights the importance of measurement redundancy (see Section 6.2) in that access to coincident multiple measurements of the same quantity often leads to a more robust estimate of the true value and a better estimate of the uncertainty on that value.

The estimate for the level of confidence on any measurement is expressed as the measurement uncertainty and is a property of the measurement that combines instrumental as well as methodological uncertainties. The measurement uncertainty describes the current best knowledge of instrument performance under the conditions encountered during an observation, it describes the factors impacting a measurement as a result of operational procedures, and it makes all factors that contribute to a measurement traceable. Within GRUAN this uncertainty shall be vertically resolved and each measurement in a profile shall be treated as a single measurement result requiring both the measurement and its uncertainty. To provide the best estimate for the instrumental uncertainty, a detailed understanding of the instrumentation is required for the conditions under which it is used. Specific requirements that an observation must fulfil to serve as a reference for calibrating or validating other systems, have been defined in Immler et al. (2010).

A reference measurement typically results from a measurement procedure that provides sufficient confidence in its results by relating to well-founded physical or chemical principles, or a measurement standard that is calibrated to a recognized standard, in general a standard provided by a National Metrological Institute (NMI). For GRUAN, a reference measurement is one where the uncertainty of the calibration and the measurement itself is carefully assessed. This includes the requirement that all known biases have been identified and corrected, and, furthermore, that the uncertainty on these bias corrections has also been determined and reported. An additional requirement for a reference measurement is that the measurement method and associated uncertainties should be accepted by the user community as being appropriate for the application.

The methods by which the measurements are obtained and the data products derived shall be reproducible by any end-user at any time in the future. It should be kept in mind that these end-users are likely to use GRAUN data for decades to come. They shall be able to reproduce how measurements were made, which corrections were applied, and be informed as to what changes occurred during the observation and post-observation periods to the instruments and the algorithms. Hence maintenance of comprehensive rich metadata regarding data provenance and processing is key.

In brief, reference within GRUAN means that, at a minimum, the observations are tied to a traceable standard, that the uncertainty on the measurement (including corrections) has been determined, and that the entire measurement procedure and set of processing algorithms are properly documented and accessible.

## Managing Change

This section is being developed as a stand-alone document and once completed will be reintegrated into the manual.

GRUAN recognizes that change is inevitable – changes in instrumentation, changes in operating procedures, changes in data processing algorithms and changes in operators.  GRUAN appreciates that without change, improvement is impossible. Therefore, the goal is to manage changes that have scientific value in a way that does not compromise the integrity of the long-term climate records being measured while, as far as practical, minimizing those changes that have no likely scientific benefits. Given the wide range of observing systems that potentially may be deployed as part of GRUAN operations, the protocols for (a) assuring high stability and (b) deciding when an improvement merits a change to the GRUAN methods of observations will need to be developed as required by WG-ARO and developed by the appropriate instrumentation specialists, given guidance on the user requirements when required by GANTDOR.

With the radiosonde observations, the standard procedures recommended by the LEAD centre shall be used, and the equipment and methods of observation in daily use shall not be changed, without agreement from WG-ARO, as advised by the GRUAN Lead centre. Improvements  to performance  can be developed at GRUAN sites, but the evidence that the improvement justifies changing the GRUAN radiosonde protocol must be rigorously assessed, before any change to GRUAN observations is considered by WG-ARO.

Managing change within GRUAN is a key component of maintaining network homogeneity. If changes are implemented unilaterally at a single location, and even if those changes are implemented such that the long-term homogeneity of the measurement record at that station is preserved, the change may introduce inconsistency with other stations in the network. Therefore, changes in measurement systems at GRUAN stations shall be conducted in such a way that the homogeneity of the resultant GRUAN data products is not compromised.

The first focus in managing change as detailed in the GCOS climate monitoring principles is to ensure that when transitioning from older to newer instrumentation, that a sample of coincident measurements, sufficient to quantify any biases between the two systems, is obtained before the evidence is presented to the appropriate GRUAN task team. For example flying dual ozonesondes has proven to be useful when shifting from one ozonesonde system to another or from one standard operating procedure to another (Boyd et al., 1998). The length of time for which the older and newer systems should be run in parallel, and the frequency with which coincident measurements should be made, will depend on the instruments used and on an in-depth understanding of the measurement technique. The overlap period may also depend on the site since seasonal variability may differ between sites such that a site experiencing greater atmospheric variability may take longer to derive a robust estimate of differences between measurement systems than a site experiencing lower atmospheric variability. Such decisions shall be informed by robust scientific investigations. Until the results of such research are available, sites should err on the side of caution and undertake an approach that maximizes overlap so that sub-sampling can be undertaken later to determine a minimum safe level of overlap required to preserve the record (see approach used at Tateno in Section 2.2.1).

As new and more in-depth knowledge of various measurement systems is gained, reprocessing of historical data will be necessary. Data reduction processes and data archiving within GRUAN shall be designed with this in mind i.e. that the original raw data (which must always be archived) can be easily and regularly reprocessed, as required, to form a single homogeneous time series that is then provided to end-users. Each change in instrumentation, operating procedure or data processing algorithm is likely to require reprocessing of historical raw data to ensure a consistent product. Protocols shall be established to indicate when reprocessing of the full measurement record is justified. Every reprocessing generating a new homogeneous time series over the complete measurement period shall be reflected in an increment in the data version as prescribed in the data versioning protocols developed by the Lead Centre. Such updates shall be communicated to users who have accessed earlier versions of the data (see Section 8.6). For this reason it is also important that all older versions of any data set are always made available through the GRUAN archives.

A discussion of specific sources of changes is presented below but in general this requires dealing with breakpoints in the measurement time series. It is far more preferable that these changes are identified a priori through the available meta-data that identifies such changes. However, it is also possible to identify breakpoints in measurement time series based on the statistical behaviour of the data themselves. Significant resources and techniques have already been developed within the surface climate community (see e.g. http://www.homogenization.org) and upper air climate community around this issue (e.g. Haimberger, 2007) although homogenizing upper-air records is more challenging that homogenizing surface climate records.

These techniques must be grounded in quantitative understanding of the causes of offsets and drifts between two different measurement systems i.e. the reliance should not be on the implementation of signal processing techniques that identify and correct for offsets and drifts in time series. This quantitative understanding in turn should emerge from the meta-data associated with each measurement and from in-depth knowledge of each measurement system.

### Changes in instrumentation

Changes in instrumentation are both inevitable and desirable if they lead to more precise measurements of the true atmospheric state or if they can save on the cost of making the measurement (without compromising measurement quality). Instrument changes will also often be driven by the necessities of production engineering (when instrument components become unavailable or too expensive) and decisions will have to be made as to what level of component change requires additional change testing. Formal instrument intercomparisons will be essential for developing the in-depth understanding required to manage changes from one instrument to another and for informing decisions on the relative advantages and disadvantages of changing instrumentation. For this reason, participation in formal intercomparisons should be a pre-requisite for the adoption of any instrument within the GRUAN network. Outcomes from such intercomparisons would form an important component of the meta-data archived at the GRUAN Lead Centre. GRUAN need not necessarily organise these intercomparisons themselves. WMO and partner networks (e.g. NDACC) often run instrument intercomparison campaigns and GRUAN should participate in these and share the data where possible. Such participation would be mutually beneficial to both communities. GRUAN needs to work closely with CBS and CIMO to gain maximum benefit for all parties from these intercomparisons. In addition to intercomparisons of similar instruments (e.g. radiosondes), intercomparisons between different instruments measuring the same ECV will also be highly informative (e.g. comparisons of ozonesondes, ozone lidars and ozone microwave radiometers at a single site). A number of case studies exist which can be used as examples of how to manage changes in instrumentation. For example the impacts of changes from the Meisei RS2-91 type radiosonde to the Vaisala RS92-SGPJ type GPS sonde at Tateno were quantified by conducting dual sonde flights during four intensive observation periods in December 2009, and in March, June and September/October 2010.

Following a scientifically robust replacement strategy that maximises the maintenance of long-term climate records will be important for ensuring the integrity of the GRUAN data products in the face of change. A goal within the 'Management of Change' research topic of the GATNDOR team is to provide scientific bases to develop operational practices to better manage instrument changes at GRUAN sites and to accurately merge disparate data segments to create a homogeneous time series. Consideration will need to be given to the desired strategy when more than one station in the network is making an identical (or very similar) change with respect to timing, sharing of data, and whether certain sites will act as pioneers. This will be especially important where the change is forced by a supply issue.

Measurement redundancy (see Section 6.2) has significant benefits for managing instrument change as a second instrument measuring the same ECV can be used as a common reference against which both old and new instruments can be compared over an extended period. This benefit increases further when three or more instruments measure the same ECV and any changes are substantially staggered. An ideal aim that assures the record is therefore at least triple redundancy. The same advantages could be achieved through the use of a travelling standard instrument. For in situ balloon-borne instruments, consistent ground-check routines between new and old instruments will minimize changes in procedural uncertainty contributions.

Dealing with changes in instrumentation will require GRUAN to establish close two-way links to instrument manufacturers. Inclusion of the Association of Hydro-Meteorological Equipment Industry (HMEI) in discussions of instrument change within GRUAN would be advantageous. A productive point of interaction with the different vendors and manufacturers would be the periodic GRUAN participation in the CIMO multi-sensor field campaigns. Engaging the manufacturers in these field campaigns will assist GRUAN not only in evaluating the different sensors but also as a point of interaction with the vendors apart from the limited HMEI attendance at GRUAN meetings. A close cooperation between GRUAN and instrument suppliers will also help GRUAN to better understand industry capabilities and to better quantify instrumental uncertainties. This cooperation will also help suppliers to better understand GRUAN requirements, and the industry would be able to advise GRUAN of its current and prospective abilities to meet these requirements. For many of the parameters of interest (as instruments of required accuracy do not yet exist), GRUAN aims to further their development in cooperation with instrument manufacturers. HMEI has suggested that a workshop specifically for manufacturers and open to all HMEI members would be helpful.

Detailed archiving of instrument meta-data will be vital to managing changes in instrumentation. This will allow later reprocessing of the raw data as 'deep' as possible. Since it is not always known in advance which meta-data are likely to be required for reprocessing at a later date, GRUAN operators should err on the side of collating as much meta-data as possible about measurement systems even if no immediate use for those data can be envisaged. In all cases sufficient meta-data must be available to tie the new instrument via a comparable traceability chain back to the same recognized standard as the old instrument.

### Changes in operating procedures

Even if instruments themselves do not change, changes in the operating procedures for an instrument may also introduce breakpoints in a measurement time series. For the most part, changes in operating procedures should be dealt with in a fashion similar to changes in instrumentation e.g. reprocessing of historical data to homogenize the time series and redistribution of the data with an updated version number will almost certainly be required. The expectation is that standard operating procedures for all instrument types within GRUAN will be archived at the Lead Centre and that this body of material will be used to advise sites through transitions in operating procedures. As discussed in Section 5.4 the extent to which I site is prepared to conform to GRUAN standard operating procedures will be one of the criteria used when evaluating the potential inclusion of the site in GRUAN.

### Changes in data processing algorithms

New knowledge and resultant improvements in reduction of raw data to useful measurements are likely to lead to changes in data processing algorithms. As for changes in operating procedures, such changes in data processing algorithms should be dealt with in a fashion similar to changes in instrumentation. At the very least every change in data processing algorithm must be reflected in a change in version number of the final data product. Because raw data from various GRUAN sites will be processed at one location and one location only (either the Lead Centre or some other GRUAN site with particular expertise in that measurement), changes in data processing algorithms shall be implemented uniformly across the network. To achieve homogeneity across the network it is important that individual sites do not independently implement changes in data processing algorithms for data submitted as GRUAN data even if those changes are well documented and follow the prescriptions listed above. This more central, 'top-down' approach to data processing is different from the more decentralized approach employed in other networks. While such enforced conformity incurs an operational cost, the advantage is that end-users of the GRUAN data products will see data homogeneity not only in time for single stations, but also between stations. In support of maintaining consistency in the use of data processing algorithms within GRUAN, the Lead Centre shall maintain an archive of data processing algorithms which then also comprises an important part of the meta-data archive for GRUAN.

Tension may arise where a site may wish to implement a non-standard (at least non-standard for GRUAN) data processing algorithm for some purpose e.g. to create a data product that is tailored for a specific need. Such eventualities can be accommodated by having a central processing facility for each GRUAN product (see above) where a common data processing procedure is applied to the ‘rawest’ form of data collected. This would not preclude a site from implementing non-standard processing of the raw data and serving this for their own purposes.

### Change in operators

Ideally the quality of the measurements should be immune from changes in operators. This is more likely achievable if standard operating procedures are developed where there is reduced opportunity for idiosyncrasies of operators to affect the measurements. Meta-data should include codes (not names to protect the privacy of operators) to denote where different operators have been responsible for measurements.

# MEASUREMENT UNCERTAINTY

## Estimating measurement uncertainty

No measurement can be made perfectly, and estimating measurement uncertainty is a central tenet in GRUAN’s operations. A common GRUAN definition of measurement uncertainty and a common procedure to establish measurement uncertainties is required to homogenize uncertainty estimates across the network. It is also needed to make the steps leading to the determination of measurement uncertainty traceable. This common definition should, ideally, be adopted by instrument providers as well.

Achieving a useful estimate of measurement uncertainty may require as much, if not more, effort than making the measurement itself. However, such effort is necessary to achieve the goal of GRUAN to provide reference measurements from the surface to the upper stratosphere. The availability of an estimate of the measurement uncertainty for every measurement made within GRUAN will significantly improve the utility of the measurements and will elevate the GRUAN measurements above what is currently available.

The availability of sufficiently detailed meta-data is vital to quantifying random errors and biases in measurements. The more detailed the meta-data, the 'deeper' the measurement uncertainty can be traced. The approach that should be followed is that where some calibration, reference standard, application of an operating procedure, or use of a data processing algorithm introduces a source of uncertainty into a measurement, complete details about that uncertainty source must be available through the meta-data tagged to that measurement. Such sources of meta-data may include (Immler et al., 2010) previous measurement data, experience with or general knowledge of the behaviour and properties of relevant materials and instruments, manufacturer’s specifications, data provided in calibration and other certificates, and uncertainties assigned to reference data taken from handbooks. It is vital that all sources of measurement uncertainty are made transparently available to end-users of GRUAN measurements.

A particular challenge for GRUAN in estimating measurement uncertainty is that for in situ measurements of upper-air ECVs, the instrumentation operates in conditions that are difficult to replicate in a controlled environment (e.g., a test chamber). Calibration of the instrument in its operating environment where e.g. transient influences of changes in solar radiation and/or clouds are likely to affect sensor characteristics is generally not possible. Furthermore, the staple instruments for much of GRUAN, viz. balloon-borne sondes, are used for measurements of single profiles. The well calibrated instruments with quantified measurement errors are discarded after each profile measurement and re-calibration or re-characterization after a measurement is often not possible even if the instrument is recovered. The emphasis is then on employing standards that ensure stability, traceability, and uniformity between instruments and across the GRUAN network as a whole.

Because one of GRUAN’s primary goals is to detect long-term climate trends in the upper atmosphere, the primary consideration might be to work towards reducing changes in the systematic bias in measurements. However, because GRUAN data are likely to be used for other purposes such as satellite validation, acting as a reference for GUAN, or as input to global meteorological reanalyses, reducing biases to achieve the best possible accuracy also needs to be a priority. Therefore the aim should be to identify and minimize both biases and random errors, and to include the effects of both when calculating measurement uncertainties.

The GRUAN policy for dealing with measurement uncertainty shall be:

- Describe/Analyze all sources of measurement uncertainty.

- Quantify/Synthesize the contribution of each source of uncertainty to the total measurement uncertainty.

- Verify that the derived net uncertainty is a faithful representation of the true uncertainty.

### Describe/Analyze sources of measurement uncertainty

The first step in the process of deriving an uncertainty associated with any measurement is to first fully explore and describe each source of uncertainty in the form of biases and random errors. Contributions to the net measurement uncertainty are likely to include sensor calibration, sensor integration, sensor performance and external influences to operational routines such as sensor preparation and sensor ground-checks. While a specific sensor might perform well, if its value depends in some way on another sensor that performs less well, this source of uncertainty needs to be accounted for. For example, if a very precise and accurate temperature measurement is made but the vertical coordinate for that measurement is a less precise pressure measurement, in the presence of large (T/(p, the uncertainty in pressure can introduce significant uncertainty in the temperature measurement. Therefore uncertainty in the geo-location and time coordinates associated with each measurement shall also be considered when identifying and describing sources of measurement uncertainty. A full list of sources of measurement uncertainty will be defined in the GRUAN common definition of measurement uncertainty terms. Every GRUAN station shall measure, collect, and provide all information necessary to establish an uncertainty budget for every measurement.

### Quantify/Synthesize sources of uncertainty

The second step is, where possible, to quantify and correct for any measurement biases. Uncertainty in such bias corrections, which shall also be diagnosed, documented and quantified, then contributes to the random error on the measurement. Once all biases have been corrected for, and assuming all remaining random errors are normally distributed about the mean, the resultant net uncertainty on the measurement can be reported as a single value i.e. the first standard deviation of the distribution (1σ errors). Where systematic biases cannot be determined, or perhaps can be determined but cannot be corrected for, or when remaining random errors are not normally distributed about the mean, a different approach is required for quantifying the net uncertainty on the measurement. In such cases, because the net uncertainty is no longer represented by a Gaussian distribution, it cannot be reported as a single value. Techniques to fully describe the shape of the error distribution must then be developed and higher order moments of the distribution (e.g. the skewness or kurtosis) would need to be reported as part of the measurement uncertainty description. If a measurement process can be simulated, and if the probability distribution functions (PDFs) of the various sources of uncertainty are well known, a Monte Carlo approach can be used to generate a large ensemble of ‘virtual’ measurements from which measurement uncertainty statistics can be calculated. This approach can be used no matter how structured or asymmetrical the individual PDFs might be. This approach has been used to estimate asymmetric errors in ozonesonde measurements (Bodeker et al., 1998).

### Verify measurement uncertainties

The uncertainty budget for every GRUAN measurement should be verified at regular intervals using redundant observations from complementary instruments (see Section 6.2). If coincident observations of the same ECV are available and are subjected to the same uncertainty analysis, the degree to which the measurements agree within their stated uncertainties is indicative of the validity of the measurement uncertainties. If measurements agree within their uncertainties, the error estimates on the measurements are more likely to be correct. Formal methods have been developed to achieve this (Immler et al., 2010). For example, if two large sets of data are compared and more than 4.5% of the data are statistically significantly different within their error bars, then either a systematic effect in either or both measurement sets has been overlooked, or the uncertainties have been under-estimated. On the other hand, if much less that 32% of measurement differences are smaller than the RMS of the uncertainties, then the measurement uncertainties have probably been over-estimated. This verification by itself does not provide a statement about the usefulness of a measurement; it only provides information about the completeness of an uncertainty analysis. Including such comparisons in operational data processing can act as a flag for where error analysis within the processing may not be complete.

GRUAN includes both in situ and remote sensing methods. In the case of in situ methods, the instrument is generally calibrated directly to the geophysical quantity of interest. In the case of remote sensing methods, the calibrated data are in physical units of radiance and/or frequency, which are then analyzed to provide estimates of the underlying climate variable of interest. Validation of data products, which is equivalent to verifying measurement uncertainties, is therefore a two-step process whereby the accuracy of both the instrument calibration and the analysis algorithm, are validated.

## Reporting measurement uncertainty

An overarching principle for the operation of GRUAN is that no measurement should be provided without also providing an estimate of the measurement uncertainty. Where all sources of systematic error in the measurement have been identified and corrected for, the measurement uncertainty can be quoted as the standard deviation of the random error. As discussed above, where biases remain in the measurement, or where the net random error in the measurement does not follow a Gaussian distribution, alternative methods for reporting the measurement uncertainty must be considered. This may be in the form of establishing 1σ upper and lower bounds on the measurement uncertainty to denote that the uncertainty is asymmetric – generally reported aswhere X is the measurement, u is the 1σ uncertainty in the positive direction and l is the 1σ uncertainty in the negative direction. For more complex distributions of measurement uncertainty it may be necessary to quote the most likely value i.e. the peak in the PDF for the measurement and parameters that detail the shape of the PDF (or a pointer to the PDF itself).

## Reducing measurement uncertainty

Changes in instrumentation or operating procedures may lead to reductions in measurement uncertainty. In such circumstances it is important that the same detail of uncertainty analysis is conducted for the new instrument/operating procedure as has been done for the instrument/operating procedure to be replaced.

In some circumstances, e.g. in the presence of high natural variability (such as for temperature and water vapour), reducing measurement uncertainty has little impact on derived trends since the primary source of the variability in the trend estimate might be the measurand noise on the signal being analyzed. It is therefore important that scientific analyses guide where reducing measurement uncertainties is most likely to lead to reductions in uncertainties in trend estimates.

## Reducing operational uncertainty

Operational uncertainty includes uncertainties related to instrument set-up, sampling rates and the application of algorithms for data analysis. The contribution of operational uncertainty to the total measurement uncertainty in GRUAN is likely to be significantly reduced if the ‘rawest’ form of measurement data is submitted to a central GRUAN data processing facility (see Section 8.1) where a single verified, validated and well described data processing algorithm is applied to the raw data. Similarly, the adoption of an identical standard operating procedure for each instrument type across the network, would reduce the operational uncertainties related to instrument set-up. To this end, optimal standard operating procedures shall be developed at the GRUAN Lead Centre and then disseminated to all sites making that particular measurement and adopted where practical with exceptions clearly documented and agreed with the WG-ARO.

## Validating measurement uncertainty

Once the uncertainty on a measurement has been calculated, the question then becomes: how well does this measure of uncertainty represent the degree of confidence we should have in this measurement? Two approaches are available for validating the derived uncertainty on any measurement, viz.:

### Comparison of redundant measurements

A traditional way of validating measurement uncertainty is to measure the quantity of interest through two (or more) techniques, based on physically different measurement principles. Because the different techniques are subject to unique measurement uncertainties, comparisons yield a robust and continuous demonstration of measurement accuracy. Where simultaneous measurements of the same quantity are made using two different techniques, and disagree within their stated measurement uncertainties it suggests that either one or both of the measurements are erroneous, or that the measurement uncertainties are under-estimated. In this way, complementary measurement techniques with different susceptibilities to local conditions can be chosen to maximize the accuracy of the data record. Additionally, uncertainty budgets validated in this way may help identify other error sources that cannot be compensated for by complementary sensors, but may be monitored in situ.

### Laboratory analysis of the measurement system

The ability to simulate a specific measurement in the laboratory can permit an in-depth investigation of the various sources of uncertainty in the measurement. For example, the environmental simulation facility at the Research Centre Juelich (Smit et al., 2007) has provided information to validate measurement uncertainty in ozonesondes. Similar laboratory facilities are available for radiosonde at Lindenberg.

# ESSENTIAL CLIMATE VARIABLES MEASURED IN GRUAN

Since GRUAN’s goal is not only to provide long-term high quality climate records, but also the ancillary data required to interpret those records, a number of parameters in addition to the fundamental atmospheric state variables of temperature, pressure, and humidity would add value to the GRUAN database. High quality measurements of atmospheric state variables, trace gas concentrations, the atmospheric radiation environment, and cloud and aerosol properties will significantly aid the interpretation of priority 1 variables (see below). Many of these parameters have been identified by GCOS as Essential Climate Variables (ECVs; GCOS-92). A subset of ECVs has been selected as the most scientifically important and most tractable for GRUAN (see Appendix 1 of GCOS-112). As scientific research into the underlying causes of observed changes in upper-air climate advances, and as the capabilities of GRUAN sites expand, this list is likely to grow.

## Justification and context for Essential Climate Variables

The purpose of this section is to provide scientific justification and context, and more general guidelines for the measurement requirements for those ECVs listed as priority 1 for GRUAN, viz. temperature, pressure, and water vapour. The complete list of ECVs targeted by GRUAN is given in Appendix 1 of GCOS-112.A summary of material for the priority 2, 3 and 4 variables is provided in Appendix A.

The desired performance requirements  for each of the ECVs are based on the scientific requirements of the data and not on current instrument performance, so they may not be currently achievable.. In such cases the WG-ARO and Lead Centre shall provide possible incremental approaches to achieving the target attributes for each measurement.. Therefore, as stated in GCOS-112, these GRUAN requirements should be interpreted as eventual measurement goals of any given network site.

The values in Appendix 1 of GCOS-112 describe what is required of the measurements to meet specific research goals and a distinction needs to be made between what is desirable and what is feasible. While they may not be currently achievable, as measurement technology advances, attaining such targets should become more likely. In no case should an inability to achieve these targets result in the exclusion of a site or a measurement programme from the GRUAN network. However, a GRUAN site shall use currently available equipment in a manner ensuring optimum performance from that system. Development and improvement  of systems at GRUAN sites is to be encouraged, but these developments should be performed in a manner that does not interfere with the stability of GRUAN network observations.

The measurement ranges prescribed in Appendix 1 of GCOS-112 should cover the range of values likely to be encountered over the vertical range of interest so that any proposed instrument, or set of instruments, would need to be able to operate throughout that range. Measurement precision refers to the repeatability of the measurement as measured by the standard deviation of random errors. However, measurement precision is closely tied to the frequency of observations since observations are often averaged and the greater the sample size, the less stringent the required precision in terms of the uncertainty on the mean. Measurement frequencies are not specified because they may vary over time (see Section 7.2). Measurement accuracy refers to the systematic bias in a measurement (the difference between the measured or derived value, and the true value). It is not directly specified for many variables for which variations, and not absolute values, are needed to understand processes. Measurement accuracy is directly related to long-term stability, the maximum tolerable change in bias over time, which is a critical aspect of the reference network.

## Priority 1 ECVs

### Temperature

Scientific justification: Upper-air temperatures are a key dataset for the detection and attribution of tropospheric and stratospheric climate change since they represent the first order connection between natural and anthropogenically driven changes in radiative forcing and changes in other climate variables at the surface. Furthermore, the vertical structure of temperature trends is important information for climate change attribution since increases in atmospheric long-lived greenhouse gas (GHG) concentrations warm the troposphere but cool the stratosphere steepening vertical temperature gradients in extra-tropical regions. Other drivers of atmospheric temperature changes, e.g. changes in solar output, would not have the same vertical profile fingerprint. Remaining discrepancies between temperature trends derived from satellite-based measurements and from radiosondes weaken the attribution of changes in temperatures to changes in climate forcing agents. High quality temperature measurements within GRUAN will contribute to the resolution of these discrepancies.

Radiosondes remain a primary workhorse within the global upper-air network for the measurement of temperature, pressure and humidity, it is imperative that GRUAN sites establish state-of-the-art radiosonde measurement programmes that match the optimum stability of performance obtainable to date. In addition, efforts should continue to improve the quality of radiosonde measurements, where it is known there are significant limitations in performance for use in climatological obervations, see WMO(2011). Other measurement techniques can and should be developed to extend the height range of the temperature profile measurements and to reduce the random error and bias on the measurements. However, these should always be quantitatively inter-compared with collocated radiosonde measurements to provide a traceable link to the radiosonde measurements made within GUAN. Temperatures measured by high-quality radiosondes are needed to:

- Monitor the vertical structure of local temperature trends.

- Correlate changes in other parameters, especially water vapour (see below), with changes in temperature.

- Provide a reference against which satellite-based temperature measurements can be calibrated and adjusted so that long-term changes can be estimated globally with greater confidence.

- Validate temperature trends simulated by climate models.

- Provide input to global meteorological reanalyses such as NCEP, ECMWF, NASA, JMA.

- Provide input to numerical weather prediction models if and when submitted shortly after the measurement. Upper-air measurements of temperature and relative humidity are two of the basic measurements used in the initialization of numerical weather prediction models for operational weather forecasting. In turn, feedback from the numerical analysis potentially provides a useful metadata element in the final GRUAN measurement (see Section 9).

There are a number of satellite-based measurements of this ECV. Stratospheric Sounding Unit (SSU) instruments have been flown on TIROS-N and NOAA-6 to NOAA-14 operational meteorological since 1979. Three SSU channels with weighting functions of 10-15 km width, peak at 29, 38 and 44 km providing sampling of the middle and upper stratosphere until October 2005. SSU provides the only near-global source of data on temperature trends above the lower stratosphere over such a long period; it has been extensively used in assessments of those trends, and their possible causes (Randel et al., 2009). Global satellite observations from the Microwave Sounding Unit (MSU) Channel 4 and the Advanced Microwave Sounding Unit (AMSU) channel 9 provide a weighted layer mean temperature of the lower stratosphere between 13 and 22 km. This weighting primarily covers the stratosphere in the extra-tropics, but spans the upper troposphere and lower stratosphere in the tropics (the weighting function peaks near the tropical tropopause). The MSU/AMSU time series are derived by combining measurements from the series of satellite instruments that have been operational over 1979–2007. Since the mid-1990s,  radio occultation (RO) measurements of temperature have also become available (Steiner et al., 2009).

{Ed Note: This isn’t true as stated and is best left out}It is particularly important that trends in tropical cold point tropopause temperatures are accurately detected since this is thought to control the flux of water vapour into the stratosphere (Gettelman et al., 2002; Fueglistaler and Haynes, 2005) and changes in stratospheric water vapour influence radiative forcing and temperatures both in the lower stratosphere but also in the upper troposphere (Forster et al., 2007). At present temperature trend uncertainties in the lower stratosphere and upper troposphere remain large, particularly in the tropics. For this ECV, addressing trends in tropical cold point temperatures should be a focus for GRUAN. To this end establishing close working ties between the tropical GRUAN sites at Manus and Nauru with the sites within the SHADOZ network (Thompson et al., 2007)  and with the GUAN stations operating in the tropics would be particularly advantageous.

Measurement range: Ideally temperature measurements should cover the range 170 – 350 K to span the range of measurements encountered between the Earth’s surface and the upper stratosphere. Currently available technology can meet this requirement.

Vertical range: The effects of elevated concentrations of greenhouse gases on atmospheric temperatures are seen most clearly in the upper stratosphere (Shine et al., 2003). Ideally GRUAN measurements of the vertical temperature profile should extend from the surface to ~50 km. Vertical temperature profiles are most routinely measured using radiosondes which seldom reach above ~35 km altitude.

Vertical resolution: Given that it is primarily balloon-borne instruments that provide high resolution profiles of the vertical temperature profile in the atmosphere, a resolution of 100 m or better below 30 km altitude and a resolution of ~500 m above 30 km altitude is appropriate.

Random error: ≤0.2 K in measurement repeatability.

Bias: Biases ≤0.1 K in the troposphere and ≤0.2 K in the stratosphere.  In the dark this requirement can probably be met by several of the better operational radiosondes but not in the daytime, see WMO (2011) and the revision of chapter 11 in the  CIMO Guide, published in 2012.  The most accurate radiosonde in the day is possibly  the ‘Accurate Temperature Measuring Radiosonde’(Schmidlin, 1991), claiming an uncertainty of 0.3K throughout most of the upper troposphere and the stratosphere, but this is not yet widely available in sufficient numbers for use throughout GRUAN..

Thus, GRUAN should proceed with the best operational radiosondes available, using the methods of observation agreed with the GRUAN Lead Centre,  ensuring that sufficient  sites make a priority of temperature measurements in the dark. Development of commercially availabe new technology to achieve higher accuracy in the daytime is a priority

Stability: Better than 0.05 K/decade. The signal of change over the satellite era is in the order of 0.1–0.2K/ decade requiring long-term stability to be an order of magnitude smaller to avoid ambiguity.

### Water vapour

Scientific justification: Water vapour is the primary natural GHG and is central to global water and energy cycles. It acts primarily as a feedback, amplifying the effects of increases in other GHGs. Water vapour is the raw material for clouds and precipitation, and limited knowledge has compromised our ability to understand and predict the hydrological cycle, and understand its effect on radiative transfer (Peter et al., 2006). Water vapour is also a source of OH in the upper troposphere and stratosphere, influencing methane, ozone and halogenated GHGs. High clouds due to water vapour in the UT/LS affect both the planet's shortwave albedo and its longwave greenhouse effect, and both cloud particles and water molecules are involved in chemical reactions that govern stratospheric ozone concentrations. Fully quantifying the Earth’s radiation budget depends on an accurate assessment of the radiative properties of clouds and the water vapour continuum.

Changes in water vapour in the UT/LS exert a greater radiative forcing than changes elsewhere (Solomon et al., 2010). Standard radiosonde humidity sensors had very poor response at the low temperatures [below -50 ºC], pressures, and water vapour concentrations of the UT/LS (Wang et al., 2003) .However, there has been a lot of progress since 2003, see the results of the WMO Intercomparison of High Quality Radiosonde Systems, WMO(2011), although no operational radiosonde can be expected to measure with sufficient accuracy for climatological purposes ( see GCOS-112) in the lower stratosphere.

A number of factors, many linked to changes in climate, are likely to affect the flux of water vapour into this climatically important region of the atmosphere, viz.:

- Changes in the cold-point tropopause temperature (Zhou et al., 2001).

- Changes in convection. Convective transport of ice particles into the UT/LS can provide a path with bypasses the limitation imposed by the cold-point tropopause temperature.

- Changes in the Brewer-Dobson circulation (Austin et al., 2006).

While most of the Earth’s water vapour is contained in the lower atmosphere where it is can be measured as absolute or relative humidity, the water vapour content of the upper atmosphere is measured in parts per million and is difficult to measure accurately; the older generation of operationally-deployed balloon-borne instruments, and the satellite data record to date did not allow the measurement of water vapour in the upper troposphere and lower stratosphere to the required accuracy to be useful for climate applications (Soden et al., 2004). However, accurate water vapour measurements in the upper atmosphere are critical, especially for radiative transfer modelling. Understanding the water vapour budget throughout the atmosphere is also necessary for interpreting measurements of outgoing longwave radiation.

Satellite-based solar occultation and limb-sounding instruments can measure water vapour in the upper troposphere and stratosphere but inter-satellite differences preclude the use of earlier  data in long-term trend analyses (Rosenlof et al., 2001). High precision measurements of water vapour profiles will provide valuable input data to global meteorological reanalyses and data for validating global climate models.

Instruments such as the Cryogenic Frostpoint Hygrometer (CFH; Vömel et al. 2007b), the Fluorescent Advanced Stratospheric Hygrometer for Balloon (FLASH-B) Lyman-alpha instrument, can provide  water vapour measurements in the lower stratosphere, but are very expensive compared to operational radiosondes. The Snow White chilled mirror hygrometer is able to measure reliably in the upper troposphere at night. All these instruments require a much  higher skill level to to ensure reliable operation than an operational radiosonde. It does not appear justified to expect every GRUAN site to fly these systems once a month, where several GRUAN sites are in the same climate region, e.g. western Europe. The  variability of water vapour in the stratosphere over a given climatic region is not expected to be high, but as indicated above, a priority should be given to measurements in the tropics, when resources are available.

Modern operational radiosondes have much improved performance compared to those reported earlier and there has been a significant improvement between the WMO Radiosonde Intercomparison hosted in Mauritius in 2005 and that hosted in Yangjiang China in 2010, see WMO(2011).  The better sensors now start to become slow to respond at temperatures around -70ºC. A second source of error comes from assuming the temperature of the humidity sensor in the day is the same as that reported by the radiosonde temperature sensor. However, adjustment algorithms for this slow response have been implemented and methods of reducing the solar heating error have been implemented, so the relative humidity errors in the tropical upper troposphere are very much smaller than  in earlier operational radiosondes. The use of the better operational radiosondes in GRUAN will improve the capability to monitor changes in the upper troposphere day to day, although further development of the systems should be encouraged.

Many sites are currently developing the capability to observe and analyze data from ground-based GPS receiver, usually as part of a larger or international networks. These data provide continuous high-quality estimates of column water vapour which, in addition to being useful in their own right, can be used to partially validate the vertical humidity profile measurements; total precipitable water calculated from the radiosonde measured temperature and humidity profiles should compare well with that measurement by the GPS receiver.

Measurement range: 0.1 – 90000 ppm. The large range in values that needs to be covered by these measurements presents a challenge for instrument development and operation since no single commercially available instrument is responsive over this range. Instrument packages may therefore need to include more than one instrument, each of which covers a particular region of the atmosphere.

Vertical range: 0 to ~40 km.

Vertical resolution: 50 m below 5 km and 100 m above 5 km altitude.

Random error: 2% in mixing ratio in the troposphere and 5% in mixing ratio in the stratosphere.

Bias: 2% in mixing ratio throughout the profile. 1% for total column. This is more stringent than the 3% standard recommended in WMO-No. 8  by CIMO as currently achievable..

Stability: 0.3%/decade in mixing ratio and for the total column.

### Pressure

Scientific justification: Measurements of upper air temperature and water vapour  must be accompanied by the height/ pressure at which the measurement is made. Data used by WMO, e.g.  in numerical weather prediction primarily use standard geopotential heights, but conversion from geometric to geopotential height is straightforward, see WMO Guide to Observations(WMO No.8). Thus, to convert from geometric height as measured for instance by a GPS radiosonde, to geopotential height does not require knowledge of the vertical temperature structure.

In most numerical weather prediction models the observational data are input at levels defined by a ratio of the pressure to the surface pressure. Thus, the model must convert geopotential height into pressure, if the system does not provide pressure observations. Deducing pressure from the geopotential height requires knowledge of the temperature and water vapour structure in the vertical, and if this is not directly available from the system, the model will compute the values using its own analysis fields..

If a radiosonde uses a pressure sensor to compute geopotential height, the uncertainty in calculated geopotential heights will result from uncertainties in temperature, pressure and water vapour measurements. However , the majority of modern radiosondes are now using GPS navigation signals to measure height and when set up carefully can meet all the GRUAN requirements for pressure/height observations. The uncertainty in the GPS height measurements  has very little variation with height in the atmosphere, see WMO (2011). So CIMO recommend that GPS radiosonde are used at all GRUAN stations.  Thus , for instance problems with reproducing  structure in the vertical in the stratosphere become negligible with GPS radiosondes, e.g. with ozone profiles .

Measurement range: 1 – 1100 hPa

Vertical range: 0 – 50 km

Vertical resolution: 0.1 hPa[ better quoted in equivalent height]

Random error: 1 km       1 hPa ,      [equivalent height  error, 10m]

16km       0.3   hPa,         [equivalent height  error, 20m]

32 km      0.05 hPa ,        [equivalent height error, 30m]

48 km      0.01 hPa,         [equivalent height error,50m]

{Ed Note: the requirements as stated previously, impose far too high a requirement on pressure measurements near the ground.}

Bias: 		   Half the random error quoted above.

Stability:  Better than a quarter of the random error quoted above, per decade..

## Moving beyond priority 1 variables

This section is being developed in a separate document which will then be merged with the manual. The content from the previous version of the manual is provided below. The final version of this section will be far more comprehensive.

The emphasis to date within GRUAN has been on observations of priority 1 variables. This allows testing of the guiding principles for all reference observations before expanding the measurements at GRUAN sites to lower priority variables. A fully functioning GRUAN that serves all envisaged purposes will require measurements of all ECVs listed in Appendix 1 of GCOS-112. An approach to expanding site measurement capabilities to eventually cover as many of the specified variables as possible, whilst recognising that not all variables may be observed at all stations, is required.

# GRUAN SITES

This section is being developed in a separate document which will then be merged with the manual once approved by the WG-ARO and 66% of GRUAN sites. The content of the latest version of that document is provided below. It must be emphasized however that this section is very much a work in progress.

## Introduction

The purpose of this chapter is to define the process by which sites are assessed for GRUAN certification and the process by which that certification is maintained. Certification is essential to ensure that the sites within GRUAN operate at a level that maintains GRUAN’s status as a premier upper air climate monitoring network (Seidel et al., 2009). GRUAN is more than a collection of measurements made at individual sites. Part of the scientific benefit that will accrue from GRUAN results from the homogeneity of the reference quality standard of the measurements made at the sites comprising the network. A shortfall in maintaining that quality standard at one site reduces the confidence that users will have in the utility of the measurements made across the network as a whole. Sites therefore need to be sufficiently consistent and scientifically sound in their operation for scientific benefits to accrue. The site certification process assures all sites that they are operating to reference quality standards, that all sites in the network are operating to the same standard, and that the scientific benefits resulting from such homogeneity of quality across the network are guaranteed. This document provides pragmatic criteria that will be used to assess and certify existing sites as well as new site offers. These criteria are designed to be as transparent as possible and to minimize the overhead involved for all parties in the certification process.

Specifics regarding the site assessment and certification include:

- Site assessment and certification falls within the mandate of the Working Group on Atmospheric Reference Observations (WG-ARO) together with the GRUAN Lead Centre. Assessment and certification of sites within GRUAN is consistent with the guidance developed with the WMO Commission for Instruments and Methods of Observation (CIMO; WMO-No. 8) and the WMO Commission for Basic Systems (CBS; WMO-No. 488).

- Sites seeking to become GRUAN sites will first be assessed according to their ability to meet the mandatory operating protocols defined in Section 5.2 and then according to the added value they bring to the network, as defined in Section 5.3.

- Sites will propose specific measurement programmes for inclusion in GRUAN and it is these that will be required to conform to the operating protocols defined in Section 5.2 and which will form the basis for assessing the added value that the site brings to the network as a whole. This will enable sites to operate some, but not necessarily all, measurement programmes to GRUAN standards.

- Determining whether the operating procedures for proposed measurement programmes meet the prescribed operating protocols will be done objectively against the standards outlined in Section 5.2.

- While based on a body of sound scientific research, in assessing the value which a specific site adds to the network, the WG-ARO will exercise its discretion in evaluating the proposal against the criteria defined in Section 5.2.

- The Lead Centre and WG-ARO will provide written feedback to each site as part of the certification process.

- To identify potential issues early, sites will be reviewed annually based on their annual reports (see below) which must highlight any operational anomalies, and based on reports on data flow, site performance etc. from the Lead Centre. Such reviews will be ‘by exception’ reviews. More complete site audits will be undertaken every 3 to 4 years (see Section 5.5 below for more details).

Where site reassessments identify measurement programmes that consistently fall short of GRUAN operating standards, GRUAN certification of that programme will be suspended. If all measurement programmes at a site lose their GRUAN certification and where jointly developed recovery plans for the measurement programmes at the site have repeatedly failed to resolve outstanding issues, the site will be suspended from the GRUAN network. The WG-ARO and Lead Centre will work pro-actively with sites to remedy these issues wherever possible in a timely and cost-effective manner.

### Geographical coverage

. GRUAN should aim to sample as many as possible of the  major climatic regimes and environment types to ensure that different temperature and radiation environments are reliably calibrated. Expansion of the network should concentrate on climatic zones and regions that are under-sampled in the initial network configuration. Geographical coverage of GRUAN sites should also be tailored to meet the specific needs of end-users e.g. the satellite-based measurement community is likely to want validation data in key regions of the atmosphere, with relatively uniform surface properties surrounding the site.

### Funding

Candidate GRUAN sites able to demonstrate reasonable expectations of funding to maintain operations over many decades will be valued above sites that cannot. The WG-ARO should have in place procedures for supporting long-term funding applications to local funding agencies for sites seeking to join GRUAN.. Thus realistic aims must be set for GRUAN to ensure operating costs are kept as low aspossible. Thus it is sensible to have a sizeable  proportion of the GRUAN sites where the facilities are already mostly funded for weather observations.Having sites transiently joining and then leaving the GRUAN network could compromise the goal of ensuring data homogeneity across the network e.g. if trends in ECVs differ at two different stations which measured the ECV over different time periods it is not clear whether the differences arise from the geographical separation or from different time periods being sampled.

## Mandatory Operating Protocols

GRUAN is a heterogeneous network that includes sites from both the research community and the operational meteorological community. The mandatory requirements for sites reflect GRUAN’s primary goal of providing reference quality observations of the atmospheric column (see Section 2). The emphasis is on how the measurements are made rather than specifically on what measurements are made. These requirements define GRUAN’s unique nature while accommodating the diverse capabilities of sites within the network. These protocols also recognize that GRUAN is not the sole stakeholder at any of the sites. Recognizing these boundary conditions, sites shall:

- Provide reference quality observations as defined in Section 2. In particular every measurement must be traceable to fundamental standards and calibrations through well documented routes.

- Provide uncertainty estimates for each datum or collaborate with other sites, instrument developers, GRUAN Task Teams and the GRUAN Lead Centre to provide these estimates in a consistent manner for a given instrument across the network. Profile measurements require uncertainty estimates for each measurement point on the profile. Documentation describing the calibration methods applied to each measurement, and the sources of measurement uncertainty excluded and included in the uncertainty estimate, must be provided.

- Provide access to raw data and assure long-term storage of the raw data either at the site, another GRUAN facility, or at another internationally accessible archive.

- Provide complete metadata for each measurement as defined in the requirements document developed by the Lead Centre. Metadata need to be sufficient to allow reprocessing of raw data by an independent party and will depend on the measurement system employed.

- Provide regular traceable ground checks, independent of the manufacturer, for any instruments which provide vertical profiles extending from the surface.

- Provide regular calibration information about the measurement systems (in-situ and remote sensing).

- Provide redundant reference observations of the essential climate variables (ECVs; GCOS-138) selected for measurement at the site at intervals sufficient to validate the derivation of the uncertainty on the primary measurement (noting that this validation is generally achieved through comparison against other recognized reference observations).

- Provide annual reports summarizing GRUAN operations at the site, the extent to which standard operating procedures developed for the network as a whole have been adhered to, changes in instrumentation, how those changes were managed, improvements made etc. Present these reports at the annual network meeting.

- Selected GRUAN sites may strive  to improve measurement accuracy, actively conducing research through intercomparisons, laboratory studies, work with other GRUAN sites and/or cooperation with manufacturers to improve measurement accuracy.

- Manage change pro-actively as defined in Section 2.3.

- Have a representative on the task team of site representatives and a nominated reserve contact for GRUAN purposes.

- Actively communicate with the Lead Centre, WG-ARO, Task Teams and/or other sites, (e.g. through attendance of meetings, blog postings etc.).

These mandatory operating protocols do not replace the target requirements defined in GCOS-112 and GCOS-121, which remain the targets for GRUAN, but rather emphasize the importance of how the measurements are made, and in particular what is required to guarantee reference quality observations, rather than what physical measurements are made.

## Criteria for Assessing Added Value

Once a site has committed to operating a set of measurement programmes under the protocols defined in Section 5.2, the added value that a site brings to the GRUAN network will be a function of:

- The extent to which the site measurement programmes provide measurements in regions, or of atmospheric phenomena, which were not previously sampled. In this case, the added value will be a function of the locations and capabilities of the sites already participating in the network.

- The extent to which a site brings unique observational and/or analysis capabilities to the network as a whole and the likelihood of being able to propagate those capabilities across other sites in the network.

- The participation of that site in other international measurement networks (e.g. GUAN, GAW, NDACC,  BSRN). Participation in existing networks reduces start up costs for establishing the GRUAN site and it quantitatively links the GRUAN measurements to the measurements being made in those other networks

- The extent to which a site can commit to a multi-decade programme of measurements. While it is recognised that a multi-decade programme of measurements cannot be guaranteed, a statement of intent with documented support (e.g. from the host institution or relevant funding agency) will add to the assessment of the value that the site brings to the network.

- The extent to which a site can fulfil the following measurement programmes expected of a candidate GRUAN site:

- At least twice daily (00 and 12 LST as a preference over UTC) measurements of vertical profiles of temperature from the surface to ~30 km and water vapour in the troposphere. High quality surface measurements of these same variables are also required to provide a traceable link between the measurements at the lowest level of each profile. Where feasible, occasional soundings at both 00/12 LST and UTC could be used to establish climatologies of differences, including uncertainties, which could thereafter be used to relate measurements made at one standard time to measurements made at another.

- At least monthly observations of stratospheric water vapour to ~30 km, at nominated sites.. Given that high frequency natural variability in the lower stratosphere is relatively small, these profile measurements should be made when most practical and when the altitude coverage can be maximized (see Section 7.1.4).

- Hourly observations of integrated precipitable water vapour.

These high priority measurement programmes will be refined as the research which forms their basis progresses. At such a time this section of the manual will be updated to reflect these changes.

While weekly sampling significantly underestimates monthly standard deviations in temperature, differences between detectable trends for weekly sampling compared to 12 hourly sampling may be acceptably small (Seidel and Free, 2006). So, for example, a site that makes weekly reference quality radiosonde measurements of temperature, pressure and humidity in a large region of the globe containing no other GRUAN stations might be assessed as adding as much value to the network as a site making 12 hourly reference quality measurements but located very close to another site making the same measurements, but operational practice suggests weekly measurements by non-experts does not give reliable results, usually.

- The availability of historical measurements that conform to the GRUAN standard. Detailed documentation describing how changes in standard operating procedures, instruments, data processing algorithms and operators over the history of the measurement programmes have been managed to ensure that the historical measurements are reference quality. Where historical reference quality measurements are available, consideration will be given by the WG-ARO and Lead Centre to providing these as GRUAN data through the GRUAN data archives.

- .

- .The level of institutional support for the site (and any partner institutions) to undertake fundamental scientific research of the measurements from the site and other GRUAN sites. Because GRUAN includes aspects of both operational and research networks, a critical mass of scientific analyses is required on a continuing basis to ensure that GRUAN fulfils its role as a research network. The degree of historical or planned cooperation with other sites both within and outside the GRUAN network including other GRUAN-relevant networks. It may be the case that while a single station might not be able to provide the full range of ECV measurements required by GRUAN, a group of two or more stations, located sufficiently close together, might have the combined capability of providing the full range of measurements. Such a collection of stations may then act as a single GRUAN site. Robust analysis would be required to demonstrate the extent to which such a cluster of sites might act as a single GRUAN station (see Section 6.5)

Such assessments of added value rely on the expert judgement of the WG-ARO and Lead Centre, recognize the heterogeneity of the sites within the network, and facilitate a practical approach to expansion of the network following the 2009-2013 implementation phase for GRUAN (GCOS-134).

## The Assessment and Certification Process

Figure 2: A schematic representation of the site assessment and certification process.

A schematic of the site assessment and certification process is provided in Figure 2. Proposals for the addition of new sites to GRUAN are likely to happen through two possible routes, viz.:

- Sites being approached by the WG-ARO and/or Lead Centre and invited to become GRUAN certified sites.

- Sites being proposed externally e.g. through the National Weather Service of the host country to the Lead Centre or WG-ARO.

Once a site has been identified for possible inclusion in GRUAN, through either of the routes listed above, the following sequence of events will be used to assess the site for potential GRUAN certification:

- Provision of the GRUAN requirements document (the external document comprising this chapter), the GRUAN manual, guidelines for the operation of specific instruments in widespread use in GRUAN, as well as documentation describing data submission protocols and the procedures that must be followed when data are submitted to the internal GRUAN archives, to the candidate site by the Lead Centre.

- The response from the candidate site should include:

    - A list of the measurement programmes at the site proposed for inclusion in GRUAN. This need not necessarily include all measurement programmes at the site. If a new or existing measurement programme is later proposed for inclusion in GRUAN, a similar procedure to that defined here will be used to include that programme in the GRUAN certification for that site.

    - A complete description of how those measurement programmes will be conducted. Such information would include, for example, detailed standard operating procedures for each of the measurement programmes, including a description of data storage policies, and a description of how systematic and random uncertainties in the measurements will be derived and reported. This information must be sufficient to establish the ability of the site to meet the mandatory operating protocols detailed in Section 5.2.

    - For measurement programmes for which a GRUAN data product has not yet been well defined, the site must describe their intended strategy for developing the existing observational product into a GRUAN data product that fulfils the mandatory operating protocols defined in Section 5.2. In such instances, cooperation with other sites already in the network is highly desirable to ensure that this expertise is disseminated to similar measurement programmes in operation at other sites.

    - The management structure of the site and a general description of the manner in which the site is operated. This would include a description of current and expected future funding levels for ongoing operation of the site.

    - A description of which databases the measurements have previously been submitted to and are currently being submitted to.

    - A description of how measurements to date at the site have been processed. This will be used to assess whether the time series to date meet the standards for a GRUAN reference measurement. Particularly important in this regard will be detailed documentation around how changes in standard operating procedures over the history of the measurement programmes have been managed to derive a homogeneous time series of measurements. Since the historical database of measurements is an important aspect for assessing the added value that a site brings to the network (see Section 5.3), it is particularly important that the historical data can meet the stated GRUAN requirements for long-term homogeneity.

    - A list of the scientific experts employed at the site who would likely participate in the analyses of the data collected within GRUAN.

    - Any additional information required to assess the site against the requirements listed in Sections 5.2 and 5.3.

- There is likely to be some iteration between the Lead Centre and the candidate site to confirm specific details, fill in information gaps, and finalize the documentation from the candidate site.

- Based on the documentation received from the candidate site, the Lead Centre will then write a short recommendation. This, together with the documentation from the candidate site, will then be submitted to the WG-ARO who will evaluate the proposal within 6 calendar months against the requirements listed in Sections 5.2 and 5.3. A visit to the site by members of the WG and/or Lead Centre may be required to obtain specific additional information about the measurement programmes slated for inclusion in GRUAN at that site. If accepted, the measurement programmes defined in 5.2a will then be included in the GRUAN certification for the site.

- Regardless of the outcome, the WG-ARO and Lead Centre will provide written constructive feedback to the candidate site outlining strengths and weaknesses of their programme for GRUAN purposes and suggestions as to future improvements for GRUAN operational purposes. This feedback is non-binding but rather intended to provide useful guidance and support to site capability development and retention of current capabilities.

Sites currently within GRUAN will be assessed and certified in a similar manner before GRUAN becomes fully operational at the beginning of 2014.

## Site Auditing

Certification of GRUAN sites will not be a single event. Periodic (e.g. every 3-4 years) complete auditing of the measurement programmes included in the GRUAN certification for a site will be conducted to ensure that the programmes continue to meet GRUAN standards. Such an audit may include:

- A review of annual reports from sites on GRUAN activities.

- A written report from the site – essentially an update of the original report written to initiate the assessment and certification process.

- A site visit by selected members of the WG-ARO and the GRUAN Lead Centre. Such a visit would include discussions with the researchers responsible for the measurement programmes at the site.

It is important for external perceptions of GRUAN integrity that these audits are conducted by the WG-ARO and Lead Centre and are not based exclusively on annual station reports. In the eventuality of identified site issues the following protocols will be followed:

- Should a measurement programme at an existing GRUAN site show significantly reduced observational capability over more than a year, as evaluated by the criteria listed above, the WG-ARO and Lead Centre will investigate the circumstances at that site, and, if needed, exclude that programme from the GRUAN certification for that site. The WG-ARO and Lead Centre will work pro-actively with sites to resurrect such programmes providing what technical and in-kind support is practical.

- Should the overall contribution of a site be deemed sufficiently diminished to call into question its continued presence in the network, the site will be informed immediately in writing. The site will be given six months to form, in consultation with the Lead Centre and WG-ARO, a capabilities recovery plan. Should this plan be accepted the site will have no more than two calendar years from its acceptance to implement agreed key aspects. In the eventuality that this is not achieved, the site will be suspended with an invitation to submit anew at such time as issues are remedied.

# INSTRUMENTATION

## Instrument selection

. Periodic review of instrumentation likely to be of use within GRUAN shall be undertaken since instrument technology is constantly evolving. It must also be recognized that not all sites within GRUAN will operate the same instrumentation, e.g. a new site may decide to adopt the most recent technology while a site that has a multi-decade record using an older instrument may decide to continue to use that instrument to potentially avoid introducing a discontinuity in the measurement time series. The emphasis is therefore not on prescribing an instrument, but rather on prescribing the capabilities of an instrument and allowing individual sites to select an instrument that achieves those capabilities in the context of the scientific, programmatic, and practical constraints on the site. That said, the fewer the number of instrument types deployed within GRUAN, the more likely network homogeneity will be achieved.

A number of factors should be considered when selecting instruments for use in the GRUAN network including (Immler et al., 2010):

- Instrument heritage: How long has an instrument been in use by the community and for what purpose? In what other networks is the instrument deployed? How substantial is the body of literature documenting its performance and measurement uncertainty? How widely distributed is the knowledge base that facilitates the instrument’s successful operation?

- Sustainability: Are the costs for operating the instrument and the demands on personnel for operating the instrument consistent with the resources available at GRUAN sites? Is the commercial demand sufficient, and the technology available, to support the production and use of the instrument for sufficiently long for the expected multi-decade deployment within GRUAN?

- Robustness of uncertainty: Is the underlying accuracy claim for the instrument and its resultant data sufficiently robust i.e. is it likely to be able to meet the accuracy, precision and stability standards (see Section 4.1) required by GRUAN?

- Information content: Are the temporal and spatial resolution, dynamic range, and other characteristics of the measurements made by the instrument consistent with GRUAN requirements?

- Manufacturer support: Is the manufacturer committed to a process of improving the performance of the instrument,? Is the manufacturer prepared to actively participate in instrument intercomparisons? Is the manufacturer willing to disclose the necessary information required to form a fully traceable chain of sources of measurement uncertainty, given that in some cases this information may have to be kept from public display by GRUAN lay so as not to undermine the competitive advantage of the manufacturer? For a consistent uncertainty analysis it is imperative that the algorithms used for corrections within the data processing software are made available for GRUAN analaysis, but not necessarily for public display. Site location: Instrumentation may have to differ by climate region. For example, high-latitude sites exhibit extremely low water vapour contents in winter compared to equatorial sites. Therefore, instruments such as water vapour radiometers operating at 23.8 and 31.4 GHz, which have limited sensitivity for integrated water vapour amounts below 5 mm, would need to be augmented with more sensitive microwave radiometers operating near 183 GHz.

## Measurement redundancy

Having different instruments at GRUAN sites measuring the same atmospheric parameters will be invaluable for identifying, understanding and reducing systematic errors in measurements. A project within GATNDOR has been tasked with quantifying the value of redundant measurements and assessing optimal combinations of measurements. If successive reductions in measurement uncertainty with the addition of each coincident measurement from a different instrument can be quantified in a scientifically robust way, this provides a powerful justification for measurement redundancy at GRUAN sites. A case study underway within GATNDOR is using vertical profile measurements of temperature and water vapour at the GRUAN sites at Beltsville, Cabauw, Lindenberg, and Potenza to quantify the error reduction resulting from increasing redundancy of measurements. This requires an assessment of the uncertainty of the temperature and water vapour vertical profiles retrieved using each of the considered techniques and then the investigation of possible sensors’ synergies to reduce the uncertainty. The investigation will be carried out focusing on the most common instruments at the considered GRUAN sites: for temperature, radiosonde soundings and microwave profilers; for moisture, radiosonde soundings, Raman lidars, microwave profilers, and GPS receivers. The quantification of the value added by complementary observations should be assessed with respect to:

- Sensor calibration/inter-calibration (here the ARM Value Added Products could be considered as a model)

- Identification of possible biases

- Representativeness of measurements i.e. which horizontal and vertical region of the atmosphere does the measurement represent.

- Quality control/assurance with a focus on instrument performance in different meteorological conditions.

As for much of the other research underway to support the operational framework for GRUAN, this is work in progress and the true value of having multiple measurements of the same climate variables at GRUAN stations will become clear in time.

One important factor for GRUAN is that independent measurements of the same (or related) variables should be reported in a consistent way. The cross-checking of redundant measurements for consistency should be an essential part of the GRUAN quality assurance procedures. Since all data are to be reported with uncertainties, a consistency check is, in principle, a straight forward task.

## Surface measurements

While GRUAN is, by definition, an upper-air network, surface measurements at sites should also be made in such a way that:

- They are made according to WMO guidelines (WMO-no. 8), including traceability to SI standards. The CIMO classification for stations should be applied.

- The surface measurements provide ground-truthing for vertical profile measurements. For example, comparisons between ozonesonde measurements of ozone at the surface against a high precision standard provides essential information for quantifying uncertainties in the ozonesonde measurement.

- The measurements can, where relevant, constrain retrievals applied to remotely sensed profile data. Some remote sensing instruments that derive vertical profile data from e.g. optimal estimation techniques (Rodgers, 2000), are better constrained when a high precision surface measurement is included as input to the forward model used in the retrieval. In some cases remote sensing of column amounts of a trace gas can benefit from having collocated surface measurements of that trace gas e.g. as is done in TCCON.

While there are no formal requirements for GRUAN stations to include surface measurements, the guideline is that where such measurements would significantly add to the quality or utility of the GRUAN measurements, these surface measurements should be made.

## Upper-air measurements

### In-situ instruments

A discussed in Section 4.2.1, radiosondes will remain the primary workhorse within GUAN for the measurement of vertical profiles of temperature, pressure and humidity. The fact that these instruments are not recovered has important implications for GRUAN operations, viz.:

- The sensors are  not usually the limiting expense in the cost of a modern  operational radiosonde and good sensors can be obtained relatively cheaply. The exposure/mounting of the sensors on the radiosonde is a limiting factor on the performance of many radiosondes, so there is still scope for improvement with the current systems without investment in very expensive replacement technology.

- Maintaining long-term stability in a radiosonde measurement time series is challenging when the instrument being used to make the measurement is discarded after each measurement. Each instrument must be individually calibrated and tied to common calibration standards to ensure long-term stability. It must also be able to retain its performance throughout an ascent, and currently this is probably one of the limitations of the best operational radiosondes where do the systematic bias does  not appear stable to 0.1K during an ascent, see WMO (2011). The better manufacturers have managed to eliminate most faults that occur required by production engineering, but any given radiosonde type has shown small fluctuations in performance with time, when checked on the ground, although these variations in performance during flight may have been minimised by the ground check procedures used.

### Remote sensing instruments

Material to come in here from Task Teams on ‘Ancillary Measurements’ and ‘Measurement schedules and associated instrument-type requirements’

## Instrument co-location

As discussed in Section 5.3, some of the current GRUAN sites, and many potential sites, consist of instrument clusters spread over some region rather than single compact sites. Some of them are in geographical locations that have complex orography and/or heterogeneous surface characteristics. There remain open questions about how physically far apart measurements can be made and still represent a GRUAN site measurement of a single column. Therefore, appropriate collocation requirements for variables and instrumentation should be established to ensure the representativeness of measurements. These considerations should be site and parameter-specific.

## Calibration, validation and maintenance

### Instrument calibration

Establishing reliable calibration procedures for the instruments being used within GRUAN, and applying these uniformly across the network, will be an absolute prerequisite for achieving the GRUAN goals. In addition to establishing calibration procedures at individual sites that minimize the uncertainty introduced into the measurement chain (see Section 2.2) and avoid introducing discontinuities into the time series, it is equally important that calibration procedures do not compromise the goal of achieving homogeneity across the GRUAN network as a whole so that a measurement of a given parameter at one site is directly comparable to a measurement of the same variable at a different site. A guiding principal that will achieve this goal is that when two identical instruments are deployed at two different sites, they shall also use the same calibration procedures, preferably tied to the same absolute standards, and should also employ identical data processing algorithms. While achieving a common data processing for each instrument will be facilitated through processing the raw data at a single central data processing facility (see Sections 2.2.3 and 8.1), the same approach cannot be used for calibration procedures. To this end achieving inter-site homogeneity may be improved by developing travelling calibration standards which can be taken to different GRUAN stations to be used in on-site calibration or inter-comparisons. A current example of this would be Dobson Spectrophotometer #83 which is used in the NDACC and WOUDC networks to achieve homogeneity across the global Dobson network (see Sections 1.5.1 and 1.5.3).

Traceability to recognized measurement standards (e.g. SI standards) that can be reproduced globally and over long periods of time will be the key component enabling GRUAN to provide reference measurements useful for long-term climate observations. Traceability is a property of a measurement that is manifest by an unbroken chain of measurements back to a recognized standard, with fully documented uncertainty at each step. This then allows a robust calculation of the propagation of uncertainties from the fundamental standard to the final measurement. If common fundamental standards are available across the GRUAN network this will support the goal of achieving coherence across the network.

GRUAN stations shall maintain a “GRUAN site working standard” for each basis unit, e.g. a thermometer periodically calibrated to a National Metrology Institute or other accredited agency standard since this ensures traceability to an SI standard. A mechanism shall be implemented to address the compatibility of those systems with the rest of the network that may not be traceable to SI standards.

Use of traceable calibration standards will also aid operators to detect and quantify systematic errors in GRUAN measurements (see Section 3.2). Where the final data product of a reference observation depends on ancillary measurements, these measurements must again be traceable to standards. Traceability will also permit the network to incorporate new scientific insights and new technological developments, while maintaining the integrity of the long-term climate record. To achieve traceability, meta-data on all aspects relating to a measurement and its associated uncertainty shall be collected. Each station shall maintain accurate meta-data records and provide these to the GRUAN archives. Copies of calibration certificates shall be submitted to the GRUAN meta-database.

The schedule of field recalibration and validation procedures should be drawn initially from experience with a given sensor type, then refined according to the results of laboratory tests and intercomparisons. The date and nature of field recalibrations should be included in meta-data, so that if future experiments reveal shortcomings in schedules or methods that were in use, uncertainty estimates can be adjusted after the fact to reflect those newly-discovered issues.

### Instrument validation

Validation of the instruments used within GRUAN should include well documented and traceable calibration procedures, participation in regular intercomparisons with similar instruments used at other sites and/or intercomparisons with a travelling standard, and operational comparison of uncertainty estimates on the resultant measurements with those from other instruments (see Section 3.1.3). Most sites will likely not have identical instrumentation, with the result that instrument validation will likely be site specific. A standard recommendation for the use of redundant instrumentation and remote sensing instrumentation should be developed to aid site specific, regularly scheduled, instrument validation. The purpose is to make sharing and communication of best practices across sites seamless and continuous.

### Instrument maintenance

GRUAN sites are equipped with sophisticated, state-of-the-art instrumentation and should comply with strict requirements of station maintenance, exposure of instruments and calibration performance to avoid degradation of the quality of the measurements. To ensure that the goal of long-term high quality climate records is reached, site scientists who are leading experts for the instruments used at the respective GRUAN sites shall take responsibility for individual instruments operated at the GRUAN site. However, because all maintenance of an instrument can also introduce discontinuities in measurement series, maintenance shall not be conducted more frequently than is necessary. Maintenance schedules shall developed for all instruments. All maintenance actions on instruments shall be documented as part of the meta-data associated with the measurements made by that instrument.

# METHODS OF OBSERVATION

## Guiding principles

GRUAN is not being established as a network of completely new stations, and many of the initial stations within GRUAN have been in operation in some cases for decades. As a result, sites currently collecting data from different instruments will almost certainly use different data processing algorithms, different instrument pre-checks, different quality control/quality assurance etc.. While the goal is to achieve operational homogeneity with regard to these issues across the network as a whole, any changes in these procedures will need to be very carefully managed (Section 2.2) to avoid introducing discontinuities in the measurement record that would compromise detection of long-term trends. The sites comprising GRUAN will also not be supported by a single funding agency and therefore different levels of financial support at different sites are to be expected. A process for achieving convergence on agreed operations and maintenance procedures that will be applied across the network therefore needs to be developed. Furthermore, many of the initial sites report to numerous networks and their governance and stated aims differ substantially. It is therefore essential to have in place protocols and agreements, such as a Manual of Operations, including common quality assurance procedures that allow the required flexibility, whilst maintaining the fundamental quality of the observations necessary to meet GRUAN aims.

## Measurement scheduling

This section is being developed in a separate document which will then be merged with the manual once approved by the WG-ARO. The content of the latest version of that document is provided below. It must be emphasized however that this section is very much a work in progress.

### Responsibilities

  - WG-ARO shall work with the appropriate task team to define measurement schedules that allow the resultant data products capture all important scales of temporal variability, both for trend analysis and for process understanding.  These schedules should be conservative in the early stages of GRUAN, to allow the appropriate task teams to refine their studies, since currently there is a range of opinions of what is necessary.

  - When GRUAN operations have been implemented, the core measurement scheduless and associated instrument-type requirements shall be agreed bt the GRUAN Lead centre with individual sites, subject to the agreement of WG-ARO.

  - Subsequent changes to the GRUAN operations at a site will have to be agreed by WG-ARO and then implemented as far as possible, by negotiation between the Lead centre and the GRUAN sites.

  - .

  - Core measurement scheduling shall remain stable unless there is a clear requirement for change, which would then have to be agreed with the relevant GRUAN sites. Amendments to the GRUAN measurement scheduling protocol shall be submitted by the task team to the WG-ARO before being distributed to GRUAN sites for voluntary implementation.

  - Individual GRUAN sites will agree which measurements and measurement schedule they can sustain.

### Guiding principles

  - Measurements programmes at GRUAN sites are likely to be constrained by more than just the requirements of GRUAN. In recognition of the heterogeneity of the network, the scheduling protocols defined in this document may not apply at every GRUAN site, but any deviation from the measurement schedule must be agreed by the GRUAN Lead centre and then accepted by WG-ARO.

  - Some evolution of measurement schedules can be expected in the longer term when performance requirements from the network for climate studies become clearer, but changes must be limited in time and agreed by the WG-ARO, and the GRUAN sites.

  - The highest priority is that measurement schedules are established to achieve the four primary goals of GRUAN (see Section 1.2). Measurement scheduling shall be designed not only for the purposes of quantifying variability on a range of timescales and for long-term trend detection, but to fulfil all goals of GRUAN. Where perturbations to schedules would increase the utility of the measurements without compromising the primary goals of GRUAN, measurement schedules should be adapted to meet the needs of other end-users e.g. the timing of a daily measurement may be shifted to coincide with a satellite overpass and in this way provide valuable high quality data for satellite validationWhere possible, measurements schedules for redundant systems should be synchronized so as to avoid sampling biases when combining the measurements into a single data product..

  - Required measurement schedules may vary regionally and seasonally. In places and seasons where the parameter being measured is more variable, measurements should be made more frequently so that the effects of that variability can be accounted for in trend analyses.

  - Measurement schedules will have been set to permit a statistical separation of the different drivers of changes in the observed variable.A first step will have been to assess how the sampling affects uncertainties on derived monthly means, followed by a second step which determines how those monthly mean uncertainties affect the statistical robustness of trends derived from those monthly meansModels such as atmosphere-ocean general circulation models or coupled chemistry-climate models will have a useful role to play in guiding measurement scheduling. They can be used to provide initial estimates of the autocorrelation, the magnitude of variability, and the trend in climate variables and the composition of the atmosphere as a function of location and season. Simulating the effects of a measurement schedule by sampling model output on the same schedule as the measurements can provide an indication of how the proposed measurement schedule is likely to affect the determination of variability on a range of timescales as well as long-term trends. Where model output is not available, analysis of temperatures, e.g. from reanalyses, can often serve as a valid proxy for other climate variables since temperature responds to climate variability in a similar way to many other climate variables. As this work develops it is possible that the initial GRUAN network  measurement protocols  may change..

  - {For some measurements, scheduling with respect to UTC or Local Solar Time (LST) may be important and may result in conflicting requirements regarding different intended uses of the measurements. For example, scientifically it may be advantageous to have all GRUAN sites making measurements at the same LST (especially for variables that show strong diurnal variations or for instruments that have diurnally dependent biases that we wish to minimize), while for ensuring coincidence with GUAN stations having all measurements made at the same UTC might be more appropriate.

  - Current assimilation schemes used in NWP and reanalysis centres, e.g. 4D-var assimilation, are less sensitive to measurements being made at the same time each day compared to assimilation schemes used in the past. Therefore, consistent timing of measurements is not an issue for assimilation into NWP or for reanalysis. If, however, the variable being measured shows a strong diurnal cycle, or if the instrument being used to make a measurement has diurnally varying biases, changing measurement times would introduce additional variability which would need to be accounted for in any analysis in order to avoid sampling bias.}

### Factors affecting measurement scheduling for trend detection

For trend detection four factors should guide the development of measurement schedules (Weatherhead et al., 1998), viz.:

  - The magnitude of the variability. In most cases this will vary as a function of season. Where measurements through a month are sparse, and where monthly means of those measurements will be used in trend analysis, the day-to-day variability within the month will determine the representativeness of the sparse measurements in quantifying the true monthly mean. The variability in the monthly means themselves, or, more precisely, the variability in the monthly residuals after a regression model fit, will also determine the statistical robustness of derived trends.

  - Autocorrelation. This is a measure of the ‘momentum’ or ‘latency’ in the system. When autocorrelation is high, measurements in consecutive time periods are highly correlated. When autocorrelation is low, the signal is very noisy and consecutive measurements are largely independent of each other. Autocorrelation is also likely to vary through the year. If monthly means of measurements are being used in trend detection, the auto-correlation between those monthly means constitutes an important source of uncertainty in the trend estimate (Tiao et al., 1990). One advantage of using monthly means for the calculation of long-term trends is that the uniform temporal sampling simplifies the calculation of the autocorrelation in the signal. However, individual measurements may also be used in trend detection and methods are available for determining the autocorrelation in such potentially unequally spaced measurement time series (Bodeker et al., 1998). A clear distinction must be made between:

    - Day-to-day autocorrelation which determines, in part, the likelihood of over-sampling,

    - Day-to-day variability which determines the robustness of the monthly mean when it is calculated from sparse, isolated measurements through the month,

    - Autocorrelation in the monthly means which determines, in part, the uncertainty on calculated trends,

    - Variability in the monthly means which also contributes to the uncertainty on calculated trends.

All four of the above are likely to vary spatially and seasonally with the result that optimal measurement schedules are likely to vary between sites and with season. On initiation of a measurement programme, and where the autocorrelation is not known a priori, measurements should be made at the highest possible frequency so that a robust seasonal pattern of the autocorrelation can be established. Thereafter, measurement frequency can be relaxed during periods of expected high autocorrelation since momentum in the system being sampled will result in nearby measurements not being independent.

  - The precision of the measurement. When very precise measurements can be made, measurement frequencies can potentially be reduced, depending on the extent to which measurement precision is a factor in trend detection or in analyses of specific atmospheric phenomena. When measurement precision is poor, high frequency sampling is required to reduce the effects of the random errors on the measurements. Measurement precision can also vary with season as a result of confounding factors (such as surface albedo, humidity and temperature) which vary through the year. The derivation of measurement uncertainties within GRUAN, and how these might vary with season, must therefore play a role in determining measurement scheduling.

  - The size of the expected trend to be resolved. For large trends, the signal to noise ratio will be high and measurement frequencies can be reduced (all else being equal). The magnitude of the trend is also likely to be a function of season.

Where the precision of the measurement is a significantly smaller contributor to the uncertainty in the trend estimate than autocorrelation and natural variability, for most mid-latitude locations and for most climate variables, the autocorrelation in the system results in diminishing returns for measurements made at a frequency of higher than every 3 days. On the other hand, for most climate variables measured at mid-latitudes, sampling less frequently than every 10 days significantly increases the uncertainty on derived monthly means.

The interplay between the four factors discussed above must be accounted for when planning measurement schedules. It may be that the uncertainty on derived trends is limited by natural variability rather than by the precision of the instrument, in which case more resources should be invested in increasing measurement frequency rather than increasing measurement precision. In some cases this may require a cost-benefit analysis where the cost to detect a putative trend of X%/decade (perhaps based on projections from models) over N years is minimized. A cheaper instrument making a less precise but more frequent measurement might be selected over a more expensive instrument making a more precise but less frequent measurement, since the greater frequency leads to detection of the expected trend either in fewer years or at a lower cost. A measurement strategy might have a greater cost per year than any alternative, but if that strategy can detect a statistically robust trend in fewer years, the net cost may be reduced.

### Measurement schedules

Mandatory schedules are not defined for any of the measurement programmes within GRUAN. Rather, once a station has selected the frequency with which measurements will be made, this section provides guidelines on appropriate timing of those measurements. The frequency of measurements at sites will determine, in part, the added value that a site brings to the network (see Section 5.3).

Radiosondes

- For sites performing four radiosonde flights daily: Flights either at 00, 06, 12 and 18 UTC or at 00, 06, 12 and 18 LST.

- For sites performing twice daily radiosonde flights: One flight at 00LST and one flight between 06LST and 18LST timed to maximize coincidence with any satellite-based instrument overpass measuring the same variables. Since satellite-based measurements are more likely to be daytime measurements, the daytime radiosonde launch time is the one which is varied.

- For sites performing daily radiosonde flights: One flight at 12LST.why??

- For sites performing weekly radiosonde flights: Nominal launch times should be 12LST on the same day of the week, but allowed to vary by up to 48 hours either side to match satellite overpasses.

- For sites performing monthly radiosonde flights: Nominal launch times should be 12LST on the same day of the month, but allowed to vary by up to 5 days either side to match satellite overpasses.

Frost point hygrometers {Ed Note: No site is likely to fly a frost point hygrometer more often than once per week, so the measurements will be made at 7 day intervals, but in fact most sites will only be able to make one flight per month, if at all, so the measurements will be once per month. Does this discussion add to this basic fact?}

- For sites performing an average of N flights per week i.e. N×52 flights per year: Where the seasonal cycle in natural variability is not yet known, intervals between flights should be constrained by (4/N) < t < (10/N) where t is the interval in days. Once a climatology of the seasonal cycle in natural variability has been determined, during the 5 months of the year exhibiting highest natural variability, intervals between flights should be constrained by (3.5/N) < t < (6.5/N) where t is the interval in days; this should result in a total of N×30 flights through those 5 months. For the remaining 7 months of the year intervals between flights should be constrained by (7/N) < t < (13/N); this should result in a total of N×22 flights through those 7 months.

- For sites performing an average of N flights per month i.e. N×12 flights per year: Where the seasonal cycle in natural variability is not yet known, intervals between flights should be constrained by (20/N) < t < (40/N) where t is the interval in days. Once a climatology of the seasonal cycle in natural variability has been determined, during the 4 months of the year exhibiting highest natural variability, intervals between flights should be constrained by (15/N) < t < (25/N) where t is the interval in days; this should result in a total of N×6 flights through those 4 months. For the remaining 8 months of the year, intervals between flights should be constrained by (35/N) < t < (45/N); this should result in a total of N×6 flights through those 8 months.

Aerosol sondes

- For sites performing an average of N flights per week i.e. N×52 flights per year: Where the seasonal cycle in natural variability is not yet known, intervals between flights should be constrained by (4/N) < t < (10/N) where t is the interval in days. Once a climatology of the seasonal cycle in natural variability has been determined, during the 5 months of the year exhibiting highest natural variability, intervals between flights should be constrained by (3.5/N) < t < (6.5/N) where t is the interval in days; this should result in a total of N×30 flights through those 5 months. For the remaining 7 months of the year intervals between flights should be constrained by (7/N) < t < (13/N); this should result in a total of N×22 flights through those 7 months.

- For sites performing an average of N flights per month i.e. N×12 flights per year: Where the seasonal cycle in natural variability is not yet known, intervals between flights should be constrained by (20/N) < t < (40/N) where t is the interval in days. Once a climatology of the seasonal cycle in natural variability has been determined, during the 4 months of the year exhibiting highest natural variability, intervals between flights should be constrained by (15/N) < t < (25/N) where t is the interval in days; this should result in a total of N×6 flights through those 4 months. For the remaining 8 months of the year, intervals between flights should be constrained by (35/N) < t < (45/N); this should result in a total of N×6 flights through those 8 months.

Ozonesondes

- For sites performing an average of N flights per week i.e. N×52 flights per year: Where the seasonal cycle in natural variability is not yet known, intervals between flights should be constrained by (4/N) < t < (10/N) where t is the interval in days. Once a climatology of the seasonal cycle in natural variability has been determined, during the 5 months of the year exhibiting highest natural variability, intervals between flights should be constrained by (3.5/N) < t < (6.5/N) where t is the interval in days; this should result in a total of N×30 flights through those 5 months. For the remaining 7 months of the year intervals between flights should be constrained by (7/N) < t < (13/N); this should result in a total of N×22 flights through those 7 months.

- For sites performing an average of N flights per month i.e. N×12 flights per year: Where the seasonal cycle in natural variability is not yet known, intervals between flights should be constrained by (20/N) < t < (40/N) where t is the interval in days. Once a climatology of the seasonal cycle in natural variability has been determined, during the 4 months of the year exhibiting highest natural variability, intervals between flights should be constrained by (15/N) < t < (25/N) where t is the interval in days; this should result in a total of N×6 flights through those 4 months. For the remaining 8 months of the year, intervals between flights should be constrained by (35/N) < t < (45/N); this should result in a total of N×6 flights through those 8 months.

Raman lidar

Material still to come here from experts on this measurement system.

GPS integrated precipitable water

Material still to come here from experts on this measurement system.

Microwave profiler

Material still to come here from experts on this measurement system.

Microwave radiometer

Material still to come here from experts on this measurement system.

Fourier Transform Spectrometer

Material still to come here from experts on this measurement system.

Ceilometer

Material still to come here from experts on this measurement system.

Cloud radar

Material still to come here from experts on this measurement system.

Wind profiler

Material still to come here from experts on this measurement system.

SODAR/RASS

Material still to come here from experts on this measurement system.

Doppler lidar

Material still to come here from experts on this measurement system.

Aerosol or ozone lidar

Material still to come here from experts on this measurement system.

Radar

Material still to come here from experts on this measurement system.

## Operation and maintenance, quality standards

Standards of operation and maintenance for each instrument used in GRUAN should be developed to ensure that minimum quality standards are achieved. This will be necessary to minimize sources of error when measurements are being made using sophisticated instruments that may not always be completely familiar to the operator. This will be more likely the case when measurements are being made under operational conditions. Operation and maintenance protocols should be such that collection of detailed meta-data is mandatory as these meta-data will be vital to establishing measurement uncertainties.

# DATA MANAGEMENT

## Overview of GRUAN data flow

A schematic representation of the flow of data within GRUAN and from GRUAN to the user community is shown in Figure 3.

Figure 3:A schematic representation of the flow of data in GRUAN. Blue arrows show the standard flow of data. The red arrows show the flow of near-real time data. Data provided to end-users via red routes are not ‘GRUAN data’. Different data exchange protocols should operate for exchange of data within GRUAN (shaded green region) and from the GRUAN external data archive to end-users.

The GRUAN Data Management Manual defines four data levels, viz.:

Level 0 (L0): Original raw data. This is the ‘rawest’ form of data available e.g. measured voltages before any processing has been applied. Even for the same instrument, formats of L0 data files are likely to differ between sites.

Level 1 (L1): Converted raw data. These data are stored in a common well-described file format intended for long-term storage.

Level 2 (L2): This is a standard GRUAN data product that does not require incorporation of independent measurements.

Level 3 (L3): This is also a standard GRUAN data product but an integrated product that does require incorporation of independent measurements.

Measurements and meta-data are bound together in each of these 4 levels of data. L0 data are ingested from all GRUAN sites into the internal GRUAN data archive hosted at the Lead Centre (see Section 8.5). Direct exchange of L0 data between sites is discouraged since this circumvents the data versioning protocols and reduction of the raw data to a the common L1 file format. Similarly, direct exchange of L1 data between sites is discouraged since this circumvents network wide application of calibration techniques, and other algorithms applied to convert L0 to L1 data that would be implemented either at the Lead Centre or at a centralized GRUAN data processing site (see below). L0 data are expected to be archived in perpetuity at the measurement site to provide a failsafe backup while L1 data are expected to be archived in the internal GRUAN data archive at the Lead Centre.

Where GRUAN sites have agreed to the near-real time release of their data, these data will be made immediately available via the WIS. This will require some local site-based processing of the L0 data to create data suitable for submission to the WIS.

Processing of the L1 data held in the GRUAN internal data archive to produce L2 and L3 GRUAN data products will occur either at the Lead Centre or at a GRUAN station that specializes in processing data for a particular instrument. This processing would include applying the necessary recalibrations, corrections, and the uncertainty analysis in a consistent and traceable manner across identical instruments from different sites. The L2 and L3 data, including its meta-data and documentation, are provided to the user community through the external GRUAN data archive hosted at NCDC. A performance monitoring process (see Section 9), implemented at the Lead Centre, will provide feedback on performance to individual sites.

## GRUAN data policy

This section summarizes and expands on the GRUAN data policy document prepared by the GCOS secretariat. Since GRUAN is co-sponsored by WMO it is appropriate that any policy for release and dissemination of GRUAN data complies with WMO policy, practice and guidelines for the exchange of meteorological and related data and products. Specifically GRUAN data dissemination and use should comply with WMO Resolution 40 (Cg-XII) which calls for free and unrestricted international exchange of meteorological data and related data and products. Because most GRUAN measurements are considered ‘essential’ in the context of Resolution 40, they are required to be exchanged without charge and with no conditions on their use. GRUAN stations are likely to be providing data to other networks which may have policies in place to protect the rights of the data providers to their own data. No conflict arises here since the data being provided through other networks are not ‘GRUAN’ data and are therefore not subject to the requirements of Resolution 40.

Three levels of exchange of GRUAN data should be recognised:

- Exchange of data within the GRUAN community. This should always occur through the GRUAN Lead Centre so that the exchange can be controlled by data policies developed specifically for internal exchange of GRUAN data.

- Dissemination of GRUAN products to end-users. This should always occur through the official GRUAN data centre (see Section 8.6). A different policy should be implemented to control the dissemination of GRUAN data at this level. The only exception to disseminating GRUAN data through the lead centre would be when ‘pre-GRUAN’ data (see below) are submitted in near real-time to centres such as ECMWF.

- Dissemination of pre-GRUAN data on the WMO GTS/WIS for assimilation in NWP simulations.

A distinction should be made between 'standard data' and 'enhanced or experimental data' obtained at GRUAN sites:

- Standard data (e.g., near surface synoptic observations, radiosonde observations) have general exploitation value, common measurement technology, generally well understood, and few problems with data interpretation.

- Enhanced or experimental data (e.g., Raman LIDAR, microwave radiometer, surface radiation, GPS precipitable water) have high exploitation value, sophisticated measurement technology and/or of experimental nature, would recommend contact to site scientist for correct interpretation of data, and would require considerable efforts to maintain continuous measurements and high quality of the data.

Enhanced or experimental data are more likely to be subject to limitations on dissemination than standard data.

The primary goals of GRUAN (see Section 1.2) are not consistent with near real-time dissemination of measurements made at GRUAN sites. Generating high precision, high quality measurements with well characterized uncertainties takes a significant investment of time and effort. In GRUAN the emphasis is clearly on providing reference quality measurements rather than providing near real-time measurements. However, it is recognized that GRUAN measurements are likely to be very useful to a number of users requiring data in near real-time e.g. for initializing NWP models. Therefore, where possible, and where it does not detract from achieving the primary goals of GRUAN, GRUAN sites should submit real-time data to end-users via the GTS/WIS. These, however, are not termed ‘GRUAN’ data since they would not necessarily have been subjected to the stringent QA/QC procedures that are core to GRUAN’s operation. Rather they are what might be termed ‘pre-GRUAN’ data. The measurements for which real-time submission may be valuable are also more likely to be 'standard data' as described above. The WIS requirements, e.g. on meta-data, and the transmission of near-real time data via the GTS is strongly encouraged but is not considered a mandatory requirement for GRUAN sites (see Section 5.2). This decision to exclude near real-time submission of GRUAN data from the list of mandatory requirements for a GRUAN site is consistent with the recommendation of the AOPC who at their XIVth session stated in recommendation #29 ‘AOPC recommended that GRUAN data policy should request sites to provide all data in a free and unrestricted manner (in accordance with WMO Resolution 40 (Cg-XII)), and if possible [our emphasis] in real time, in order to be of maximum value for all applications’. Where sites do not currently have the infrastructure or expertise in making such submissions, assistance from WMO should be obtained in the form of hardware and/or training. There may be advantages to submitting data in near real-time since data assimilation algorithms are able to flag data that appear to be statistically anomalous. If such two way communication can be established between GRUAN and the NWP/data assimilation community, such information could form an important part of the measurement meta-data (Section 10). Near real-time release of standard GRUAN data will also facilitate the quality control link between GRUAN and GUAN.

When GRUAN data are used in a scientific publication, the origin of the data must be acknowledged and referenced. A minimum requirement is to reference GRUAN as a reference network of GCOS and to acknowledge the GRUAN data archive at NCDC as the source. If data from only one GRUAN site (or a limited number of sites) have been used, additional acknowledgement of those site(s) and their sponsoring institutions or organizations must be given, as specified in the meta-data associated with the data files.

Inclusion of GRUAN scientists as co-authors on papers making extensive use of GRUAN data (and in particular enhanced or experimental data) is justifiable and highly recommended, in particular if a site scientist has responded to questions raised about data quality and/or suitability for the specific study in question, or has been directly involved in contributing to the paper in other ways. Co-authorship should not be a pre-condition for release of GRUAN data. However, for enhanced or experimental data it is highly recommended that data users invite site scientists to become co-authors on resultant publications, or determine whether an acknowledgement would be sufficient. Users of enhanced or experimental GRUAN data should be encouraged to establish direct contact with site scientists for the purpose of complete interpretation and analysis of data for publication purposes. GRUAN meta-data should include all information related to acknowledgements and/or co-authorship on publications making use of the data.

## Collation of meta-data

Two different types of meta-data need to be accommodated within the GRUAN data management facilities, viz.:

- Meta-data describing the context in which the measurement was made i.e. the calibration procedures, data processing algorithms employed, traceability to SI standards, log books, etc.. This information will be relevant to a set of data and not specific to any particular datum.

- Meta-data associated with each datum. For example, for point source measurements, as opposed to column or partial column measurements, in addition to the measurement uncertainty associated with that datum, meta-data such as the exact date and time associated with the datum, as well as the exact altitude, latitude and longitude must be directly available or easily derivable from other meta-data. The provision of such meta-data recognises the fact that e.g. balloon-borne instruments drift in latitude and longitude during a flight. These data can only be used in 4D-Var assimilation of they are tagged with their 4D (time, latitude, longitude, altitude) coordinates.

## Data format

In the same way that a distinction should be made between the distribution of data within the GRUAN community and the dissemination of GRUAN data to end-users, a distinction should be made with regard to prescribed data formats for these two different aspects of data distribution, viz.:

- For distribution of data within GRUAN the emphasis should be on expediency. Different data formats for different instruments should be permitted and not discouraged. Whatever format facilitates quick and automated processing of data and its associated meta-data should be used.

- For dissemination of GRUAN data to clients, a format should be selected that is flexible enough to allow a common format across all GRUAN products, should have an existing large user-base in the client community, should easily allow the inclusion of meta-data in each data file, should be an open format/standard that requires no licensing, should be self-describing, and should have a large suite of readily available tools for manipulating the data files. Perhaps the most suitable format would be NetCDF and better still CF (Climate and Forecast) compliant NetCDF. Tools such as NCO (NetCDF operator) should then be made available for manipulating these files.

## Data submission

If sites elect to submit near real-time data to end-users, this should be done directly through the WIS or through their own portals but without a GRUAN label attached as the product does not have a robustly quantified uncertainty estimate. Otherwise all data from GRUAN sites should flow through the Lead Centre. The expectation might be that GRUAN sites submit their raw data to the GRUAN Lead Centre as soon as possible after the measurement but with the policy in place that these data will not be made available outside of the GRUAN community at this time. A facility for imposing time limits on making the data available to the end-user community for different stations should be implemented as this does not contravene WMO Resolution 40 (Cg-XII). In this way stations are more likely to be willing to make their raw data immediately available within the GRUAN community without compromising their rights to first publication of the data (some funding agencies may even insist that such a data policy is in place).

Procedures for submitting data and meta-data from GRUAN sites to the GRUAN archive should be developed in such a way as to minimize the effort required at the GRUAN sites and to harmonize the process of data collection and data quality control across the network as a whole. For example, submission of data to the GRUAN archives can be easily automated if the mode of submission is through FTP to a server based at the Lead Centre, whereas if submission must be done through a web portal this cannot be easily automated and is likely to be very time consuming for individual GRUAN sites.

Where data submission tools can be developed centrally (e.g. at the Lead Centre) and distributed for use to GRUAN sites to facilitate data submission to the GRUAN archives, this is preferable to each site independently developing such tools. The ability for sites to jointly contribute to supporting such network wide activities would be desirable.

## Data dissemination

Dissemination of GRUAN data products to end-users/customers shall occur through an official GRUAN data Centre hosted at NCDC. Access to GRUAN data through a single source will reinforce the model that GRUAN data are homogeneous both in time and across GRUAN stations.

For climate research in particular it is important that users of climate data can, if required, obtain complete information on how the data they are using were acquired. Therefore, users of GRAUN data shall have access not only to the measurements and their uncertainties, but also to the instrument, operating procedures, data reduction algorithms used, and to when changes to any of these occurred through the complete time period of the data set.

A facility should be established whereby users of GRUAN data products can voluntarily register their use of the data. This would:

- Allow the Lead Centre to maintain statistics on data usage. This would be useful when applying for funding to support GRUAN operations.

- Allow users of data to be informed if and when newer versions of the data become available.

- Facilitate reporting of potential errors/anomalies in the data by end-users.

Such a facility might need to exist independently of the GRUAN NCDC archives to avoid legal issues related to data retention by US government agencies.

As discussed above, GRUAN sites are likely to also be members of other networks and are likely to submit data to end-users through other network's archives. Data submitted through a non-GRUAN networks may be subject to different data processing, different QA/QC procedures, and different calibrations resulting in a data product that is different to the GRUAN product. This is not seen as a problem since the product delivered through other networks is not identified as ‘GRUAN’ data.

Users of GRUAN data need to know the version of any dataset they are using and whether newer versions might be available. The names of data files must therefore include the data version identifier to facilitate easy identification of the data version. An application to periodically check for updates of GRUAN data files found on a client computer with the database at NCDC needs to be developed.

## Data archiving

GRUAN does not necessarily need to build its own data archive and user interface. This is a rather costly operation for any large network and partnering with an established data archive such as NCDC with a user-friendly interface should be preferred. Because data cannot be quality assured or corrected in near real-time, additional processing steps and uncertainty estimate assignment will be required. This key processing will be allowed to grow, and thus, data versioning will be required. It is important that the GRUAN archive includes all previous versions of any given data set so that analyses using previous versions of data can be repeated if required.

## Quality control at the instrument/site level

Part of the data management within GRUAN includes feedback to the sites in the form of reports on data submission, data quality, and comprehensiveness of meta-data submitted. Existing algorithms, potentially supplemented by future algorithms to be developed, shall be used operationally to identify systematic errors, anomalies or instrumental issues. Results of such tests shall be communicated back to GRUAN sites on short timescales so that remedial action can be taken if required. Following the example of the ARM Data Quality Office, communicating quality control results to GRUAN site operators and engineers will facilitate improved instrument performance and thereby minimize the amount of unacceptable data collected.

# POST-PROCESSING ANALYSIS AND FEEDBACK

This section is still under development and far from complete.

Analysis of GRUAN data products by end-users will need to be sensitive to data versioning. As new knowledge becomes available and data are reprocessed as a result, newer versions of data sets will be provided through the GRUAN archives and end-users need to be aware of such updates and, if necessary, repeat their own analyses. Users of GRUAN data must always document the version of data used to ensure that the analyses can be independently replicated. Key to this process will be the ability to make users aware of updated versions of data sets that they previously accessed, now becoming available. The data processing centre, either the Lead Centre or the designated GRUAN site specializing in processing of that particular data set, should be tasked with data version control and ensuring that the necessary meta-data on data versions are made available to end-users.

Inevitably, algorithms change and errors in data processing occur that are not necessarily apparent until the data are used. Therefore, a facility that allows data users to report potential bugs or anomalies found in data during analyses of the data needs to be designed and implemented. This might be modelled on the ARM Program Climate Research Facility bug reporting system.

A quality system should include procedures for feeding back into the measurement and quality control process to prevent the errors from recurring. Quality assurance can be applied in real-time post measurement, and can feed into the quality control process for the next process of a quality system, but in general it tends to operate in non-real time.

Additional points to be covered in this section include

Automation and fast feedback.

Quality management is a large part of post-processing (see Section 10)

# QUALITY MANAGEMENT

This section is still under development and far from complete.

This chapter defines the principles and the methodological framework for GRUAN operations, and details how activities will be coordinated to manage and control data quality within GRUAN. Quality management within GRUAN consists of quality assurance and quality control. Quality assurance focuses on providing confidence that quality requirements will be fulfilled and includes all the planned and systematic activities implemented in a quality system so that quality requirements for a product or service will be fulfilled. Quality control is associated with those components used to ensure that the quality requirements are fulfilled and includes all the operational techniques and activities used to fulfil quality requirements. The GRUAN quality management policy is to achieve a level of data quality that allows the primary goals of GRUAN (see Section 1.2) to be met for all potential users of GRUAN data products.

Assuring the quality of the GRUAN data begins with a robust process of describing, quantifying and validating all sources of uncertainty in all GRUAN measurements and by providing rich metadata that describe all facets of the measurement process. Where total measurement uncertainties lie below some prescribed threshold this increases confidence in the quality of the GRUAN data. The use of redundant measurements, as described in Section 3.1.3, also serves to assure the quality of the GRUAN data products. Agreement of two independent measurements, preferably based in different measurement principles, provides a high degree of confidence that no significant systematic effect was disregarded and uncertainties were not under-estimated. Laboratory tests and intercomparisons are fundamental methods for establishing and confirming uncertainty estimates for GRUAN data products. Laboratory tests provide an opportunity to investigate in detail the performance of instruments under controlled conditions and to measure differences against certified references or other standards. Data from these experiments can be used to detect biases that may be corrected for and to determine calibration uncertainties. Field intercomparisons allow multiple in situ sensors and remote sensing data to be directly compared under the actual atmospheric conditions of the required measurement, including the complex environmental conditions (temperature, humidity, pressure, wind/flow rate, radiation, and chemical composition) that cannot be fully reproduced in the laboratory. These complementary activities increase confidence that measurements are subject to neither unanticipated effects nor undiscovered systematic uncertainties. Therefore field experiments are particularly useful for assuring the quality of GRUAN data products. The use of GRUAN data in data assimilation also adds to the assurance of GRUAN data quality since the measurements, with their uncertainties, can be tested for comparability with the data assimilation model values in a 4D-Var assimilation setting within the known internal variability of the system.

Quality control will be achieved through the application of the various measurement protocols defined in this manual and in related measurement system guides. Visual inspection of all data by science/instrument experts will be required for all instruments to minimize issues that slip through automated routines. The Lead Centre shall coordinate this effort, which shall be distributed across different GRUAN sites. As outlined in Section 3.1.3, vertically resolved uncertainty estimates, prepared independently for each site, will be used as a metric to compare the site-to-site quality of the observations.

Section 4.2 of this manual provides explicit requirements regarding random errors, bias, stability, resolution and representativeness for measurements made within GRUAN. Minimizing cost without compromising quality is also an implied or explicit requirement for measurements made within GRUAN. The purpose of quality management is to ensure that GRUAN data meet the requirements in terms of uncertainty, resolution, continuity, homogeneity, representativeness, timeliness, format etc. for their intended use, at a minimum practicable cost. GRUAN recognizes that all measurements are imperfect, but, if their quality is known and demonstrable, they can be used appropriately.

Quality management is required at all points in the measurement process from network planning and training, through installation and station operations to data transmission and archiving. This quality management must include feedback and follow-up provisions across a range of timescales from near real-time to annual reviews. Because of the emphasis on the provision of robust measurement uncertainties and the associated requirement for in-depth quality management, the resources required with GRUAN will likely be a significant cost of operating the network, and very likely more that the few percent overall costs typical of many observational networks. However, without this expenditure, the quality of the data will be unknown, and their usefulness diminished.

A key aspect of quality management within GRUAN will be fulfilling customer requirements. To this end systems shall be developed to:

- Inform users of GRUAN products of changes in measurements systems at specific stations.

- An incident reporting system that can flag data anomalies to users.

- Inform users of the availability of updates to previously accessed data products.

- Provide ‘help desk’ support to users of GRUAN data products.

Establishing close working relationship with instrument manufacturers will also be central to quality assurance within GRUAN.

A common component of quality assurance is quality monitoring or performance monitoring, a non-real-time activity in which the performance of the network or observing system is examined for trends and systematic deficiencies. Performance monitoring within GRUAN will primarily be the responsibility of the Lead Centre, but where other specialists may be co-opted to assist in performance assessments. The outcomes of recertification of GRUAN sites (see Section 5.4) and GRUAN site audits (see Section 5.5) will be an essential component of performance monitoring. Requests for external, independent assessments of GRUAN performance from key user groups of GRUAN data products might also serve a useful performance monitoring function. The development of quantitative performance indicators such as:

- Data downloads,

- The number of peer reviewed publications in which GRUAN data have been used,

- Scientific case studies of the added value resulting from the use of GRUAN data products,

may serve to provide year-to-year traceability of GRUAN’s impact on the climate community.

# ACRONYMS

ARM: Atmospheric Radiation Measurement programme

ACRF: ARM Program Climate Research Facility

AOD: Aerosol Optical Depth

AOPC: Atmospheric Observation Panel for Climate

CBS: WMO Commission for Basic Systems

CDR: Climate Data Record

CIMO: WMO Commission for Instruments and Methods of Observation

GATNDOR: GRUAN Analysis Team for Network Design and Operations Research

GCOS: Global Climate Observing System

GHG: Well-mixed greenhouse gas (CO2, CH4, N2O, CFCs, HFCs, PFCs, SF6, etc.)

GNSS: Global Navigation Satellite System

GOS: Global Observing System

GRUAN: GCOS reference upper air network

GSICS: Global Space-Based Intercalibration System

GTS: Global Telecommunication System

GUAN: GCOS upper air network

ICM: Implementation - Coordination Meeting (GRUAN)

ISCCP: International Satellite Cloud Climatology Project

NCDC: NOAA National Climate Data Centre

NMS: National Meteorological Service

NOAA: National Oceanic and Atmospheric Administration

NWP: Numerical Weather Prediction

PDF: Probability Distribution Function

RMS: Root Mean Square

TCCON: Total Carbon Column Observing Network

UT/LS: Upper troposphere/lower stratosphere

WIS: WMO Information System

WWW: World Weather Watch

# Appendix A – Expanded details on additional GRUAN Essential Climate Variables

### A.1.	Wind speed (priority 2)

The high accuracy of 0.5 m/s prescribed for wind speed is needed to delineate calm conditions from light winds.

### A.2.	Wind direction (priority 2)

No supplementary comments yet.

### A.3.	Ozone (priority 2)

During a discussion at the ICM-2 meeting, it was suggested that ozone should develop into a priority 1 variable for GRUAN. The consensus appears to be that it remains a priority 2 variable.

### A.4.	Methane (priority 2)

No supplementary comments yet.

### A.5.	Net radiation (priority 2)

The prescribed precision and accuracy values of 5 W/m2 match the requirements for the BSRN network.

### A.6.	Incoming short-wave radiation (priority 2)

The stated measurement range of 0 to 2000 W/m2 exceeds the solar constant (1366 W/m2) but is required since in the presence of partly cloudy skies and when the sub is not obscured by cloud, reflections off clouds can enhance surface short-wave radiation significantly. The prescribed precision and accuracy values of 3 and 5 W/m2 respectively, match the requirements for the BSRN network.

### A.7.	Outgoing short-wave radiation (priority 2)

The prescribed precision of 2 W/m2 and accuracy of 3% match the requirements for the BSRN network.

### A.8.	Incoming long-wave radiation (priority 2)

The prescribed precision and accuracy values of 1 and 3 W/m2 respectively, match the requirements for the BSRN network.

### A.9.	Outgoing long-wave radiation (priority 2)

The prescribed precision and accuracy values of 1 and 3 W/m2 respectively, match the requirements for the BSRN network.

### A.10.	Radiances (priority 2)

The stated stability requirement of 0.03%/decade is achievable through SI traceability. The precision and accuracy requirements of 0.01% and 0.15% respectively are applicable for mean seasonal radiances at ~1000 km spatial scale.

### A.11.	Aerosol optical depth (priority 2)

Measurements of all aerosol parameters should be spectrally resolved. The aerosol optical depth is the most important of the aerosol parameters. While the other aerosol parameters will scientifically useful if the aerosol optical depth is large, when the aerosol optical depth is small, measurements of other aerosol parameters become less valuable.

### A.12.	Aerosol total mass concentration (priority 2)

Size-fractionated measurements are required.

### A.13.	Aerosol chemical mass concentration (priority 2)

Size-fractionated measurements are required.

### A.14.	Aerosol light scattering (priority 2)

Size-fractionated and spectral measurements are required.

### A.15.	Aerosol light absorption (priority 2)

Size-fractionated and spectral measurements are required.

### A.16.	Cloud amount/frequency (priority 2)

The prescribed precision and accuracy ranges of 0.1%-0.3% result from cloud variations of 1-3% found in the ISCCP database. The prescribed long-term stability requirement of 0.1%-0.2% results from the 1-2%/decade trends found by Norris (2005).

### A.17.	Cloud base height (priority 2)

The prescribed measurement range of 0-20 km (1000-50 hPa) is consistent with the vertical cloud range found in Rossow and Schiffer (1999). The prescribed precision and accuracy of 100 m (10-40 hPa) is consistent with variations derived from the ISCCP database. The long-term stability requirement of 20m/decade is what would be required to detect the trend in global mean cloud base height of 44 m/decade reported by Chernykh et al. (2001).

### A.18.	Cloud layer heights and thicknesses (priority 2)

The prescribed vertical resolution of 50 m is required to resolve cloud layer thickness of ~30 m for cirrus clouds and is easily achievable with a lidar based system (Winker and Vaughan, 1994).

### A.19.	Carbon Dioxide (priority 3)

This ECV was not included in Appendix 1 of GCOS-112 but is key to understanding trends in tropospheric stratospheric temperatures and so is included here.

### A.20.	Cloud top height (priority 3)

Cloud top height measurements are also important for radiosonde temperature uncertainty analysis. When a radiosonde emerges into dryer air above a cloud ,evaporation of the condensed water cools the sensor and creates a cool bias in this region. This effect can lead to deviations up to 1K above a cloud and the data need to be flagged appropriately, e.g., by assigning a correspondingly

increased uncertainty to data in such regions.

### A.21.	Cloud top pressure (priority 3)

No supplementary comments yet.

### A.22.	Cloud top temperature (priority 3)

No supplementary comments yet.

### A.23.	Cloud particle size (priority 4)

No supplementary comments yet.

### A.24.	Cloud optical depth (priority 4)

No supplementary comments yet.

### A.25.	Cloud liquid water/ice (priority 4)

No supplementary comments yet.

# REFERENCES

Alexandersson H. and A. Moberg, Homogenization of Swedish temperature data. Part I: homogeneity test for  linear trends, Int. J. Climatol., 17, 25-34, 1997.

Austin J., J. Wilson, F. Li and H. Vömel, Evolution of Water Vapor Concentrations and Stratospheric Age of Air in Coupled Chemistry-Climate Model Simulations, J. Atmos. Sci., 64, 905-921, 2007.

Bodeker G.E., I.S. Boyd, and W.A. Matthews, Trends and variability in vertical ozone and temperature profiles measured by ozonesondes at Lauder, New Zealand: 1986-1996. J. Geophys. Res., 103 (D22), 28661-28681, 1998

Bodeker G.E., J.C. Scott, K. Kreher, and R.L. McKenzie, Global ozone trends in potential vorticity coordinates using TOMS and GOME intercompared against the Dobson network: 1978-1998, J. Geophys. Res., 106 (D19), 23029-23042, 2001.

Boyd I.S., G.E. Bodeker, B.J. Connor, D.P.J. Swart, and E.J. Brinksma, An assessment of ECC ozonesondes operated using 1% and 0.5% KI cathode solutions at Lauder, New Zealand. Geophys.Res. Lett., 25 (13), 2409-2412, 1998.

Brinksma E.J., J.B. Bergwerff, G.E. Bodeker, K.F. Boersma, I.S. Boyd, B.J. Connor, J.F. de Haan, W. Hogervorst, J.W. Hovenier, A. Parrish, J.J. Tsou, J.M. Zawodny, and D.P.J. Swart, Validation of 3 years of ozone measurements over Network for the Detection of Stratospheric Change station Lauder, New Zealand, J. Geophys. Res., 105 (D13), 17291-17306, 2000.

Chernykh I.V., O.A. Alduchov, and R.E. Eskridge, Trends in low and high cloud boundaries and errors in height determination of cloud boundaries, Bull. Amer. Meteor. Soc., 82(9), 1941-1947, 2001.

Daumont D., J. Brion, J. Charbonnier and J. Malicet, Ozone UV Spectroscopy I: Absorption Cross-Sections at Room Temperature, J. Atmos. Chem., 15, 145-155, 1992.

Forster P.M., G.E. Bodeker, R. Schofield, S. Solomon, and D.W.J. Thompson, Effects of ozone cooling in the tropical lower stratosphere and upper troposphere, Geophys. Res. Lett., 34, L23813, doi:23810.21029/22007GL031994, 2007.

Fueglistaler, S. and Haynes, P.H., Control of interannual and longer-term variability of stratospheric water vapor, J. Geophys. Res., 110, D24108, doi:24110.21029/22005JD006019, 2005.

GCOS-92, Implementation plan for the global observing system for climate in support of the UNFCCC.WMO/TD No. 1219, 2004.

GCOS-112, GCOS Reference Upper-Air Network (GRUAN): Justification, requirements, siting and instrumentation options. WMO/TD-No.1379, 2007.

GCOS-121, GCOS Reference Upper Air Network (GRUAN): Report of the GRUAN Implementation Meeting, Lindenberg, Germany, 26-28 February 2008.

GCOS-122, Fourteenth Session of the GCOS/WCRP Atmospheric Observation Panel for Climate (AOPC-XIV), Geneva, 12-25 April 2008, WMO/TD No. 1436, 2008.

GCOS-131, Report of the first GCOS reference upper air network implementation and coordination meeting (GRUAN ICM-1). WMO/TD No. 1492, 2009.

GCOS-134, GRUAN Implementation Plan 2009-2013. WMO/TD No. 1506, 2009.

GCOS-138, Implementation Plan for the Global Observing System for Climate in support of the UNFCCC (2010 Update). WMO-TD/No. 1523.

GCOS-140, Report of the Second GCOS Reference Upper Air Network Implementation and Coordination Meeting (GRUAN ICM-2), WMO/TD No. 1526, 2009.

Gettelman A., W.J. Randel, F. Wu and S.T. Massie, Transport of water vapor in the tropical tropopause layer, Geophys. Res. Lett., 29, 10.1029/2001GL013818, 2002.

Haimberger, L., Homogenization of Radiosonde Temperature Time Series Using Innovation Statistics, J. Climate, 20, 1377-1403, 2007.

Immler F.J., J. Dykema, T. Gardiner, D.N. Whiteman, P.W. Thorne, and H. Vömel, Reference Quality Upper-Air Measurements: guidance for developing GRUAN data products. Atmospheric Measurement Techniques, 3, 1217–1231, 2010.

Khaliq M.N. and T.B.M.J. Ouarda, On the critical values of the standard normal homogeneity test (SNHT), Int. J. Climatol., 27, 681-687, 2007.

Leiterer, U., H. Dier and T. Naebert, Improvements in radiosonde humidity profiles using RS80/RS90 radiosondes of Vaisala,Contrib. Atmos. Phys., 70, 319–336, 1997.

McDermid I.S., J.B. Bergwerff, G.E. Bodeker, I.S. Boyd, E.J. Brinksma, B.J. Connor, R. Farmer, M.R. Gross, P. Kimvilakani, W.A. Matthews, T.J. McGee, F.T. Ormel, A. Parrish, U. Singh, D.P.J. Swart, and J.J. Tsou, OPAL: Network for the detection of stratospheric change ozone profiler assessment at Lauder, New Zealand 2. intercomparison of revised results. J. Geophys. Res., 103 (D22), 28693-28699, 1998.

McDermid I.S., J.B. Bergwerff, G.E. Bodeker, I.S. Boyd, E.J. Brinksma, B.J. Connor, R. Farmer, M.R. Gross, P. Kimvilakani, W.A. Matthews, T.J. McGee, F.T. Ormel, A. Parrish, U. Singh, D.P.J. Swart, J.J. Tsou, P.H. Wang, and J. Zawodny, OPAL: Network for the detection of stratospheric change ozone profiler assessment at Lauder, New Zealand 1. blind intercomparison. J. Geophys. Res., 103 (D22), 28683-28692, 1998.

Norris J.R., Multidecadal changes in near-global cloud cover and estimated cloud cover radiative forcing, J. Geophys. Res., 110, D08206, doi:08210.01029/02004JD005600, 2005.

Ohring G., B. Wielicki, R. Spencer, B. Emery, and R. Datla, Satellite Instrument Calibration for Measuring Global Climate Change: Report of a Workshop. Bull. Am. Meteorol. Soc., 86, 1303-1313, 2005.

Peter T.,C. Marcolli,P. Spichtinger, T. Corti, M.B. Baker, and T. Koop, When Dry Air Is Too Humid. Science, 314, 1399–1402, DOI: 10.1126/science.1135199, 2006.

Randel, W.J.; Shine, K.P.; Austin, J.; Barnett, J.; Claud, C.; Gillett, N.P.; Keckhut, P.; Langematz, U.; Lin, R.; Long, C.; Mears, C.; Miller, A.; Nash, J.; Seidel, D.J.; Thompson, D.W.J.; Wu, F. and Yoden, S., An update of observed stratospheric temperature trends, J. Geophys. Res., 114, D02107, doi:02110.01029/02008JD010421, 2009.

Rodgers, C. D.: Inverse Methods For Atmospheric Sounding: Theory and Practice, World Scientific Publishing Co. Pte. Ltd., 2, 240, 2000.

Rosenlof K.H., S.J. Oltmans, D. Kley, J.M. Russell III, E.-W. Chiou, W.P. Chu, D.G. Johnson, K. K. Kelly, H.A. Michelsen, G.E. Nedoluha, E.E. Remsberg, G.C. Toon, and M.P. McCormick, Stratospheric water vapour increases over the past half-century, Geophys. Res. Lett., 28 (7), 1195-1198, 2001.

Rossow W. B., and R. A. Schiffer, Advances in Understanding Clouds from ISCCP, Bull. Amer. Meteor. Soc., 80(11), 2261-2287, 1999.

Schmidlin, F.J., Derivation and application of temperature corrections for the United States radiosonde, in Symposium on Meteorological Observations and Instrumentations, 7th, New Orleans,LA, 14–18 January 1991, Preprints (A92-32051 12-47), Boston, MA, American Meteorological Society, 227–231, 1991.

Seidel D.J. and M. Free, Measurement Requirements for Climate Monitoring of Upper-Air Temperature Derived from Reanalysis Data, J. Climate, 19, 854-871, 2006.

Schubert S., D. Dee, S. Uppala, J. Woollen, J. Bates, and S. Worley, Report of the Workshop on “The Development of Improved Observational Data Sets for Reanalysis: Lessons learned and Future Directions”, 31 pp., http://gmao.gsfc.nasa.gov/pubs/docs/Schubert273.doc, 2006.

Seidel D.J., and I. Durre, Comments on “Trends in low and high cloud boundaries and errors in height determination of cloud boundaries”, Bull. Amer. Meteor. Soc., 84(2), 237-240, 2003.

Seidel D. J., F. H. Berger, H. J. Diamond, J. Dykema, D. Goodrich, F. Immler, W. Murray, T. Peterson, D. Sisterson, M. Sommer, P. Thorne, H. Vömel, and J. Wang, Reference upper-air observations for climate: rationale, progress, and plans, Bull. Amer. Meteor. Soc., 3, 361-369, 2009.

Shine K.P., M.S. Bourqui, P.M.D. Forster, S.H.E. Hare, U. Langematz, P. Braesicke, V. Grewe, M. Ponater, C. Schnadt, C.A. Smith, J.D. Haigh, J. Austin, N. Butchart, D.T. Shindell, W.J.
Randel, T. Nagashima, R.W. Portmann, S. Solomon, D.J. Seidel, J. Lanzante, S. Klein, V. Ramaswamy, and M.D. Schwarzkopf,A comparison of model-simulated trends in stratospheric temperatures. Quarterly Journal of the Royal Meteorological Society, 129 (590), 1565–1588, 2003.

Smit H.G.J., W. Straeter, B.J. Johnson, S.J. Oltmans, J. Davies, D.W. Tarasick, B. Hoegger, R. Stubi, F.J. Schmidlin, T. Northam, A.M. Thompson, J.C. Witte, I. Boyd, and F. Posny, Assessment of the performance of ECC-ozonesondes under quasi-flight conditions in the environmental simulation chamber: Insights from the Juelich Ozone Sonde Intercomparison Experiment (JOSIE). J. Geophys. Res., 112, D19306, doi:10.1029/2006JD007308, 2007.

Soden B.J., D.D. Turner, B.M. Lesht, and L.M. Miloshevich, An analysis of satellite, radiosonde, and lidar observations of upper tropospheric water vapor from the Atmospheric Radiation Measurement Program. J. Geophys. Res., 109, D04105, doi:10.1029/2003JD003828, 2004.

Solomon S., K.H. Rosenlof, R.W. Portmann, J.S. Daniel, S.M. Davis, T.J. Sanford and G.-K.Plattner, Contributions of Stratospheric Water Vapor to Decadal Changes in the Rate of Global Warming, Science, 327, 1219-1223, 2010.

Steiner, A.K.; Kirchengast, G.; Lackner, B.C.; Pirscher, B.; Borsche, M. and Foelsche, U., Atmospheric temperature change detection with GPS radio occultation 1995 to 2008, Geophys. Res. Lett., 36, L18702, doi:18710.11029/12009GL039777, 2009.

.

Thompson A.M., J.C. Witte, H.G.J. Smit, S.J. Oltmans, B.J. Johnson, V.W.J.H. Kirchhoff and F.J. Schmidlin, Southern Hemisphere Additional Ozonesondes (SHADOZ) 1998–2004 tropical ozone climatology: 3. Instrumentation, station-to-station variability, and evaluation with simulated flight profiles, J. Geophys. Res., 112, D03304, doi:10.1029/2005JD007042, 2007.

Tiao, G. C., G. C. Reinsel, D. Xu, J. H. Pedrick, X. Zhu, A. J. Miller, J. J. DeLuisi, C. L. Mateer, and D. J. Wuebbles (1990), Effects of autocorrelation and temporal sampling schemes on  estimates of trend and spatial correlation, J. Geophys. Res., 95(D12), 20507-20517.

Trenberth, The need for a new upper air baseline network, [details from Kevin to come in here], 2003.

Vömel H., J.E. Barnes, R.N. Forno, M. Fujiwara, F. Hasebe, S. Iwasaki, R. Kivi, N. Komala, E. Kyrö, T. Leblanc, B. Morel, S.-Y. Ogino, W.G. Read, S.C. Ryan, S. Saraspriya, H. Selkirk, M. Shiotani, J.V. Canossa and D.N. Whiteman, Validation of Aura Microwave Limb Sounder water vapor by balloonborne Cryogenic Frost point Hygrometer measurements, J. Geophys. Res., 112, D24S37, doi:10.1029/2007JD008698, 2007a.

Vömel H., D.E. David and K. Smith, Accuracy of tropospheric and stratospheric water vapor measurements by the cryogenic frost point hygrometer: Instrumental details and observations, J. Geophys. Res., 112, D08305, doi:10.1029/2006JD007224, 2007b.

Wang, J., D.J. Carlson, D.B. Parsons, T.F. Hock, D. Lauritsen, H.L. Cole, K. Beierle, and E. Chamberlain, Performance of operational radiosonde humidity sensors in direct comparison with a chilled mirror dew-point hygrometer and its climate implication. Geophys. Res. Lett., 30, 1860, doi:10.1029/2003GL016985, 2003.

Weatherhead, E. C., G. C. Reinsel, G. C. Tiao, X.-L. Meng, D. Choi, W.-K. Cheang, T. Keller, J. DeLuisi, D. J. Wuebbles, J. B. Kerr, A. J. Miller, S. J. Oltmans, and J. E. Frederick (1998), Factors affecting the detection of trends: Statistical considerations and applications to environmental data, J. Geophys. Res., 103(D14), 17149-17161.

Winker, D.M., and M.A. Vaughan, Vertical distribution of clouds over Hampton, Virginia observed by lidar under the ECLIPS and FIRE ETO programs, Atmos. Res., 34, 117-133, 1994.

WMO-No. 8, Guide to Meteorological Instruments and Methods of Observation, 7th edition, 2008.

WMO-No. 488, Guide to the Global Observing System, 2010.

World Meteorological Organization, 2011: WMO Intercomparison of High Quality Radiosonde Systems, Yangjiang, China (J. Nash, T. Oakley, H. Vömel and Li Wei) Instruments and Observing Methods Report No. 107, WMO/TD-No 1580, Geneva

Zhou S., M.A. Geller and M. Zhang, Cooling trend of the tropical cold point tropopause temperatures and its implications, J. Geophys.Res., 106, 1511-1522, 2001.

# Footnotes

00/12 UTC observations are no longer as important for NWP since 4D data assimilation is now more common. Where higher priority considerations require sites to measure at 00/12 UTC rather than 00/12 LST, this will not count against the site.

Available from http://www.dwd.de/bvbw/generator/DWDWWW/Content/Projekte/Gruan/Downloads/documents/gruan-td-1__data__management__v04__doc,templateId=raw,property=publicationFile.pdf/gruan-td-1_data_management_v04_doc.pdf

Available from http://www.dwd.de/bvbw/generator/DWDWWW/Content/Projekte/Gruan/Downloads/GRUAN__LC/gruan__data__policy,templateId=raw,property=publicationFile.pdf/gruan_data_policy.pdf

http://nco.sourceforge.net/

http://dq.arm.gov/

Trends reported in Chernykh have been questioned by Seidel and Durre (2003)